## 4-24
Failure Mode | Hazard | Hazardous Situation (HS) | failure causes | Harm | Severity | Probability | Risk Controls (RC) | Information for Safety
Air in line | Improper Flow or Infusion | Air infused into the body while connected to pump | Air in line due to broken, loose, unsealed fluid path | Death | 5 | 1 | Device designed and processed in a manner to be free from cracks, splits, fractures, and discoloration; Device designed remain securely attached to mating devices; Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use; Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector
Air in line | Improper Flow or Infusion | Air infused into the body while connected to pump | Air in line due to broken, loose, unsealed fluid path | Pulmonary Embolus; Stroke; Myocardial infarction | 4 | 1 | Device designed and processed in a manner to be free from cracks, splits, fractures, and discoloration; Device designed remain securely attached to mating devices; Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use; Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector
Air in line | Improper Flow or Infusion | Air infused into the body while connected to pump | Air in line due to broken, loose, unsealed fluid path | Mild to Moderate Respiratory Distress | 3 | 2 | Device designed and processed in a manner to be free from cracks, splits, fractures, and discoloration; Device designed remain securely attached to mating devices; Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use; Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector
Air in line | Improper Flow or Infusion | Air infused into the body while connected to pump | Air in line due to broken, loose, unsealed fluid path | Temporary discomfort | 1 | 2 | Device designed and processed in a manner to be free from cracks, splits, fractures, and discoloration; Device designed remain securely attached to mating devices; Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use; Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector
Air in line | Improper Flow or Infusion | Air infused into the body while connected to pump | Air in line due to improper priming | Death | 5 | 1 | Device designed to allow air to escape during priming | Labeling provides instructions to prime device before use; Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Air in line | Improper Flow or Infusion | Air infused into the body in gravity mode | Air in line due to broken, loose, unsealed fluid path | Mild to Moderate Respiratory Distress | 3 | 2 | Device designed and processed in a manner to be free from cracks, splits, fractures, and discoloration; Device designed remain securely attached to mating devices; Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use; Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector
Air in line | Improper Flow or Infusion | Air infused into the body in gravity mode | Air in line due to broken, loose, unsealed fluid path | Temporary discomfort | 1 | 2 | Device designed and processed in a manner to be free from cracks, splits, fractures, and discoloration; Device designed remain securely attached to mating devices; Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use; Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector
Air in line | Improper Flow or Infusion | Air infused into the body in gravity mode | Air in line due to improper priming | Mild to Moderate Respiratory Distress | 3 | 2 | Device designed to allow air to escape during priming | Labeling provides instructions to prime device before use; Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Allergenic substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Contact dermatitis (skin inflammation) with systemic allergic reaction; Eye inflammation | 3 | 3 | Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use
Allergenic substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Contact dermatitis with localized allergic reaction | 2 | 3 | Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use
Allergenic substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Temporary discomfort | 1 | 3 | Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use
Allergenic substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Fluid leak due to defective product (E.g. Crack, broken, holes, cuts, incorrect solvent bonding) | Contact dermatitis (skin inflammation) with systemic allergic reaction; Eye inflammation | 3 | 3 | Device designed not to leak during normal use; Device designed not to come apart during usage; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to reseat following 10 activations; Device designed to tolerate 10 activations to simulate a 24 hour period use life; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use
Allergenic substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Incompatible device (needle or non-compatible luer) inserted into Spiros | Contact dermatitis with localized allergic reaction | 2 | 3 | N/A | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Allergenic substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Patient movement causes set to disconnect at a connection point | Contact dermatitis (skin inflammation) with systemic allergic reaction; Eye inflammation | 3 | 3 | Device designed to prevent the unintended disconnection with mating device | N/A
Allergenic substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Set materials are not biocompatible | Temporary discomfort | 1 | 3 | Sterile barrier system and device materials should be compliant with ISO biocompatibility requirements; Materials of construction should comply with regulatory requirements | N/A
Allergenic Substances | Biocompatibility Issue | Infusion of allergenic substances | Allergenic substances in the fluid path | Death; Anaphylaxis | 5 | 1 | Device materials selected without the use of latex; Production controls designed to keep product non-pyrogenic; Sterile barrier system and device materials should be compliant with ISO biocompatibility requirements | Labeling instruct the user to follow aseptic technique in opening the package; Labeling includes precautions and instructions to disinfect before each access; Labeling includes information to inform user that the product is sterile and nonpyrogenic in unopened, undamaged package; Labeling includes caution to not use if package is damaged or opened prior to use; Labeling includes caution to not reuse or resterilize
Allergenic Substances | Biocompatibility Issue | Infusion of allergenic substances | Allergenic substances in the fluid path | Severe Respiratory distress; Stevens Johnson Syndrome | 4 | 1 | Device materials selected without the use of latex; Production controls designed to keep product non-pyrogenic; Sterile barrier system and device materials should be compliant with ISO biocompatibility requirements | Labeling instruct the user to follow aseptic technique in opening the package; Labeling includes precautions and instructions to disinfect before each access; Labeling includes information to inform user that the product is sterile and nonpyrogenic in unopened, undamaged package; Labeling includes caution to not use if package is damaged or opened prior to use; Labeling includes caution to not reuse or resterilize
Allergenic Substances | Biocompatibility Issue | Infusion of allergenic substances | Allergenic substances in the fluid path | Mild to Moderate Respiratory Distress | 3 | 2 | Device materials selected without the use of latex; Production controls designed to keep product non-pyrogenic; Sterile barrier system and device materials should be compliant with ISO biocompatibility requirements | Labeling instruct the user to follow aseptic technique in opening the package; Labeling includes precautions and instructions to disinfect before each access; Labeling includes information to inform user that the product is sterile and nonpyrogenic in unopened, undamaged package; Labeling includes caution to not use if package is damaged or opened prior to use; Labeling includes caution to not reuse or resterilize
Allergenic Substances | Biocompatibility Issue | Infusion of allergenic substances | Allergenic substances in the fluid path | Urticaria | 2 | 2 | Device materials selected without the use of latex; Production controls designed to keep product non-pyrogenic; Sterile barrier system and device materials should be compliant with ISO biocompatibility requirements | Labeling instruct the user to follow aseptic technique in opening the package; Labeling includes precautions and instructions to disinfect before each access; Labeling includes information to inform user that the product is sterile and nonpyrogenic in unopened, undamaged package; Labeling includes caution to not use if package is damaged or opened prior to use; Labeling includes caution to not reuse or resterilize
Allergenic Substances | Biocompatibility Issue | Infusion of allergenic substances | Allergenic substances in the fluid path | Temporary discomfort | 1 | 2 | Device materials selected without the use of latex; Production controls designed to keep product non-pyrogenic; Sterile barrier system and device materials should be compliant with ISO biocompatibility requirements | Labeling instruct the user to follow aseptic technique in opening the package; Labeling includes precautions and instructions to disinfect before each access; Labeling includes information to inform user that the product is sterile and nonpyrogenic in unopened, undamaged package; Labeling includes caution to not use if package is damaged or opened prior to use; Labeling includes caution to not reuse or resterilize
Allergenic Substances | Biocompatibility Issue | Infusion of allergenic substances | Loss of packaging integrity allows foreign material to enter into the sterile fluid path | Death; Anaphylaxis | 5 | 1 | Device and packaging designed to prevent protective cap from detaching during sterilization or shipping; Sterile barrier system shall be designed to keep the product sterile and contamination free | Labeling includes information to inform user that the product is sterile and nonpyrogenic in unopened, undamaged package; Labeling includes caution to not use if package is damaged or opened prior to use; Labeling includes caution to not reuse or resterilize; Labeling includes instruction to use aseptic technique in opening of package and removal of Spiros from packaging; Labeling instruct the user to follow aseptic technique in opening the package; Labeling includes precautions and instructions to disinfect before each access
Allergenic Substances | Biocompatibility Issue | Infusion of allergenic substances | Set materials are incompatible to prolonged exposure to disinfectant caps | Severe Respiratory distress; Stevens Johnson Syndrome | 4 | 1 | Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcoho | Labeling includes a precaution for the user to change the device at recommended intervals per facility protocol
Allergenic Substances | Biocompatibility Issue | Infusion of allergenic substances | Set materials are not biocompatible | Mild to Moderate Respiratory Distress | 3 | 2 | Materials of construction should comply with regulatory requirements; Sterile barrier system and device materials should be compliant with ISO biocompatibility requirements | N/A
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Severe Burn; Sepsis; Severe Pain | 4 | 3 | Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Moderate Burn; Moderate Pain | 3 | 3 | Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Mild Burn; Mild Pain | 2 | 3 | Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Temporary discomfort | 1 | 3 | Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Fluid leak due to defective product (E.g. Crack, broken, holes, cuts, incorrect solvent bonding) | Severe Burn; Sepsis; Severe Pain | 4 | 3 | Device designed not to leak during normal use; Device designed not to come apart during usage; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to reseat following 10 activations; Device designed to tolerate 10 activations to simulate a 24 hour period use life; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution statement for proper accessing Spiros by pushing straight on without angled or sliding entry and do not twist or bend connector when accessing; Labeling includes a precaution statement to inspect for leaks before use
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Incompatible device (needle or non-compatible luer) inserted into Spiros | Moderate Burn; Moderate Pain | 3 | 3 | N/A | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Patient movement causes set to disconnect at a connection point | Mild Burn; Mild Pain | 2 | 3 | Device designed to prevent the unintended disconnection with mating device | N/A
Chemical-Other toxin | Biocompatibility Issue | Intravenous exposure to non-hemocompatible materials | Non-hemocompatible material in the fluid path | Death; Cardiopulmonary Arrest | 5 | 1 | Sterile barrier system and device materials should be compliant with ISO biocompatibility requirements; Materials of construction should comply with regulatory requirements | N/A
Chemical-Other toxin | Biocompatibility Issue | Intravenous exposure to non-hemocompatible materials | Non-hemocompatible material in the fluid path | Severe Hemolytic Anemia; Severe Hyperkalemia; Hypotension; Congestive Heart Failure; Acute Kidney Injury | 4 | 1 | Sterile barrier system and device materials should be compliant with ISO biocompatibility requirements; Materials of construction should comply with regulatory requirements | N/A
Chemical-Other toxin | Biocompatibility Issue | Intravenous exposure to non-hemocompatible materials | Non-hemocompatible material in the fluid path | Mild to Moderate Hemolytic Anemia; Mild to Moderate Hyperkalemia; Fever | 3 | 1 | Sterile barrier system and device materials should be compliant with ISO biocompatibility requirements; Materials of construction should comply with regulatory requirements | N/A
Chemical-Other toxin | Biocompatibility Issue | Intravenous exposure to non-hemocompatible materials | Non-hemocompatible material in the fluid path | Temporary discomfort | 1 | 1 | Sterile barrier system and device materials should be compliant with ISO biocompatibility requirements; Materials of construction should comply with regulatory requirements | N/A
BPA leached | Biocompatibility Issue | Intravenous exposure to BPA | BPA from the device material leached into the fluid path | Organ damage | 4 | 1 | Materials of construction should comply with regulatory requirements | Labeling provides information that the device is not made with DEHP
BPA leached | Biocompatibility Issue | Intravenous exposure to BPA | BPA from the device material leached into the fluid path | Reproductive system dysfunction | 3 | 1 | Materials of construction should comply with regulatory requirements | Labeling provides information that the device is not made with DEHP
BPA leached | Biocompatibility Issue | Intravenous exposure to BPA | BPA from the device material leached into the fluid path | Temporary discomfort | 1 | 1 | Materials of construction should comply with regulatory requirements | Labeling provides information that the device is not made with DEHP
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Fluid path contains toxins (E.g. systemic, mutagens, carcinogens) | Death; Cardiopulmonary Arrest | 5 | 2 | Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol; Production controls designed to keep product non-pyrogenic | Labeling includes information to inform user that the product is sterile and nonpyrogenic in unopened, undamaged package; Labeling includes caution to not use if package is damaged or opened prior to use; Labeling includes caution to not reuse or resterilize; Labeling includes instruction to use aseptic technique in opening of package and removal of Spiros from packaging
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Fluid path contains toxins (E.g. systemic, mutagens, carcinogens) | Paralysis; Organ failure/injury; Malignancy; Severe Hypotension | 4 | 2 | Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol; Production controls designed to keep product non-pyrogenic | Labeling includes information to inform user that the product is sterile and nonpyrogenic in unopened, undamaged package; Labeling includes caution to not use if package is damaged or opened prior to use; Labeling includes caution to not reuse or resterilize; Labeling includes instruction to use aseptic technique in opening of package and removal of Spiros from packaging
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Fluid path contains toxins (E.g. systemic, mutagens, carcinogens) | Mild to Moderate Hypotension; Moderate Inflammation (e.g. Phlebitis, Organ); Moderate Pain; Fever | 3 | 2 | Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol; Production controls designed to keep product non-pyrogenic | Labeling includes information to inform user that the product is sterile and nonpyrogenic in unopened, undamaged package; Labeling includes caution to not use if package is damaged or opened prior to use; Labeling includes caution to not reuse or resterilize; Labeling includes instruction to use aseptic technique in opening of package and removal of Spiros from packaging
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Fluid path contains toxins (E.g. systemic, mutagens, carcinogens) | Mild Inflammation (e.g. Phlebitis, Organ); Mild Pain | 2 | 1 | Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol; Production controls designed to keep product non-pyrogenic | Labeling includes information to inform user that the product is sterile and nonpyrogenic in unopened, undamaged package; Labeling includes caution to not use if package is damaged or opened prior to use; Labeling includes caution to not reuse or resterilize; Labeling includes instruction to use aseptic technique in opening of package and removal of Spiros from packaging
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Fluid path contains toxins (E.g. systemic, mutagens, carcinogens) | Temporary discomfort | 1 | 1 | Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol; Production controls designed to keep product non-pyrogenic | Labeling includes information to inform user that the product is sterile and nonpyrogenic in unopened, undamaged package; Labeling includes caution to not use if package is damaged or opened prior to use; Labeling includes caution to not reuse or resterilize; Labeling includes instruction to use aseptic technique in opening of package and removal of Spiros from packaging
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Cap falls off allowing foreign material to enter into the sterile fluid path | Death; Cardiopulmonary Arrest | 5 | 2 | Device and packaging designed to prevent protective cap from detaching during sterilization or shipping; Protective caps designed to be able to be securely attached, but easily removed | Labeling includes information to inform user that the product is sterile and nonpyrogenic in unopened, undamaged package; Labeling includes caution to not use if package is damaged or opened prior to use; Labeling includes caution to not reuse or resterilize; Labeling includes instruction to use aseptic technique in opening of package and removal of Spiros from packaging
Open fluid path | Fluid Leak | Patient blood loss | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Death | 5 | 1 | Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use
Open fluid path | Fluid Leak | Patient blood loss | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Hypovolemic shock | 4 | 2 | Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use
Open fluid path | Fluid Leak | Patient blood loss | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Mild to Moderate Hypotension; Mild to Moderate Hypovolemia | 3 | 2 | Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use
Open fluid path | Fluid Leak | Patient blood loss | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Temporary Discomfort | 1 | 3 | Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use
Open fluid path | Fluid Leak | Patient blood loss | Blood loss due to defective product (E.g. Crack, broken, holes, cuts, incorrect solvent bonding) | Mild to Moderate Hypotension; Mild to Moderate Hypovolemia | 3 | 2 | Device designed and processed in a manner to be free from cracks, splits, fractures, and discoloration; Device designed not to leak during normal use; Device designed not to come apart during usage; Device designed to prevent the unintended disconnection with mating device; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol; Device designed to prevent the leak when subject to 36'' head height back pressure | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use
Open fluid path | Fluid Leak | Patient blood loss | Difficult / Unable to connect / disconnect the device with other mating devices (e.g. Container, access site) | Hypovolemic shock | 4 | 2 | Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use
Open fluid path | Fluid Leak | Patient blood loss | Incompatible device (needle or non-compatible luer) inserted into Spiros | Mild to Moderate Hypotension; Mild to Moderate Hypovolemia | 3 | 2 | N/A | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Open fluid path | Fluid Leak | Patient blood loss | Patient movement causes set to disconnect at a connection point | Temporary Discomfort | 1 | 3 | Device designed to prevent the unintended disconnection with mating device | N/A
Open fluid path | Fluid Leak | Patient blood loss | Set materials are incompatible to prolonged exposure to disinfectant caps | Mild to Moderate Hypotension; Mild to Moderate Hypovolemia | 3 | 2 | Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution for the user to change the device at recommended intervals per facility protocol
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 2 | Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Mild to Moderate Pain | 2 | 2 | Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Temporary discomfort/ User discomfort | 1 | 2 | Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Difficult / Unable to connect / disconnect the device with other mating devices (e.g. Container, access site) | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 2 | Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Difficult / Unable to open the packaging | Mild to Moderate Pain | 2 | 2 | Sterile barrier system shall provide product in an aseptic manner | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Difficult / Unable to remove Cap / cover / hood from the product | Mild to Moderate Pain | 2 | 2 | Protective caps designed to be able to be securely attached, but easily removed | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Fluid leak due to defective product (E.g. Crack, broken, holes, cuts, incorrect solvent bonding) | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 2 | Device designed not to leak during normal use; Device designed not to come apart during usage; Device designed to prevent the leak when subject to 36'' head height back pressure; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution statement for proper accessing Spiros by pushing straight on without angled or sliding entry and do not twist or bend connector when accessing; Labeling includes a precaution statement to inspect for leaks before use
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Fluid path occlusion | Temporary discomfort/ User discomfort | 1 | 2 | Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Incompatible device (needle or non-compatible luer) inserted into Spiros | Temporary discomfort/ User discomfort | 1 | 2 | N/A | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | No fluid flow | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 2 | Device designed to maintain required flow rate; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Problem identifying / selecting set | Temporary discomfort/ User discomfort | 1 | 2 | Labeling designed to meet customer and regulatory requirements; Labelling designed to compatible with sterilization, shipping and storage processes | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Problem setting up / priming the device | Temporary discomfort/ User discomfort | 1 | 2 | Device designed and processed in a manner to be free from cracks, splits, fractures, and discoloration; Device designed remain securely attached to mating devices; Device designed to allow air to escape during priming; Protective caps designed to be able to be securely attached, but easily removed | Labeling includes a precaution statement to tighten and secure non-bonded connections and inspect for leaks before use, and to not attempt to tighten or disconnect if connections are bonded.
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Set materials are incompatible to prolonged exposure to disinfectant caps | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe pain | 3 | 2 | Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution for the user to change the device at recommended intervals per facility protocol
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Product appears unsafe to use (E.g.. Open package, Missing / Loose / Damaged components, Discoloration, Visible particulate etc.) | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe pain | 3 | 2 | Device and packaging designed to prevent protective cap from detaching during sterilization or shipping; Device designed and processed in a manner to be free from cracks, splits, fractures, and discoloration; Device designed to function until expiry date; Sterile barrier system shall be designed to keep the product sterile and contamination free | Labeling includes caution to use by expiration date on package label; Labeling includes caution to not use if package is damaged or opened prior to use; Labeling includes information to inform user that the product is sterile and nonpyrogenic in unopened, undamaged package; Labeling includes caution to not use if package is damaged or opened prior to use; Labeling includes caution to not reuse or resterilize; Labeling includes instruction to use aseptic technique in opening of package and removal of Spiros from packaging
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Catheter pull out when exposed to MRI | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe pain | 3 | 2 | Material should be non-ferrous | N/A
setup delayed | Device Difficult to Setup or Prepare | Blood/Blood product rendered unusable during setup | Incompatible device (needle or non-compatible luer) inserted into Spiros | Death Cardiopulmonary Arrest | 5 | 1 | N/A | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
setup delayed | Device Difficult to Setup or Prepare | Blood/Blood product rendered unusable during setup | Incompatible device (needle or non-compatible luer) inserted into Spiros | Stroke | 4 | 1 | N/A | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
setup delayed | Device Difficult to Setup or Prepare | Blood/Blood product rendered unusable during setup | Incompatible device (needle or non-compatible luer) inserted into Spiros | Mild to Moderate Hypotension; Mild to Moderate Tachycardia | 3 | 2 | N/A | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
setup delayed | Device Difficult to Setup or Prepare | Blood/Blood product rendered unusable during setup | Incompatible device (needle or non-compatible luer) inserted into Spiros | Temporary discomfort | 1 | 1 | N/A | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Air in line due to broken, loose, unsealed fluid path | Mild to Moderate Hypotension; Mild to Moderate Tachycardia | 3 | 2 | Device designed and processed in a manner to be free from cracks, splits, fractures, and discoloration; Device designed to prevent the unintended disconnection with mating device; Device designed remain securely attached to mating devices | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Air in line due to broken, loose, unsealed fluid path | Pain, Mild | 2 | 2 | N/A | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Air in line due to broken, loose, unsealed fluid path | Temporary discomfort | 1 | 2 | N/A | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use; Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Temporary discomfort | 1 | 2 | Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Difficult / Unable to connect / disconnect the device with other mating devices (e.g. Container, access site) | Temporary discomfort | 1 | 2 | Device designed to prevent the unintended disconnection with mating device | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Difficult / Unable to open the packaging | Temporary discomfort | 1 | 2 | Sterile barrier system shall provide product in an aseptic manner | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Fluid leak due to defective product (E.g. Crack, broken, holes, cuts, incorrect solvent bonding) | Temporary discomfort | 1 | 2 | Device designed not to leak during normal use; Device designed not to come apart during usage; Device designed to prevent the leak when subject to 36'' head height back pressure; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution statement for proper accessing Spiros by pushing straight on without angled or sliding entry and do not twist or bend connector when accessing; Labeling includes a precaution statement to inspect for leaks before use
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Problem identifying / selecting set | Mild Pain | 2 | 2 | Labeling designed to meet customer and regulatory requirements; Labelling designed to compatible with sterilization, shipping and storage processes | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Problem setting up / priming the device | Mild Pain | 2 | 2 | Device designed and processed in a manner to be free from cracks, splits, fractures, and discoloration; Device designed remain securely attached to mating devices; Device designed to allow air to escape during priming; Protective caps designed to be able to be securely attached, but easily removed | Labeling includes a precaution statement to tighten and secure non-bonded connections and inspect for leaks before use, and to not attempt to tighten or disconnect if connections are bonded.
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Product appears unsafe to use (E.g.. Open package, Missing / Loose / Damaged components, Discoloration, Visible particulate etc.) | Mild to Moderate Hypotension; Mild to Moderate Tachycardia | 3 | 2 | Device and packaging designed to prevent protective cap from detaching during sterilization or shipping; Device designed and processed in a manner to be free from cracks, splits, fractures, and discoloration; Device designed to function until expiry date; Sterile barrier system shall be designed to keep the product sterile and contamination free | Labeling includes information to inform user that the product is sterile and nonpyrogenic in unopened, undamaged package; Labeling includes caution to not use if package is damaged or opened prior to use; Labeling includes caution to not reuse or resterilize; Labeling includes instruction to use aseptic technique in opening of package and removal of Spiros from packaging; Labeling includes caution to use by expiration date on package label;
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Death; Cardiopulmonary Arrest | 5 | 1 | Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Severe Arrhythmias; Severe Hypertension; Severe Hypotension; Status Epilepticus | 4 | 2 | Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 2 | Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Mild to Moderate Pain | 2 | 2 | Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Temporary discomfort | 1 | 3 | Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Difficult / Unable to connect / disconnect the device with other mating devices (e.g. Container, access site) | Death; Cardiopulmonary Arrest | 5 | 1 | Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Difficult / Unable to connect / disconnect the device with other mating devices (e.g. Container, access site) | Severe Arrhythmias; Severe Hypertension; Severe Hypotension; Status Epilepticus | 4 | 2 | Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Difficult / Unable to connect / disconnect the device with other mating devices (e.g. Container, access site) | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 2 | Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Fluid leak due to defective product (E.g. Crack, broken, holes, cuts, incorrect solvent bonding) | Mild to Moderate Pain | 2 | 2 | Device designed not to leak during normal use; Device designed not to come apart during usage; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to reseat following 10 activations; Device designed to tolerate 10 activations to simulate a 24 hour period use life; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution statement for proper accessing Spiros by pushing straight on without angled or sliding entry and do not twist or bend connector when accessing; Labeling includes a precaution statement to inspect for leaks before use
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Fluid leak due to defective product (E.g. Crack, broken, holes, cuts, incorrect solvent bonding) | Temporary discomfort | 1 | 3 | Device designed not to leak during normal use; Device designed not to come apart during usage; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to reseat following 10 activations; Device designed to tolerate 10 activations to simulate a 24 hour period use life; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution statement for proper accessing Spiros by pushing straight on without angled or sliding entry and do not twist or bend connector when accessing; Labeling includes a precaution statement to inspect for leaks before use
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Fluid path occlusion | Death; Cardiopulmonary Arrest | 5 | 1 | Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Incompatible device (needle or non-compatible luer) inserted into Spiros | Mild to Moderate Pain | 2 | 2 | N/A | N/A
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Patient movement causes set to disconnect at a connection point | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 2 | Device designed to prevent the unintended disconnection with mating device | N/A
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Set materials are incompatible to prolonged exposure to disinfectant caps | Death | 5 | 1 | Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution for the user to change the device at recommended intervals per facility protocol
Blood clot | Occlusion within Device | Embolization into the vascular system | Blood clot introduced into the device | Death; Cardiopulmonary Arrest | 5 | 1 | Device designed to maintain required flow rate | N/A
Blood clot | Occlusion within Device | Embolization into the vascular system | Blood clot introduced into the device | Pulmonary embolism; Myocardial infarction; Stroke | 4 | 1 | Device designed to maintain required flow rate | N/A
Blood clot | Occlusion within Device | Embolization into the vascular system | Blood clot introduced into the device | Extremity swelling Pain; Moderate Paresthesias | 3 | 1 | Device designed to maintain required flow rate | N/A
Blood clot | Occlusion within Device | Embolization into the vascular system | Blood clot introduced into the device | Thrombophlebitis or pain | 2 | 1 | Device designed to maintain required flow rate | N/A
Blood clot | Occlusion within Device | Embolization into the vascular system | Blood clot introduced into the device | Temporary Discomfort | 1 | 1 | Device designed to maintain required flow rate | N/A
particulate presented | Particulates | Embolization into the vascular system | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Death; Cardiopulmonary Arrest | 5 | 1 | Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use
particulate presented | Particulates | Embolization into the vascular system | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Pulmonary embolism; Myocardial infarction; Stroke | 4 | 2 | Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use
particulate presented | Particulates | Embolization into the vascular system | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Extremity swelling; Moderate Pain; Paresthesias | 3 | 2 | Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use
particulate presented | Particulates | Embolization into the vascular system | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Thrombophlebitis; Mild Pain | 2 | 2 | Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use
particulate presented | Particulates | Embolization into the vascular system | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Temporary Discomfort | 1 | 3 | Device designed to prevent the unintended disconnection with mating device | Labeling includes a general warning on the use of a damaged device or improper use and to inspect device and read instructions for use prior to usage; Labeling includes precaution for the user to properly tighten and do not disconnect bonded connections. Doing so may damage or break the connection; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to inspect for leaks before use
particulate presented | Particulates | Embolization into the vascular system | Fluid leak due to defective product (E.g. Crack, broken, holes, cuts, incorrect solvent bonding) | Thrombophlebitis; Mild Pain | 2 | 2 | Device designed not to leak during normal use; Device designed not to come apart during usage; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to reseat following 10 activations; Device designed to tolerate 10 activations to simulate a 24 hour period use life; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution statement for proper accessing Spiros by pushing straight on without angled or sliding entry and do not twist or bend connector when accessing; labeling includes a precaution statement to inspect for leaks before use
particulate presented | Particulates | Embolization into the vascular system | Incompatible device (needle or non-compatible luer) inserted into Spiros | Pulmonary embolism; Myocardial infarction; Stroke | 4 | 2 | N/A | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
particulate presented | Particulates | Embolization into the vascular system | Cap falls off allowing foreign material to enter into the sterile fluid path | Death | 5 | 1 | Protective caps designed to be able to be securely attached, but easily removed | Labeling includes information to inform user that the product is sterile and nonpyrogenic in unopened, undamaged package; Labeling includes caution to not use if package is damaged or opened prior to use; Labeling includes caution to not reuse or resterilize; Labeling includes instruction to use aseptic technique in opening of package and removal of Spiros from packaging
particulate presented | Particulates | Embolization into the vascular system | Particulate in the fluid path | Pulmonary embolism; Myocardial infarction; Stroke | 4 | 2 | Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol; Products are manufactured under conditions that minimize particulate contamination; Sterile barrier system shall be designed to keep the product sterile and contamination free | N/A
particulate presented | Particulates | Embolization into the vascular system | Set materials are incompatible to prolonged exposure to disinfectant caps | Pulmonary embolism; Myocardial infarction; Stroke | 4 | 2 | Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution for the user to change the device at recommended intervals per facility protocol
particulate presented | Particulates | Infusing particulates outside of the vessel | Cap falls off allowing foreign material to enter into the sterile fluid path | Edema; Local tissue necrosis; Severe Pain | 4 | 1 | Protective caps designed to be able to be securely attached, but easily removed | Labeling includes information to inform user that the product is sterile and nonpyrogenic in unopened, undamaged package; Labeling includes caution to not use if package is damaged or opened prior to use; Labeling includes caution to not reuse or resterilize; Labeling includes instruction to use aseptic technique in opening of package and removal of Spiros from packaging
particulate presented | Particulates | Infusing particulates outside of the vessel | Cap falls off allowing foreign material to enter into the sterile fluid path | Hematoma; Edema; Moderate Pain | 3 | 2 | Protective caps designed to be able to be securely attached, but easily removed | Labeling includes information to inform user that the product is sterile and nonpyrogenic in unopened, undamaged package; Labeling includes caution to not use if package is damaged or opened prior to use; Labeling includes caution to not reuse or resterilize; Labeling includes instruction to use aseptic technique in opening of package and removal of Spiros from packaging
particulate presented | Particulates | Infusing particulates outside of the vessel | Cap falls off allowing foreign material to enter into the sterile fluid path | Burning sensation; Edema; Mild Pain | 2 | 2 | Protective caps designed to be able to be securely attached, but easily removed | Labeling includes information to inform user that the product is sterile and nonpyrogenic in unopened, undamaged package; Labeling includes caution to not use if package is damaged or opened prior to use; Labeling includes caution to not reuse or resterilize; Labeling includes instruction to use aseptic technique in opening of package and removal of Spiros from packaging
particulate presented | Particulates | Infusing particulates outside of the vessel | Cap falls off allowing foreign material to enter into the sterile fluid path | Temporary discomfort | 1 | 2 | Protective caps designed to be able to be securely attached, but easily removed | Labeling includes information to inform user that the product is sterile and nonpyrogenic in unopened, undamaged package; Labeling includes caution to not use if package is damaged or opened prior to use; Labeling includes caution to not reuse or resterilize; Labeling includes instruction to use aseptic technique in opening of package and removal of Spiros from packaging
particulate presented | Particulates | Infusing particulates outside of the vessel | Particulate in the fluid path | Hematoma; Edema; Moderate Pain | 3 | 2 | Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol; Products are manufactured under conditions that minimize particulate contamination; Sterile barrier system shall be designed to keep the product sterile and contamination free | N/A
Inappropriate disposal of contaminated sets | Environmental Control or Utility Issue | Environmental exposure to toxic substances from contaminated sets | Contaminated sets contains toxic materials | Moderate toxic effects such as fatigue, sleepiness, headaches, eye irritation, skin rashes | 3 | 1 | N/A | Product labeling contains precautions to prepare hazardous drugs according to current, recognized guidelines; Product labeling contains instructions for proper disposal per facility protocol
Inappropriate disposal of contaminated sets | Environmental Control or Utility Issue | Environmental exposure to toxic substances from contaminated sets | Contaminated sets contains toxic materials | Mild toxic effects such as fatigue, sleepiness, headaches, eye irritation, skin rashes | 2 | 1 | N/A | Product labeling contains precautions to prepare hazardous drugs according to current, recognized guidelines; Product labeling contains instructions for proper disposal per facility protocol
Inappropriate disposal of contaminated sets | Environmental Control or Utility Issue | Environmental exposure to toxic substances from contaminated sets | Contaminated sets contains toxic materials | Temporary discomfort | 1 | 1 | N/A | Product labeling contains precautions to prepare hazardous drugs according to current, recognized guidelines; Product labeling contains instructions for proper disposal per facility protocol
Inappropriate disposal of contaminated sets | Environmental Control or Utility Issue | Environmental exposure to toxic substances from contaminated sets | Proper disposal procedure not followed | Moderate toxic effects such as fatigue, sleepiness, headaches, eye irritation, skin rashes | 3 | 1 | N/A | Product labeling contains precautions to prepare hazardous drugs according to current, recognized guidelines; Product labeling contains instructions for proper disposal per facility protocol
Wrong route | Misconnection | Incorrect setup of epidural therapy | Fluid infused through wrong route (Misconnection) | Death; Cardiopulmonary Arrest | 5 | 2 | Luer connectors are designed to not connect with non-IV application sets | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Wrong route | Misconnection | Incorrect setup of epidural therapy | Fluid infused through wrong route (Misconnection) | Severe Arrhythmias; Severe Respiratory Depression; Severe Hypotension; Abnormal mental status- Coma; Status Epilepticus | 4 | 1 | Luer connectors are designed to not connect with non-IV application sets | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Wrong route | Misconnection | Incorrect setup of epidural therapy | Fluid infused through wrong route (Misconnection) | Mild to Moderate Arrhythmias; Mild to Moderate Respiratory Depression; Mild to Moderate Hypotension; Seizures; CNS irritability, tremors, or restlessness | 3 | 2 | Luer connectors are designed to not connect with non-IV application sets | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Wrong route | Misconnection | Incorrect setup of epidural therapy | Fluid infused through wrong route (Misconnection) | Temporary discomfort | 1 | 1 | Luer connectors are designed to not connect with non-IV application sets | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Wrong Route | Misconnection | Incorrect setup of parenteral therapy | Fluid infused through wrong route (Misconnection) | Death; Cardiopulmonary Arrest | 5 | 1 | Luer connectors are designed to not connect with non-IV application sets | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Wrong Route | Misconnection | Incorrect setup of parenteral therapy | Fluid infused through wrong route (Misconnection) | Severe Arrhythmias; Severe Respiratory Distress or Depression; Severe Hypotension; Severe Hypertension; Severe Hyperglycemia; Severe Hypoglycemia | 4 | 1 | Luer connectors are designed to not connect with non-IV application sets | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Wrong Route | Misconnection | Incorrect setup of parenteral therapy | Fluid infused through wrong route (Misconnection) | Moderate Arrhythmias; Moderate Respiratory Distress or Depression; Moderate Hypotension; Moderate Hypertension; Moderate Hyperglycemia; Moderate Hypoglycemia | 3 | 1 | Luer connectors are designed to not connect with non-IV application sets | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Wrong Route | Misconnection | Incorrect setup of parenteral therapy | Fluid infused through wrong route (Misconnection) | Mild Arrhythmias; Mild Respiratory Distress or Depression; Mild Hypotension; Mild Hypertension; Mild Hyperglycemia; Mild Hypoglycemia | 2 | 2 | Luer connectors are designed to not connect with non-IV application sets | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Wrong Route | Misconnection | Incorrect setup of parenteral therapy | Fluid infused through wrong route (Misconnection) | Temporary discomfort | 1 | 2 | Luer connectors are designed to not connect with non-IV application sets | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Wrong route | Misconnection | Incorrect access connection | Fluid infused through wrong route (Misconnection) | Death; Cardiopulmonary Arrest | 5 | 1 | Luer connectors are designed to not connect with non-IV application sets | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Wrong route | Misconnection | Incorrect access connection | Fluid infused through wrong route (Misconnection) | Stroke; Organ injury or failure; Pulmonary embolism; Severe Respiratory distress; Severe Hypertension; Severe Hypotension; Severe Hyperglycemia; Severe Hypoglycemia | 4 | 2 | Luer connectors are designed to not connect with non-IV application sets | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Wrong route | Misconnection | Incorrect access connection | Fluid infused through wrong route (Misconnection) | Mild to Moderate Respiratory Distress; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Moderate Hyperglycemia; Moderate Hypoglycemia | 3 | 2 | Luer connectors are designed to not connect with non-IV application sets | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Wrong route | Misconnection | Incorrect access connection | Fluid infused through wrong route (Misconnection) | Mild Hyperglycemia; Mild Hypoglycemia | 2 | 1 | Luer connectors are designed to not connect with non-IV application sets | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Wrong route | Misconnection | Incorrect access connection | Fluid infused through wrong route (Misconnection) | Temporary discomfort | 1 | 1 | Luer connectors are designed to not connect with non-IV application sets | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Aseptic techniques not followed | Death; Cardiopulmonary Arrest | 5 | 2 | N/A | Labeling instruct the user to follow aseptic technique in opening the package; Labeling includes precautions and instructions to disinfect before each access
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Aseptic techniques not followed | Septic shock; Organ injury | 4 | 2 | N/A | Labeling instruct the user to follow aseptic technique in opening the package; Labeling includes precautions and instructions to disinfect before each access
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Aseptic techniques not followed | Bacteremia; Moderate Pain | 3 | 2 | N/A | Labeling instruct the user to follow aseptic technique in opening the package; Labeling includes precautions and instructions to disinfect before each access
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Aseptic techniques not followed | Mild Phlebitis; Local reaction/ infection; Mild Pain | 2 | 3 | N/A | Labeling instruct the user to follow aseptic technique in opening the package; Labeling includes precautions and instructions to disinfect before each access
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Aseptic techniques not followed | Temporary discomfort | 1 | 2 | N/A | Labeling instruct the user to follow aseptic technique in opening the package; Labeling includes precautions and instructions to disinfect before each access
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Device not sterilized | Death | 5 | 2 | The packaging system and labeling system are designed to allow for worst case sterilization | Labeling includes information to inform user that the product is sterile and nonpyrogenic in unopened, undamaged package; Labeling includes caution to not use if package is damaged or opened prior to use; Labeling includes caution to not reuse or resterilize; Labeling includes instruction to use aseptic technique in opening of package and removal of Spiros from packaging
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Bad connection / disconnection between container / set / Other mating devices (E.g. Loose connection, improper mating) | Bacteremia; Moderate Pain | 3 | 2 | Device designed to prevent the unintended disconnection with mating device | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Fluid path sterility compromised due to broken, loose or unsealed fluid path | Bacteremia; Moderate Pain | 3 | 2 | Device designed not to leak during normal use; Device designed not to come apart during usage; Device designed to function until expiry date; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the unintended disconnection with mating device; Device designed to reseat following 10 activations; Device designed to tolerate 10 activations to simulate a 24 hour period use life; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes information to inform user that the product is sterile and nonpyrogenic in unopened, undamaged package; Labeling includes caution to not use if package is damaged or opened prior to use; Labeling includes caution to not reuse or resterilize; Labeling includes instruction to use aseptic technique in opening of package and removal of Spiros from packaging; Labeling includes caution to use by expiration date on package label; Labeling instruct the user to follow aseptic technique in opening the package; Labeling includes precautions and instructions to disinfect before each access; Labeling includes a precaution statement for proper accessing Spiros by pushing straight on without angled or sliding entry and do not twist or bend connector when accessing; Labeling includes a precaution statement to inspect for leaks before use
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Incompatible device (needle or non-compatible luer) inserted into Spiros | Bacteremia; Moderate Pain | 3 | 2 | N/A | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Reuse of the same device on multiple patients | Bacteremia; Moderate Pain | 3 | 2 | Labeling designed to meet customer and regulatory requirements; Packaging designed to not be reusable/resealable | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA); Labeling provides information to the user that the device is for single patient use only; Labeling provide cautions not reuse product which are labeled for single use only
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Set materials are incompatible to prolonged exposure to disinfectant caps | Bacteremia; Moderate Pain | 3 | 2 | Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution for the user to change the device at recommended intervals per facility protocol
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Product is used for a longer period than specified | Bacteremia; Moderate Pain | 3 | 2 | N/A | Labeling includes a precaution for the user to change the device at recommended intervals per facility protocol
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Sterile field is compromised | Bacteremia; Moderate Pain | 3 | 2 | Sterile barrier system shall provide product in an aseptic manner | Labeling includes information to inform user that the product is sterile and nonpyrogenic in unopened, undamaged package; Labeling includes caution to not use if package is damaged or opened prior to use; Labeling includes caution to not reuse or resterilize; Labeling instruct the user to follow aseptic technique in opening the package; Labeling includes precautions and instructions to disinfect before each access
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Sterility compromised due to loss of packaging integrity | Temporary discomfort | 1 | 2 | Device designed to function until expiry date; Sterile barrier system shall be designed to keep the product sterile and contamination free | Labeling includes caution to use by expiration date on package label; Labeling includes caution to not use if package is damaged or opened prior to use; Labeling includes information to inform user that the product is sterile and nonpyrogenic in unopened, undamaged package; Labeling includes caution to not use if package is damaged or opened prior to use; Labeling includes caution to not reuse or resterilize; Labeling instruct the user to follow aseptic technique in opening the package; Labeling includes precautions and instructions to disinfect before each access
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Inadequate sterilization of the set | Bacteremia; Moderate Pain | 3 | 2 | Device designed to meet sterility requirements | N/A
Biological contaminants | Device Contamination with Biological Material | Skin/Mucous membranes exposure to biological contaminants | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Death; Cardiopulmonary Arrest | 5 | 1 | Device designed to prevent the unintended disconnection with mating device | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Biological contaminants | Device Contamination with Biological Material | Skin/Mucous membranes exposure to biological contaminants | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Septic; shock; Organ injury | 4 | 1 | Device designed to prevent the unintended disconnection with mating device | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Biological contaminants | Device Contamination with Biological Material | Skin/Mucous membranes exposure to biological contaminants | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Bacteremia; Moderate Cellulitis; Organ Inflammation (reversible) | 3 | 2 | Device designed to prevent the unintended disconnection with mating device | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Biological contaminants | Device Contamination with Biological Material | Skin/Mucous membranes exposure to biological contaminants | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | mild Cellulitis | 2 | 3 | Device designed to prevent the unintended disconnection with mating device | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Biological contaminants | Device Contamination with Biological Material | Skin/Mucous membranes exposure to biological contaminants | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Localized skin swelling; redness | 2 | 1 | Device designed to prevent the unintended disconnection with mating device | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Biological contaminants | Device Contamination with Biological Material | Skin/Mucous membranes exposure to biological contaminants | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Temporary discomfort | 1 | 1 | Device designed to prevent the unintended disconnection with mating device | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Biological contaminants | Device Contamination with Biological Material | Skin/Mucous membranes exposure to biological contaminants | Incompatible device (needle or non-compatible luer) inserted into Spiros | Localized skin swelling; redness | 2 | 1 | N/A | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Biological contaminants | Device Contamination with Biological Material | Skin/Mucous membranes exposure to biological contaminants | Patient movement causes set to disconnect at a connection point | Localized skin swelling; redness | 2 | 1 | Device designed to prevent the unintended disconnection with mating device | N/A
Biological contaminants | Device Contamination with Biological Material | Skin/Mucous membranes exposure to biological contaminants | Proper disposal procedure not followed | mild Cellulitis | 2 | 3 | N/A | Product labeling contains instructions for proper disposal per facility protocol; Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Systemic infection from virus, bacteria, or other infectious agents | 3 | 2 | Device designed to prevent the unintended disconnection with mating device | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Mild Cellulitis | 2 | 2 | Device designed to prevent the unintended disconnection with mating device | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Temporary discomfort | 1 | 2 | Device designed to prevent the unintended disconnection with mating device | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Systemic infection from virus, bacteria, or other infectious agents | 3 | 2 | Device designed not to leak during normal use; Device designed not to come apart during usage; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to reseat following 10 activations; Device designed to tolerate 10 activations to simulate a 24 hour period use life; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution statement for proper accessing Spiros by pushing straight on without angled or sliding entry and do not twist or bend connector when accessing; Labeling includes a precaution statement to inspect for leaks before use
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Mild Cellulitis | 2 | 2 | Device designed not to leak during normal use; Device designed not to come apart during usage; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to reseat following 10 activations; Device designed to tolerate 10 activations to simulate a 24 hour period use life; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution statement for proper accessing Spiros by pushing straight on without angled or sliding entry and do not twist or bend connector when accessing; Labeling includes a precaution statement to inspect for leaks before use
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Fluid leak due to defective product (E.g. Crack, broken, holes, cuts, incorrect solvent bonding) | Temporary discomfort | 1 | 2 | Device designed not to leak during normal use; Device designed not to come apart during usage; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to reseat following 10 activations; Device designed to tolerate 10 activations to simulate a 24 hour period use life; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution statement for proper accessing Spiros by pushing straight on without angled or sliding entry and do not twist or bend connector when accessing; Labeling includes a precaution statement to inspect for leaks before use
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Fluid leak due to inappropriate disposal of the set | Systemic infection from virus, bacteria, or other infectious agents | 3 | 2 | N/A | Product labeling contains instructions for proper disposal per facility protocol; Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Incompatible device (needle or non-compatible luer) inserted into Spiros | Mild Cellulitis | 2 | 2 | N/A | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Patient movement causes set to disconnect at a connection point | Temporary discomfort | 1 | 2 | Device designed to prevent the unintended disconnection with mating device | N/A
Biologic - Bacteria, virus, other agent | Device Contamination with Biological Material | Exposure to blood products caused by fluid leaks | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Infections from viral, bacteria, protozoan, and other infectious agents | 3 | 2 | Device designed to prevent the unintended disconnection with mating device | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Biologic - Bacteria, virus, other agent | Device Contamination with Biological Material | Exposure to blood products caused by fluid leaks | Fluid leak due to defective product (E.g. Crack, broken, holes, cuts, incorrect solvent bonding) | Infections from viral, bacteria, protozoan, and other infectious agents | 3 | 2 | Device designed not to leak during normal use; Device designed not to come apart during usage; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to reseat following 10 activations; Device designed to tolerate 10 activations to simulate a 24 hour period use life; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution statement for proper accessing Spiros by pushing straight on without angled or sliding entry and do not twist or bend connector when accessing; Labeling includes a precaution statement to inspect for leaks before use
Biologic - Bacteria, virus, other agent | Device Contamination with Biological Material | Exposure to blood products caused by fluid leaks | Incompatible device (needle or non-compatible luer) inserted into Spiros | Infections from viral, bacteria, protozoan, and other infectious agents | 3 | 2 | N/A | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Biologic - Bacteria, virus, other agent | Device Contamination with Biological Material | Exposure to blood products caused by fluid leaks | Patient movement causes set to disconnect at a connection point | Infections from viral, bacteria, protozoan, and other infectious agents | 3 | 2 | Device designed to prevent the unintended disconnection with mating device | N/A
Ergonomic stress | Inadequate User Interface | Exposure to ergonomic stress | Difficult / Unable to connect / disconnect the device with other mating devices (e.g. Container, access site) | Temporary discomfort | 1 | 1 | Device designed to prevent the unintended disconnection with mating device | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Ergonomic stress | Inadequate User Interface | Exposure to ergonomic stress | Difficult / Unable to open the packaging | Temporary discomfort | 1 | 1 | Sterile barrier system shall provide product in an aseptic manner | N/A
Ergonomic stress | Inadequate User Interface | Exposure to ergonomic stress | Difficult / Unable to remove Cap / cover / hood from the product | Temporary discomfort | 1 | 1 | Protective caps designed to be able to be securely attached, but easily removed | N/A
Fluid leakage | Fluid Leak | Slips/Falls | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Death; Cardiopulmonary Arrest | 5 | 1 | Device designed to prevent the unintended disconnection with mating device | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Fluid leakage | Fluid Leak | Slips/Falls | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Head Trauma with CNS; injury | 4 | 1 | Device designed to prevent the unintended disconnection with mating device | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Fluid leakage | Fluid Leak | Slips/Falls | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Fracture; Sprain; Strain; Laceration; Moderate Pain | 3 | 2 | Device designed to prevent the unintended disconnection with mating device | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Fluid leakage | Fluid Leak | Slips/Falls | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Bruise; Scrape; Skin Abrasion | 2 | 3 | Device designed to prevent the unintended disconnection with mating device | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Fluid leakage | Fluid Leak | Slips/Falls | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Temporary discomfort | 1 | 2 | Device designed to prevent the unintended disconnection with mating device | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Fluid leakage | Fluid Leak | Slips/Falls | Difficult / Unable to connect / disconnect the device with other mating devices (e.g. Container, access site) | Temporary discomfort | 1 | 2 | Device designed to prevent the unintended disconnection with mating device | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Fluid leakage | Fluid Leak | Slips/Falls | Fluid leak due to defective product (E.g. Crack, broken, holes, cuts, incorrect solvent bonding) | Bruise; Scrape; Skin Abrasion | 2 | 3 | Device designed not to leak during normal use; Device designed not to come apart during usage; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to reseat following 10 activations; Device designed to tolerate 10 activations to simulate a 24 hour period use life; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol; Device designed to function until expiry date | Labeling includes a precaution statement for proper accessing Spiros by pushing straight on without angled or sliding entry and do not twist or bend connector when accessing; Labeling includes a precaution statement to inspect for leaks before use
Fluid leakage | Fluid Leak | Slips/Falls | Fluid leak due to improper disconnection from the infuser / pump | Temporary discomfort | 1 | 2 | Device designed to prevent the unintended disconnection with mating device | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Fluid leakage | Fluid Leak | Slips/Falls | Fluid leak due to inappropriate disposal of the set | Bruise; Scrape; Skin Abrasion | 2 | 3 | N/A | Product labeling contains instructions for proper disposal per facility protocol; Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Fluid leakage | Fluid Leak | Slips/Falls | Fluid leak due to incompatible cap (priming cap) used for drug transport | Temporary discomfort | 1 | 2 | N/A | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Fluid leakage | Fluid Leak | Slips/Falls | Incompatible device (needle or non-compatible luer) inserted into Spiros | Bruise; Scrape; Skin Abrasion | 2 | 3 | N/A | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Fluid leakage | Fluid Leak | Slips/Falls | Incorrect/wrong product is used | Temporary discomfort | 1 | 2 | Labeling designed to meet customer and regulatory requirements | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA); Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Fluid leakage | Fluid Leak | Slips/Falls | Patient movement causes set to disconnect at a connection point | Bruise; Scrape; Skin Abrasion | 2 | 3 | Device designed to prevent the unintended disconnection with mating device | N/A
Fluid leakage | Fluid Leak | Slips/Falls | Reuse of the same device on multiple patients | Temporary discomfort | 1 | 2 | Labeling designed to meet customer and regulatory requirements; Packaging designed to not be reusable/resealable | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA); Labeling provides information that the device is for single patient use only; Labeling provide cautions not reuse product which are labeled for single use only
Fluid leakage | Fluid Leak | Slips/Falls | Product is used for a longer period than specified | Temporary discomfort | 1 | 2 | N/A | Labeling includes a precaution for the user to change the device at recommended intervals per facility protocol
Fluid leakage | Fluid Leak | Slips/Falls | Use of expired product | Temporary discomfort | 1 | 2 | Labeling designed to meet customer and regulatory requirements | Labeling includes caution to use by expiration date on package label; Labeling includes caution to not use if package is damaged or opened prior to use
Fluid leakage | Fluid Leak | Slips/Falls | Catheter pull out when exposed to MRI | Bruise; Scrape; Skin Abrasion | 2 | 3 | Material should be non-ferrous | N/A
Fluid leakage | Fluid Leak | Slips/Falls | Incompatibility of set materials and solutions | Bruise; Scrape; Skin Abrasion | 2 | 3 | Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | N/A
Fluid leakage | Fluid Leak | Slips/Falls | Set materials are incompatible to prolonged exposure to disinfectant caps | Temporary discomfort | 1 | 2 | Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution for the user to change the device at recommended intervals per facility protocol
Sharp edge | Material Puncture | Exposure to sharp edge/pinch point/abrasive surface | Product contains sharp corners/edges | Skin Abrasion; Mild Bruising Pain; Skin integrity compromised | 2 | 2 | Device designed and processed in a manner to be free from cracks, splits, fractures, and discoloration; Production controls designed to limit the amount of flash on products | N/A
Sharp edge | Material Puncture | Exposure to sharp edge/pinch point/abrasive surface | Product contains sharp corners/edges | Temporary discomfort | 1 | 2 | Device designed and processed in a manner to be free from cracks, splits, fractures, and discoloration; Production controls designed to limit the amount of flash on products | N/A
Device Tension | Retraction Problem | Patient access device is pulled out | Catheter pull out when exposed to MRI | Hematoma; Moderate Pain; Moderate Cellulitis | 3 | 2 | Material should be non-ferrous | N/A
Device Tension | Retraction Problem | Patient access device is pulled out | Catheter pull out when exposed to MRI | Bruise; Skin Abrasion; Mild Pain | 3 | 2 | Material should be non-ferrous | N/A
Device Tension | Retraction Problem | Patient access device is pulled out | Catheter pull out when exposed to MRI | Temporary discomfort | 1 | 2 | Material should be non-ferrous | N/A
Device Tension | Retraction Problem | Patient access device is pulled out | Patient movement causes set to disconnect at a connection point | Hematoma; Moderate Pain; Moderate Cellulitis | 3 | 2 | Device designed to prevent the unintended disconnection with mating device | N/A
Device Tension | Retraction Problem | Patient access device is pulled out | Patient movement causes set to disconnect at a connection point | Bruise; Skin Abrasion; Mild Pain | 3 | 2 | Device designed to prevent the unintended disconnection with mating device | N/A
Device Tension | Retraction Problem | Patient access device is pulled out | Patient movement causes set to disconnect at a connection point | Temporary discomfort | 1 | 2 | Device designed to prevent the unintended disconnection with mating device | N/A
Removable component | Detachment of Device Component | User ingests removable component | Removable components left in patient bed or placed into mouth | Death; Cardiopulmonary Arrest | 5 | 2 | N/A | Product labeling contains instructions for proper disposal per facility protocol
Removable component | Detachment of Device Component | User ingests removable component | Removable components left in patient bed or placed into mouth | Hypoxic organ injury; Severe Respiratory Distress | 4 | 1 | N/A | Product labeling contains instructions for proper disposal per facility protocol
Removable component | Detachment of Device Component | User ingests removable component | Removable components left in patient bed or placed into mouth | Esophageal or respiratory airway obstruction; Gastric foreign body | 3 | 2 | N/A | Product labeling contains instructions for proper disposal per facility protocol
Removable component | Detachment of Device Component | User ingests removable component | Removable components left in patient bed or placed into mouth | Coughing or gagging | 2 | 2 | N/A | Product labeling contains instructions for proper disposal per facility protocol
Removable component | Detachment of Device Component | User ingests removable component | Removable components left in patient bed or placed into mouth | Temporary discomfort | 1 | 1 | N/A | Product labeling contains instructions for proper disposal per facility protocol
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/ fumes | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Mild to Moderate Respiratory Distress; Moderate toxic effects such as seizures, abdominal pain | 3 | 2 | Device designed to prevent the unintended disconnection with mating device | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/ fumes | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Mild Pain; Mild toxic effects such as fatigue, sleepiness, headaches, eye irritation, skin rashes | 2 | 2 | Device designed to prevent the unintended disconnection with mating device | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/ fumes | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Temporary discomfort | 1 | 2 | Device designed to prevent the unintended disconnection with mating device | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/ fumes | Fluid leak due to defective product (E.g. Crack, broken, holes, cuts, incorrect solvent bonding) | Mild to Moderate Respiratory Distress; Moderate toxic effects such as seizures, abdominal pain | 3 | 2 | Device designed not to leak during normal use; Device designed not to come apart during usage; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to reseat following 10 activations; Device designed to tolerate 10 activations to simulate a 24 hour period use life; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution statement for proper accessing Spiros by pushing straight on without angled or sliding entry and do not twist or bend connector when accessing; Labeling includes a precaution statement to inspect for leaks before use
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/ fumes | Fluid leak due to defective product (E.g. Crack, broken, holes, cuts, incorrect solvent bonding) | Mild Pain; Mild toxic effects such as fatigue, sleepiness, headaches, eye irritation, skin rashes | 2 | 2 | Device designed not to leak during normal use; Device designed not to come apart during usage; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to reseat following 10 activations; Device designed to tolerate 10 activations to simulate a 24 hour period use life; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution statement for proper accessing Spiros by pushing straight on without angled or sliding entry and do not twist or bend connector when accessing; Labeling includes a precaution statement to inspect for leaks before use
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/ fumes | Fluid leak due to defective product (E.g. Crack, broken, holes, cuts, incorrect solvent bonding) | Temporary discomfort | 1 | 2 | Device designed not to leak during normal use; Device designed not to come apart during usage; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to reseat following 10 activations; Device designed to tolerate 10 activations to simulate a 24 hour period use life; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution statement for proper accessing Spiros by pushing straight on without angled or sliding entry and do not twist or bend connector when accessing; Labeling includes a precaution statement to inspect for leaks before use
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/ fumes | Fluid leak due to inappropriate disposal of the set | Mild Pain; Mild toxic effects such as fatigue, sleepiness, headaches, eye irritation, skin rashes | 2 | 2 | N/A | Product labeling contains instructions for proper disposal per facility protocol; Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/ fumes | Fluid leak due to inappropriate disposal of the set | Temporary discomfort | 1 | 2 | N/A | Product labeling contains instructions for proper disposal per facility protocol; Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/ fumes | Incompatible device (needle or non-compatible luer) inserted into Spiros | Mild Pain; Mild toxic effects such as fatigue, sleepiness, headaches, eye irritation, skin rashes | 2 | 2 | N/A | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/ fumes | Incompatible device (needle or non-compatible luer) inserted into Spiros | Temporary discomfort | 1 | 2 | N/A | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/ fumes | Exposure to solution from device after disconnection | Mild Pain; Mild toxic effects such as fatigue, sleepiness, headaches, eye irritation, skin rashes | 2 | 2 | Device designed to minimize the residual moisture during disconnection | N/A
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/ fumes | Exposure to solution from device after disconnection | Temporary discomfort | 1 | 2 | Device designed to minimize the residual moisture during disconnection | N/A
Blood Reflux / Reverse flow | Reflux within Device | Patient blood loss due to backflow of blood into set/device | Fluid leak due to defective product (E.g. Crack, broken, holes, cuts, incorrect solvent bonding) | Death | 5 | 1 | Device designed not to leak during normal use; Device designed not to come apart during usage; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to reseat following 10 activations; Device designed to tolerate 10 activations to simulate a 24 hour period use life; Device designed to function until expiry date; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution statement for proper accessing Spiros by pushing straight on without angled or sliding entry and do not twist or bend connector when accessing; Labeling includes a precaution statement to inspect for leaks before use; Labeling includes caution to use by expiration date on package label; Labeling includes caution to not use if package is damaged or opened prior to use
Blood Reflux / Reverse flow | Reflux within Device | Patient blood loss due to backflow of blood into set/device | Fluid leak due to defective product (E.g. Crack, broken, holes, cuts, incorrect solvent bonding) | Hypovolemic shock | 4 | 2 | Device designed not to leak during normal use; Device designed not to come apart during usage; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to reseat following 10 activations; Device designed to tolerate 10 activations to simulate a 24 hour period use life; Device designed to function until expiry date; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution statement for proper accessing Spiros by pushing straight on without angled or sliding entry and do not twist or bend connector when accessing; Labeling includes a precaution statement to inspect for leaks before use; Labeling includes caution to use by expiration date on package label; Labeling includes caution to not use if package is damaged or opened prior to use
Blood Reflux / Reverse flow | Reflux within Device | Patient blood loss due to backflow of blood into set/device | Fluid leak due to defective product (E.g. Crack, broken, holes, cuts, incorrect solvent bonding) | Mild to Moderate Hypotension; Mild to Moderate Hypovolemia; Significant bleeding requiring transfusion | 3 | 2 | Device designed not to leak during normal use; Device designed not to come apart during usage; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to reseat following 10 activations; Device designed to tolerate 10 activations to simulate a 24 hour period use life; Device designed to function until expiry date; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution statement for proper accessing Spiros by pushing straight on without angled or sliding entry and do not twist or bend connector when accessing; Labeling includes a precaution statement to inspect for leaks before use; Labeling includes caution to use by expiration date on package label; Labeling includes caution to not use if package is damaged or opened prior to use
Blood Reflux / Reverse flow | Reflux within Device | Patient blood loss due to backflow of blood into set/device | Fluid leak due to defective product (E.g. Crack, broken, holes, cuts, incorrect solvent bonding) | Temporary Discomfort | 1 | 3 | Device designed not to leak during normal use; Device designed not to come apart during usage; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to reseat following 10 activations; Device designed to tolerate 10 activations to simulate a 24 hour period use life; Device designed to function until expiry date; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution statement for proper accessing Spiros by pushing straight on without angled or sliding entry and do not twist or bend connector when accessing; Labeling includes a precaution statement to inspect for leaks before use; Labeling includes caution to use by expiration date on package label; Labeling includes caution to not use if package is damaged or opened prior to use
Blood Reflux / Reverse flow | Reflux within Device | Patient blood loss due to backflow of blood into set/device | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Hypovolemic shock | 4 | 2 | Device designed to prevent the unintended disconnection with mating device | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Blood Reflux / Reverse flow | Reflux within Device | Patient blood loss due to backflow of blood into set/device | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Mild to Moderate Hypotension; Mild to Moderate Hypovolemia; Significant bleeding requiring transfusion | 3 | 2 | Device designed to prevent the unintended disconnection with mating device | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Blood Reflux / Reverse flow | Reflux within Device | Patient blood loss due to backflow of blood into set/device | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Temporary Discomfort | 1 | 3 | Device designed to prevent the unintended disconnection with mating device | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Drug with altered activity | Medication Error | Infusion of drug with altered activity | Incompatible drug solutions infused concurrently | Death; Cardiopulmonary Arrest | 5 | 1 | Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | N/A
Drug with altered activity | Medication Error | Infusion of drug with altered activity | Incompatible drug solutions infused concurrently | Pulmonary embolism; Septic shock; Severe Hyperglycemia; Severe Hypertension; Severe Hypoglycemia; Severe Hypotension; Severe Arrhythmias; Severe Pain | 4 | 1 | Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | N/A
Drug with altered activity | Medication Error | Infusion of drug with altered activity | Incompatible drug solutions infused concurrently | Mild to Moderate Arrhythmia; Moderate Hyperglycemia; Mild to Moderate Hypertension; Moderate Hypoglycemia; Mild to Moderate Hypotension; Moderate Pain | 3 | 2 | Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | N/A
Drug with altered activity | Medication Error | Infusion of drug with altered activity | Incompatible drug solutions infused concurrently | Mild Pain; Mild Hyperglycemia; Mild Hypoglycemia | 2 | 1 | Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | N/A
Drug with altered activity | Medication Error | Infusion of drug with altered activity | Incompatible drug solutions infused concurrently | Temporary discomfort | 1 | 1 | Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | N/A
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Death; Cardiopulmonary Arrest | 5 | 1 | Device designed to prevent the unintended disconnection with mating device | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Pulmonary embolism; Septic shock; Severe Hyperglycemia; Severe Hypoglycemia; Severe Hypertension; Severe Hypotension; Severe Arrhythmias; Severe Pain | 4 | 2 | Device designed to prevent the unintended disconnection with mating device | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Mild to Moderate Hypotension; Moderate Hyperglycemia; Moderate Hypoglycemia; Moderate Pain | 3 | 3 | Device designed to prevent the unintended disconnection with mating device | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Mild Pain; Mild Hyperglycemia; Mild Hypoglycemia | 2 | 3 | Device designed to prevent the unintended disconnection with mating device | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Temporary discomfort | 1 | 2 | Device designed to prevent the unintended disconnection with mating device | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Fluid leak due to defective product (E.g. Crack, broken, holes, cuts, incorrect solvent bonding) | Death; Cardiopulmonary Arrest | 5 | 1 | Device designed not to leak during normal use; Device designed not to come apart during usage; Device designed to prevent the leak when subject to 36'' head height back pressure; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution statement for proper accessing Spiros by pushing straight on without angled or sliding entry and do not twist or bend connector when accessing; Labeling includes a precaution statement to inspect for leaks before use
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Fluid leak due to defective product (E.g. Crack, broken, holes, cuts, incorrect solvent bonding) | Pulmonary embolism; Septic shock; Severe Hyperglycemia, Severe Hypoglycemia; Severe Hypertension; Severe Hypotension; Severe Arrhythmias; Severe Pain | 4 | 2 | Device designed not to leak during normal use; Device designed not to come apart during usage; Device designed to prevent the leak when subject to 36'' head height back pressure; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution statement for proper accessing Spiros by pushing straight on without angled or sliding entry and do not twist or bend connector when accessing; Labeling includes a precaution statement to inspect for leaks before use
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Fluid leak due to defective product (E.g. Crack, broken, holes, cuts, incorrect solvent bonding) | Mild to Moderate Hypotension; Moderate Hyperglycemia; Moderate Hypoglycemia; Moderate Pain | 3 | 3 | Device designed not to leak during normal use; Device designed not to come apart during usage; Device designed to prevent the leak when subject to 36'' head height back pressure; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution statement for proper accessing Spiros by pushing straight on without angled or sliding entry and do not twist or bend connector when accessing; Labeling includes a precaution statement to inspect for leaks before use
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Fluid leak due to defective product (E.g. Crack, broken, holes, cuts, incorrect solvent bonding) | Mild Pain; Mild Hyperglycemia; Mild Hypoglycemia | 2 | 3 | Device designed not to leak during normal use; Device designed not to come apart during usage; Device designed to prevent the leak when subject to 36'' head height back pressure; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution statement for proper accessing Spiros by pushing straight on without angled or sliding entry and do not twist or bend connector when accessing; Labeling includes a precaution statement to inspect for leaks before use
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Fluid leak due to defective product (E.g. Crack, broken, holes, cuts, incorrect solvent bonding) | Temporary discomfort | 1 | 2 | Device designed not to leak during normal use; Device designed not to come apart during usage; Device designed to prevent the leak when subject to 36'' head height back pressure; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution statement for proper accessing Spiros by pushing straight on without angled or sliding entry and do not twist or bend connector when accessing; Labeling includes a precaution statement to inspect for leaks before use
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Incompatibility of set materials and solutions | Mild to Moderate Hypotension; Moderate Hyperglycemia; Moderate Hypoglycemia; Moderate Pain | 3 | 3 | Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | N/A
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Incompatibility of set materials and solutions | Mild Pain; Mild Hyperglycemia; Mild Hypoglycemia | 2 | 3 | Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | N/A
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Incompatible device (needle or non-compatible luer) inserted into Spiros | Mild Pain; Mild Hyperglycemia; Mild Hypoglycemia | 2 | 3 | N/A | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Incompatible device (needle or non-compatible luer) inserted into Spiros | Temporary discomfort | 1 | 2 | N/A | Labeling contains precautions not to use needles for luer activated injection access sites; Labeling contains precautions not to use luer style caps on Spiros as this will activate the connector; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Catheter pull out when exposed to MRI | Mild Pain; Mild Hyperglycemia; Mild Hypoglycemia | 2 | 3 | Material should be non-ferrous | N/A
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Catheter pull out when exposed to MRI | Temporary discomfort | 1 | 2 | Material should be non-ferrous | N/A
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Incorrect/wrong product is used | Mild Pain; Mild Hyperglycemia; Mild Hypoglycemia | 2 | 3 | Labeling designed to meet customer and regulatory requirements | N/A
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Incorrect/wrong product is used | Temporary discomfort | 1 | 2 | Labeling designed to meet customer and regulatory requirements | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | No fluid flow | Mild Pain; Mild Hyperglycemia; Mild Hypoglycemia | 2 | 3 | Device designed to maintain required flow rate; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | N/A
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | No fluid flow | Temporary discomfort | 1 | 2 | Device designed to maintain required flow rate; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | N/A
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Problem delivering blood products at the expected rate | Death; Cardiopulmonary Arrest | 5 | 1 | Device designed not to leak during normal use; Device designed not to come apart during usage; Device designed to function until expiry date; Device designed to maintain required flow rate; Device designed to prevent the leak when subject to 36'' head height back pressure; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution statement for proper accessing Spiros by pushing straight on without angled or sliding entry and do not twist or bend connector when accessing; Labeling includes a precaution statement to inspect for leaks before use; Label; Labeling includes caution to use by expiration date on package label; Labeling includes caution to not use if package is damaged or opened prior to use
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Problem delivering blood products at the expected rate | Pulmonary embolism; Septic shock; Severe Hyperglycemia; Severe Hypoglycemia; Severe Hypertension; Severe Hypotension; Severe Arrhythmias; Severe Pain | 4 | 2 | Device designed not to leak during normal use; Device designed not to come apart during usage; Device designed to function until expiry date; Device designed to maintain required flow rate; Device designed to prevent the leak when subject to 36'' head height back pressure; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution statement for proper accessing Spiros by pushing straight on without angled or sliding entry and do not twist or bend connector when accessing; Labeling includes a precaution statement to inspect for leaks before use; Label; Labeling includes caution to use by expiration date on package label; Labeling includes caution to not use if package is damaged or opened prior to use
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Problem delivering blood products at the expected rate | Mild to Moderate Hypotension; Moderate Hyperglycemia; Moderate Hypoglycemia; Moderate Pain | 3 | 3 | Device designed not to leak during normal use; Device designed not to come apart during usage; Device designed to function until expiry date; Device designed to maintain required flow rate; Device designed to prevent the leak when subject to 36'' head height back pressure; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution statement for proper accessing Spiros by pushing straight on without angled or sliding entry and do not twist or bend connector when accessing; Labeling includes a precaution statement to inspect for leaks before use; Label; Labeling includes caution to use by expiration date on package label; Labeling includes caution to not use if package is damaged or opened prior to use
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Problem delivering blood products at the expected rate | Mild Pain; Mild Hyperglycemia; Mild Hypoglycemia | 2 | 3 | Device designed not to leak during normal use; Device designed not to come apart during usage; Device designed to function until expiry date; Device designed to maintain required flow rate; Device designed to prevent the leak when subject to 36'' head height back pressure; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution statement for proper accessing Spiros by pushing straight on without angled or sliding entry and do not twist or bend connector when accessing; Labeling includes a precaution statement to inspect for leaks before use; Label; Labeling includes caution to use by expiration date on package label; Labeling includes caution to not use if package is damaged or opened prior to use
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Problem delivering blood products at the expected rate | Temporary discomfort | 1 | 2 | Device designed not to leak during normal use; Device designed not to come apart during usage; Device designed to function until expiry date; Device designed to maintain required flow rate; Device designed to prevent the leak when subject to 36'' head height back pressure; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution statement for proper accessing Spiros by pushing straight on without angled or sliding entry and do not twist or bend connector when accessing; Labeling includes a precaution statement to inspect for leaks before use; Label; Labeling includes caution to use by expiration date on package label; Labeling includes caution to not use if package is damaged or opened prior to use
IV Catheter out of vein | Retraction Problem | Infusing solution outside of the vessel | Catheter pull out when exposed to MRI | Severe Burn; Edema; Local tissue necrosis; Severe Pain | 4 | 1 | Material should be non-ferrous | N/A
IV Catheter out of vein | Retraction Problem | Infusing solution outside of the vessel | Catheter pull out when exposed to MRI | Hematoma; Moderate Burn; Edema; Moderate Pain | 3 | 2 | Material should be non-ferrous | N/A
IV Catheter out of vein | Retraction Problem | Infusing solution outside of the vessel | Catheter pull out when exposed to MRI | Burning sensation; Mild Burn; Edema; Mild Pain | 2 | 2 | Material should be non-ferrous | N/A
IV Catheter out of vein | Retraction Problem | Infusing solution outside of the vessel | Catheter pull out when exposed to MRI | Temporary discomfort | 1 | 2 | Material should be non-ferrous | N/A
Vein clotted off | Use of Device Issue | Exposure to physical injury at infusion site | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Hematoma; Moderate Pain; Skin irritation; bruising; Mild to Moderate Cellulitis | 3 | 3 | Device designed to prevent the unintended disconnection with mating device | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Vein clotted off | Use of Device Issue | Exposure to physical injury at infusion site | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Hematoma Pain, Mild | 2 | 3 | Device designed to prevent the unintended disconnection with mating device | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Vein clotted off | Use of Device Issue | Exposure to physical injury at infusion site | Bad connection / disconnection between container / set / other mating devices (E.g. Loose connection, improper mating) | Temporary discomfort | 1 | 2 | Device designed to prevent the unintended disconnection with mating device | Labeling provides information about the precautions (E.g.., Usage of set/device, packaging conditions); Labeling includes cautions that the device is intended for use by trained clinical personnel only or per RX only (per FDA guidance for USA)
Vein clotted off | Use of Device Issue | Exposure to physical injury at infusion site | Problem delivering blood products at the expected rate | Hematoma; Moderate Pain; Skin irritation; bruising; Mild to Moderate Cellulitis | 3 | 3 | Device designed not to leak during normal use; Device designed not to come apart during usage; Device designed to function until expiry date; Device designed to maintain required flow rate; Device designed to prevent the leak when subject to 36'' head height back pressure; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution statement for proper accessing Spiros by pushing straight on without angled or sliding entry and do not twist or bend connector when accessing; Labeling includes a precaution statement to inspect for leaks before use; Labeling includes caution to use by expiration date on package label;Labeling includes caution to not use if package is damaged or opened prior to use
Vein clotted off | Use of Device Issue | Exposure to physical injury at infusion site | Problem delivering blood products at the expected rate | Hematoma; Mild Pain | 2 | 3 | Device designed not to leak during normal use; Device designed not to come apart during usage; Device designed to function until expiry date; Device designed to maintain required flow rate; Device designed to prevent the leak when subject to 36'' head height back pressure; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution statement for proper accessing Spiros by pushing straight on without angled or sliding entry and do not twist or bend connector when accessing; Labeling includes a precaution statement to inspect for leaks before use; Labeling includes caution to use by expiration date on package label;Labeling includes caution to not use if package is damaged or opened prior to use
Vein clotted off | Use of Device Issue | Exposure to physical injury at infusion site | Problem delivering blood products at the expected rate | Temporary discomfort | 1 | 2 | Device designed not to leak during normal use; Device designed not to come apart during usage; Device designed to function until expiry date; Device designed to maintain required flow rate; Device designed to prevent the leak when subject to 36'' head height back pressure; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | Labeling includes a precaution statement for proper accessing Spiros by pushing straight on without angled or sliding entry and do not twist or bend connector when accessing; Labeling includes a precaution statement to inspect for leaks before use; Labeling includes caution to use by expiration date on package label;Labeling includes caution to not use if package is damaged or opened prior to use
Undetected occlusion | Insufficient Flow or Underinfusion | Under-delivery to the patient | Blood clot introduced into the device | Death; Cardiopulmonary Arrest | 5 | 1 | Device designed to maintain required flow rate | N/A
Undetected occlusion | Insufficient Flow or Underinfusion | Under-delivery to the patient | Blood clot introduced into the device | Pulmonary embolism; Septic shock; Severe Hyperglycemia; Severe Hypertension; Severe Hypoglycemia; Severe Hypotension; Severe Arrhythmias; Severe Pain | 4 | 2 | Device designed to maintain required flow rate | N/A
Undetected occlusion | Insufficient Flow or Underinfusion | Under-delivery to the patient | Blood clot introduced into the device | Mild to Moderate Arrhythmia; Moderate Hyperglycemia; Mild to Moderate Hypertension; Moderate Hypoglycemia; Mild to Moderate Hypotension; Moderate Pain | 3 | 2 | Device designed to maintain required flow rate | N/A
Undetected occlusion | Insufficient Flow or Underinfusion | Under-delivery to the patient | Blood clot introduced into the device | Mild Pain; Mild Hyperglycemia; Mild Hypoglycemia | 2 | 2 | Device designed to maintain required flow rate | N/A
Undetected occlusion | Insufficient Flow or Underinfusion | Under-delivery to the patient | Blood clot introduced into the device | Temporary Discomfort | 1 | 2 | Device designed to maintain required flow rate | N/A
Undetected occlusion | Insufficient Flow or Underinfusion | Under-delivery to the patient | Fluid path occlusion | Death; Cardiopulmonary Arrest | 5 | 1 | Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | N/A
Undetected occlusion | Insufficient Flow or Underinfusion | Under-delivery to the patient | Fluid path occlusion | Pulmonary embolism; Septic shock; Severe Hyperglycemia; Severe Hypertension; Severe Hypoglycemia; Severe Hypotension; Severe Arrhythmias; Severe Pain | 4 | 2 | Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | N/A
Undetected occlusion | Insufficient Flow or Underinfusion | Under-delivery to the patient | Fluid path occlusion | Mild to Moderate Arrhythmia; Moderate Hyperglycemia; Mild to Moderate Hypertension; Moderate Hypoglycemia; Mild to Moderate Hypotension; Moderate Pain | 3 | 2 | Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | N/A
Undetected occlusion | Insufficient Flow or Underinfusion | Under-delivery to the patient | Fluid path occlusion | Mild Pain; Mild Hyperglycemia; Mild Hypoglycemia | 2 | 2 | Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | N/A
Undetected occlusion | Insufficient Flow or Underinfusion | Under-delivery to the patient | No fluid flow | Mild to Moderate Arrhythmia; Moderate Hyperglycemia; Mild to Moderate Hypertension; Moderate Hypoglycemia; Mild to Moderate Hypotension; Moderate Pain | 3 | 2 | Device designed to maintain required flow rate; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | N/A
Undetected occlusion | Insufficient Flow or Underinfusion | Under-delivery to the patient | No fluid flow | Mild Pain; Mild Hyperglycemia; Mild Hypoglycemia | 2 | 2 | Device designed to maintain required flow rate; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | N/A
Undetected occlusion | Insufficient Flow or Underinfusion | Under-delivery to the patient | No fluid flow | Temporary Discomfort | 1 | 2 | Device designed to maintain required flow rate; Device materials selected that are resistant to cracking and leakage under clinical stress and in the presence of stress relieving fluids and alcohol | N/A
Inappropriate disposal | Environmental Control or Utility Issue | Environmental (Landfill/Water Supply) exposure to toxic substances from device material | Inappropriate material selection | Moderate toxic effects such as seizures, abdominal pain | 3 | 1 | Materials of construction should comply with regulatory requirements | N/A
Inappropriate disposal | Environmental Control or Utility Issue | Environmental (Landfill/Water Supply) exposure to toxic substances from device material | Inappropriate material selection | Mild toxic effects such as fatigue, sleepiness, headaches, eye irritation, skin rashes | 2 | 1 | Materials of construction should comply with regulatory requirements | N/A
Inappropriate disposal | Environmental Control or Utility Issue | Environmental (Landfill/Water Supply) exposure to toxic substances from device material | Inappropriate material selection | Temporary discomfort | 1 | 1 | Materials of construction should comply with regulatory requirements | N/A
Air in line | Improper Flow or Infusion | Air infused into the body | Air in line due to improper priming | Death | 5 | 1 | Device will allow for multiple activations to comply with priming per facility protocols | Labeling provides instructions to prime device before use
Air in line | Improper Flow or Infusion | Air infused into the body | Air in line due to improper priming | Pulmonary Embolus; Stroke | 4 | 1 | Device will allow for multiple activations to comply with priming per facility protocols | Labeling provides instructions to prime device before use
Air in line | Improper Flow or Infusion | Air infused into the body | Air in line due to improper priming | Mild to Moderate Respiratory Distress | 3 | 1 | Device will allow for multiple activations to comply with priming per facility protocols | Labeling provides instructions to prime device before use
Air in line | Improper Flow or Infusion | Air infused into the body | Air in line due to improper priming | Temporary discomfort | 1 | 1 | Device will allow for multiple activations to comply with priming per facility protocols | Labeling provides instructions to prime device before use
Air in line | Improper Flow or Infusion | Air infused into the body | Air in line due to broken, loose or unsealed fluid path | Death | 5 | 1 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards); Device is designed with a self-sealing septum | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning; slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to not use needles, blunt cannulas, or luer-type caps on Tego
Air in line | Improper Flow or Infusion | Air infused into the body | Air in line due to broken, loose or unsealed fluid path | Pulmonary Embolus; Stroke | 4 | 1 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards); Device is designed with a self-sealing septum | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning; slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to not use needles, blunt cannulas, or luer-type caps on Tego
Air in line | Improper Flow or Infusion | Air infused into the body | Air in line due to broken, loose or unsealed fluid path | Mild to Moderate Respiratory Distress | 3 | 1 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards); Device is designed with a self-sealing septum | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning; slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to not use needles, blunt cannulas, or luer-type caps on Tego
Air in line | Improper Flow or Infusion | Air infused into the body | Air in line due to broken, loose or unsealed fluid path | Temporary discomfort | 1 | 1 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards); Device is designed with a self-sealing septum | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning; slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to not use needles, blunt cannulas, or luer-type caps on Tego
Allergenic Substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Allergenic substances on packaging/device material | Contact dermatitis with localized allergic reaction | 2 | 2 | Device is designed to be NON-DEHP and not contain latex | Labeling provides information that the device does not contain materials with DEHP or natural rubber latex in the manufacturing of the device
Allergenic Substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Allergenic substances on packaging/device material | Temporary discomfort | 1 | 2 | Device is designed to be NON-DEHP and not contain latex | Labeling provides information that the device does not contain materials with DEHP or natural rubber latex in the manufacturing of the device
Allergenic Substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Fluid leak due to defective device (Eg Crack, holes, cuts) | Contact dermatitis with localized allergic reaction | 2 | 2 | Devices is designed not to leak when connected; Device and packaging is designed to withstand shipping and distribution; Device and packaging are designed to withstand the sterilization process | Labeling includes a caution to discard if device packaging is damaged or open; Labeling includes a general warning on the use of a damaged device or improper use; Labeling includes a general warning to inspect device and read instructions for use prior to usage; Labeling includes a precaution statement to clamp line before disconnecting from Tego, and between dialysis sessions per established clinical and catheter manufacturer’s practices; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination
Allergenic Substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Fluid leak due to defective device (Eg Crack, holes, cuts) | Temporary discomfort | 1 | 2 | Devices is designed not to leak when connected; Device and packaging is designed to withstand shipping and distribution; Device and packaging are designed to withstand the sterilization process | Labeling includes a caution to discard if device packaging is damaged or open; Labeling includes a general warning on the use of a damaged device or improper use; Labeling includes a general warning to inspect device and read instructions for use prior to usage; Labeling includes a precaution statement to clamp line before disconnecting from Tego, and between dialysis sessions per established clinical and catheter manufacturer’s practices; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination
Allergenic Substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Contact dermatitis with localized allergic reaction | 2 | 2 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Allergenic Substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Temporary discomfort | 1 | 2 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Allergenic Substances | Biocompatibility Issue | Infusion of allergenic substances | Aseptic techniques not followed | Death; Anaphylaxis | 5 | 1 | Device is packaged with protective cap to maintain sterile barrier | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling provides instructions to remove Tego from package using aseptic technique; Labeling includes instructions/precaution to disinfect Tego prior to each use; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Allergenic Substances | Biocompatibility Issue | Infusion of allergenic substances | Aseptic techniques not followed | Severe Respiratory distress, Stevens Johnson Syndrome | 4 | 1 | Device is packaged with protective cap to maintain sterile barrier | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling provides instructions to remove Tego from package using aseptic technique; Labeling includes instructions/precaution to disinfect Tego prior to each use; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Allergenic Substances | Biocompatibility Issue | Infusion of allergenic substances | Aseptic techniques not followed | Mild to Moderate Respiratory Distress | 3 | 2 | Device is packaged with protective cap to maintain sterile barrier | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling provides instructions to remove Tego from package using aseptic technique; Labeling includes instructions/precaution to disinfect Tego prior to each use; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Allergenic Substances | Biocompatibility Issue | Infusion of allergenic substances | Aseptic techniques not followed | Urticaria | 2 | 2 | Device is packaged with protective cap to maintain sterile barrier | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling provides instructions to remove Tego from package using aseptic technique; Labeling includes instructions/precaution to disinfect Tego prior to each use; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Allergenic Substances | Biocompatibility Issue | Infusion of allergenic substances | Aseptic techniques not followed | Temporary discomfort | 1 | 2 | Device is packaged with protective cap to maintain sterile barrier | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling provides instructions to remove Tego from package using aseptic technique; Labeling includes instructions/precaution to disinfect Tego prior to each use; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Allergenic Substances | Biocompatibility Issue | Infusion of allergenic substances | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Death; Anaphylaxis | 5 | 1 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Allergenic Substances | Biocompatibility Issue | Infusion of allergenic substances | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Severe Respiratory distress; Stevens Johnson Syndrome | 4 | 1 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Allergenic Substances | Biocompatibility Issue | Infusion of allergenic substances | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Mild to Moderate Respiratory Distress | 3 | 2 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Allergenic Substances | Biocompatibility Issue | Infusion of allergenic substances | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Urticaria | 2 | 2 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Allergenic Substances | Biocompatibility Issue | Infusion of allergenic substances | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Temporary discomfort | 1 | 2 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Allergenic Substances | Biocompatibility Issue | Infusion of allergenic substances | Allergenic substances in the fluid path | Death; Anaphylaxis | 5 | 1 | Device is designed with a self-sealing septum; Device is packaged with protective cap to maintain sterile barrier | Labeling includes a precaution statement to not use luer-type caps on Tego; Labeling provides information that the device does not contain materials with DEHP or natural rubber latex in the manufacturing of the device; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling provides instructions to remove Tego from package using aseptic technique; Labeling includes instructions/precaution to disinfect Tego prior to each use; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Allergenic Substances | Biocompatibility Issue | Infusion of allergenic substances | Allergenic substances in the fluid path | Severe Respiratory distress; Stevens Johnson Syndrome | 4 | 1 | Device is designed with a self-sealing septum; Device is packaged with protective cap to maintain sterile barrier | Labeling includes a precaution statement to not use luer-type caps on Tego; Labeling provides information that the device does not contain materials with DEHP or natural rubber latex in the manufacturing of the device; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling provides instructions to remove Tego from package using aseptic technique; Labeling includes instructions/precaution to disinfect Tego prior to each use; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Allergenic Substances | Biocompatibility Issue | Infusion of allergenic substances | Allergenic substances in the fluid path | Mild to Moderate Respiratory Distress | 3 | 2 | Device is designed with a self-sealing septum; Device is packaged with protective cap to maintain sterile barrier | Labeling includes a precaution statement to not use luer-type caps on Tego; Labeling provides information that the device does not contain materials with DEHP or natural rubber latex in the manufacturing of the device; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling provides instructions to remove Tego from package using aseptic technique; Labeling includes instructions/precaution to disinfect Tego prior to each use; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Allergenic Substances | Biocompatibility Issue | Infusion of allergenic substances | Allergenic substances in the fluid path | Urticaria | 2 | 2 | Device is designed with a self-sealing septum; Device is packaged with protective cap to maintain sterile barrier | Labeling includes a precaution statement to not use luer-type caps on Tego; Labeling provides information that the device does not contain materials with DEHP or natural rubber latex in the manufacturing of the device; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling provides instructions to remove Tego from package using aseptic technique; Labeling includes instructions/precaution to disinfect Tego prior to each use; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Allergenic Substances | Biocompatibility Issue | Infusion of allergenic substances | Allergenic substances in the fluid path | Temporary discomfort | 1 | 2 | Device is designed with a self-sealing septum; Device is packaged with protective cap to maintain sterile barrier | Labeling includes a precaution statement to not use luer-type caps on Tego; Labeling provides information that the device does not contain materials with DEHP or natural rubber latex in the manufacturing of the device; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling provides instructions to remove Tego from package using aseptic technique; Labeling includes instructions/precaution to disinfect Tego prior to each use; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Biologic - Bacteria, virus, other agent | Device Contamination with Biological Material | Exposure to blood products caused by fluid leaks | Fluid leak due to defective device (Eg Crack, holes, cuts) | Patient is hemodynamically stable. | 3 | 1 | Devices is designed not to leak when connected; Device and packaging is designed to withstand shipping and distribution; Device and packaging are designed to withstand the sterilization process | Labeling includes a caution to discard if device packaging is damaged or open; Labeling includes a general warning on the use of a damaged device or improper use; Labeling includes a general warning to inspect device and read instructions for use prior to usage; Labeling includes a precaution statement to clamp line before disconnecting from Tego, and between dialysis sessions per established clinical and catheter manufacturer’s practices; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Biologic - Bacteria, virus, other agent | Device Contamination with Biological Material | Exposure to blood products caused by fluid leaks | Fluid leak due to inappropriate disposal of the device | Patient is hemodynamically stable. | 3 | 1 | N/A | Labeling includes instructions to dispose device in accordance with facility protocol
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Loss of packaging integrity allows foreign material in the sterile device | Death; Cardiopulmonary Arrest | 5 | 1 | Device and packaging are designed to withstand the sterilization process | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Loss of packaging integrity allows foreign material in the sterile device | Septic shock; Organ injury | 4 | 1 | Device and packaging are designed to withstand the sterilization process | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Loss of packaging integrity allows foreign material in the sterile device | Bacteremia; Moderate Pain | 3 | 1 | Device and packaging are designed to withstand the sterilization process | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Loss of packaging integrity allows foreign material in the sterile device | Mild Phlebitis; Local reaction/ infection; Mild Pain | 2 | 1 | Device and packaging are designed to withstand the sterilization process | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Loss of packaging integrity allows foreign material in the sterile device | Temporary discomfort | 1 | 1 | Device and packaging are designed to withstand the sterilization process | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Sterility compromised due to loss of packaging integrity | Death; Cardiopulmonary Arrest | 5 | 1 | Device and packaging are designed to withstand the sterilization process | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Sterility compromised due to loss of packaging integrity | Septic shock; Organ injury | 4 | 1 | Device and packaging are designed to withstand the sterilization process | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Sterility compromised due to loss of packaging integrity | Bacteremia; Moderate Pain | 3 | 1 | Device and packaging are designed to withstand the sterilization process | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Sterility compromised due to loss of packaging integrity | Mild Phlebitis; Local reaction/infection; Mild Pain | 2 | 1 | Device and packaging are designed to withstand the sterilization process | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Sterility compromised due to loss of packaging integrity | Temporary discomfort | 1 | 1 | Device and packaging are designed to withstand the sterilization process | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Aseptic techniques not followed | Death; Cardiopulmonary Arrest | 5 | 1 | Device is packaged with protective cap to maintain sterile barrier | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling provides instructions to remove Tego from package using aseptic technique; Labeling includes instructions/precaution to disinfect Tego prior to each use; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Aseptic techniques not followed | Septic shock; Organ injury | 4 | 1 | Device is packaged with protective cap to maintain sterile barrier | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling provides instructions to remove Tego from package using aseptic technique; Labeling includes instructions/precaution to disinfect Tego prior to each use; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Aseptic techniques not followed | Bacteremia; Moderate Pain | 3 | 1 | Device is packaged with protective cap to maintain sterile barrier | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling provides instructions to remove Tego from package using aseptic technique; Labeling includes instructions/precaution to disinfect Tego prior to each use; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Aseptic techniques not followed | Mild Phlebitis; Local reaction/ infection; Mild Pain | 2 | 1 | Device is packaged with protective cap to maintain sterile barrier | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling provides instructions to remove Tego from package using aseptic technique; Labeling includes instructions/precaution to disinfect Tego prior to each use; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Aseptic techniques not followed | Temporary discomfort | 1 | 1 | Device is packaged with protective cap to maintain sterile barrier | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling provides instructions to remove Tego from package using aseptic technique; Labeling includes instructions/precaution to disinfect Tego prior to each use; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Fluid path sterility compromised due to broken, loose or unsealed fluid path | Death; Cardiopulmonary Arrest | 5 | 1 | The device is designed for functional and microbiological efficacy for up to seven days; Cap is designed to remain securely on Tego but be easily removable; Device is designed to be compatible with common disinfectants | Labeling includes precaution statement that the Tego has been validated for functional and microbiological efficacy for up to seven days; Labeling includes a precaution statement to disinfect Tego before each access using 70% isopropyl alcohol in aggressive circular motion for three (3) seconds or per facility protocol.
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Fluid path sterility compromised due to broken, loose or unsealed fluid path | Septic shock; Organ injury | 4 | 1 | The device is designed for functional and microbiological efficacy for up to seven days; Cap is designed to remain securely on Tego but be easily removable; Device is designed to be compatible with common disinfectants | Labeling includes precaution statement that the Tego has been validated for functional and microbiological efficacy for up to seven days; Labeling includes a precaution statement to disinfect Tego before each access using 70% isopropyl alcohol in aggressive circular motion for three (3) seconds or per facility protocol.
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Fluid path sterility compromised due to broken, loose or unsealed fluid path | Bacteremia; Moderate Pain | 3 | 1 | The device is designed for functional and microbiological efficacy for up to seven days; Cap is designed to remain securely on Tego but be easily removable; Device is designed to be compatible with common disinfectants | Labeling includes precaution statement that the Tego has been validated for functional and microbiological efficacy for up to seven days; Labeling includes a precaution statement to disinfect Tego before each access using 70% isopropyl alcohol in aggressive circular motion for three (3) seconds or per facility protocol.
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Fluid path sterility compromised due to broken, loose or unsealed fluid path | Mild Phlebitis; Local reaction/ infection; Mild Pain | 2 | 1 | The device is designed for functional and microbiological efficacy for up to seven days; Cap is designed to remain securely on Tego but be easily removable; Device is designed to be compatible with common disinfectants | Labeling includes precaution statement that the Tego has been validated for functional and microbiological efficacy for up to seven days; Labeling includes a precaution statement to disinfect Tego before each access using 70% isopropyl alcohol in aggressive circular motion for three (3) seconds or per facility protocol.
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Fluid path sterility compromised due to broken, loose or unsealed fluid path | Temporary discomfort | 1 | 1 | The device is designed for functional and microbiological efficacy for up to seven days; Cap is designed to remain securely on Tego but be easily removable; Device is designed to be compatible with common disinfectants | Labeling includes precaution statement that the Tego has been validated for functional and microbiological efficacy for up to seven days; Labeling includes a precaution statement to disinfect Tego before each access using 70% isopropyl alcohol in aggressive circular motion for three (3) seconds or per facility protocol.
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Sterility of device is compromised | Death; Cardiopulmonary Arrest | 5 | 1 | Device and packaging are designed to withstand the sterilization process; Device and packaging are designed to withstand the sterilization process; EO sterilization residuals are below the following maximum allowable limits. | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Sterility of device is compromised | Septic shock; Organ injury | 4 | 1 | Device and packaging are designed to withstand the sterilization process; Device and packaging are designed to withstand the sterilization process; EO sterilization residuals are below the following maximum allowable limits. | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Sterility of device is compromised | Bacteremia; Moderate Pain | 3 | 1 | Device and packaging are designed to withstand the sterilization process; Device and packaging are designed to withstand the sterilization process; EO sterilization residuals are below the following maximum allowable limits. | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Sterility of device is compromised | Mild Phlebitis; Local reaction/ infection; Mild Pain | 2 | 1 | Device and packaging are designed to withstand the sterilization process; Device and packaging are designed to withstand the sterilization process; EO sterilization residuals are below the following maximum allowable limits. | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Sterility of device is compromised | Temporary discomfort | 1 | 1 | Device and packaging are designed to withstand the sterilization process; Device and packaging are designed to withstand the sterilization process; EO sterilization residuals are below the following maximum allowable limits. | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Fluid leak due to defective device (Eg Crack, holes, cuts) | Temporary discomfort | 3 | 1 | Devices is designed not to leak when connected; Device and packaging is designed to withstand shipping and distribution; Device and packaging are designed to withstand the sterilization process | Labeling includes a caution to discard if device packaging is damaged or open; Labeling includes a general warning on the use of a damaged device or improper use; Labeling includes a general warning to inspect device and read instructions for use prior to usage; Labeling includes a precaution statement to clamp line before disconnecting from Tego, and between dialysis sessions per established clinical and catheter manufacturer’s practices; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Fluid leak due to defective device (Eg Crack, holes, cuts) | Systemic infection from virus, bacteria, or other infectious agents | 3 | 1 | Devices is designed not to leak when connected; Device and packaging is designed to withstand shipping and distribution; Device and packaging are designed to withstand the sterilization process | Labeling includes a caution to discard if device packaging is damaged or open; Labeling includes a general warning on the use of a damaged device or improper use; Labeling includes a general warning to inspect device and read instructions for use prior to usage; Labeling includes a precaution statement to clamp line before disconnecting from Tego, and between dialysis sessions per established clinical and catheter manufacturer’s practices; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Fluid leak due to defective device (Eg Crack, holes, cuts) | Mild Cellulitis | 2 | 1 | Devices is designed not to leak when connected; Device and packaging is designed to withstand shipping and distribution; Device and packaging are designed to withstand the sterilization process | Labeling includes a caution to discard if device packaging is damaged or open; Labeling includes a general warning on the use of a damaged device or improper use; Labeling includes a general warning to inspect device and read instructions for use prior to usage; Labeling includes a precaution statement to clamp line before disconnecting from Tego, and between dialysis sessions per established clinical and catheter manufacturer’s practices; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Fluid leak due to defective device (Eg Crack, holes, cuts) | Temporary discomfort | 1 | 1 | Devices is designed not to leak when connected; Device and packaging is designed to withstand shipping and distribution; Device and packaging are designed to withstand the sterilization process | Labeling includes a caution to discard if device packaging is damaged or open; Labeling includes a general warning on the use of a damaged device or improper use; Labeling includes a general warning to inspect device and read instructions for use prior to usage; Labeling includes a precaution statement to clamp line before disconnecting from Tego, and between dialysis sessions per established clinical and catheter manufacturer’s practices; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Fluid leak due to inappropriate disposal of the device | Systemic infection from virus, bacteria, or other infectious agents | 3 | 1 | N/A | Labeling includes instructions to dispose device in accordance with facility protocol
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Fluid leak due to inappropriate disposal of the device | Mild Cellulitis | 2 | 1 | N/A | Labeling includes instructions to dispose device in accordance with facility protocol
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Fluid leak due to inappropriate disposal of the device | Temporary discomfort | 1 | 1 | N/A | Labeling includes instructions to dispose device in accordance with facility protocol
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Systemic infection from virus, bacteria, or other infectious agents | 3 | 1 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Mild Cellulitis | 2 | 1 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Temporary discomfort | 1 | 1 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Biological contaminants | Device Contamination with Biological Material | Skin/Mucous membranes exposure to biological contaminants | Aseptic techniques not followed | Death; Cardiopulmonary Arrest | 5 | 1 | Device is packaged with protective cap to maintain sterile barrier | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling provides instructions to remove Tego from package using aseptic technique; Labeling includes instructions/precaution to disinfect Tego prior to each use; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Biological contaminants | Device Contamination with Biological Material | Skin/Mucous membranes exposure to biological contaminants | Aseptic techniques not followed | Septic shock; Organ injury | 4 | 1 | Device is packaged with protective cap to maintain sterile barrier | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling provides instructions to remove Tego from package using aseptic technique; Labeling includes instructions/precaution to disinfect Tego prior to each use; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Biological contaminants | Device Contamination with Biological Material | Skin/Mucous membranes exposure to biological contaminants | Aseptic techniques not followed | Bacteremia; Moderate Cellulitis; Organ Inflammation (reversible) | 3 | 1 | Device is packaged with protective cap to maintain sterile barrier | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling provides instructions to remove Tego from package using aseptic technique; Labeling includes instructions/precaution to disinfect Tego prior to each use; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Biological contaminants | Device Contamination with Biological Material | Skin/Mucous membranes exposure to biological contaminants | Aseptic techniques not followed | Mild Cellulitis | 2 | 2 | Device is packaged with protective cap to maintain sterile barrier | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling provides instructions to remove Tego from package using aseptic technique; Labeling includes instructions/precaution to disinfect Tego prior to each use; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Biological contaminants | Device Contamination with Biological Material | Skin/Mucous membranes exposure to biological contaminants | Aseptic techniques not followed | Localized skin swelling, redness | 2 | 1 | Device is packaged with protective cap to maintain sterile barrier | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling provides instructions to remove Tego from package using aseptic technique; Labeling includes instructions/precaution to disinfect Tego prior to each use; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Biological contaminants | Device Contamination with Biological Material | Skin/Mucous membranes exposure to biological contaminants | Aseptic techniques not followed | Temporary discomfort | 1 | 1 | Device is packaged with protective cap to maintain sterile barrier | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling provides instructions to remove Tego from package using aseptic technique; Labeling includes instructions/precaution to disinfect Tego prior to each use; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Biological contaminants | Device Contamination with Biological Material | Skin/Mucous membranes exposure to biological contaminants | Fluid leak due to defective device (Eg Crack, holes, cuts) | Death; Cardiopulmonary Arrest | 5 | 1 | Devices is designed not to leak when connected; Device and packaging is designed to withstand shipping and distribution; Device and packaging are designed to withstand the sterilization process | Labeling includes a caution to discard if device packaging is damaged or open; Labeling includes a general warning on the use of a damaged device or improper use; Labeling includes a general warning to inspect device and read instructions for use prior to usage; Labeling includes a precaution statement to clamp line before disconnecting from Tego, and between dialysis sessions per established clinical and catheter manufacturer’s practices; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Biological contaminants | Device Contamination with Biological Material | Skin/Mucous membranes exposure to biological contaminants | Fluid leak due to defective device (Eg Crack, holes, cuts) | Septic shock; Organ injury | 4 | 1 | Devices is designed not to leak when connected; Device and packaging is designed to withstand shipping and distribution; Device and packaging are designed to withstand the sterilization process | Labeling includes a caution to discard if device packaging is damaged or open; Labeling includes a general warning on the use of a damaged device or improper use; Labeling includes a general warning to inspect device and read instructions for use prior to usage; Labeling includes a precaution statement to clamp line before disconnecting from Tego, and between dialysis sessions per established clinical and catheter manufacturer’s practices; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Biological contaminants | Device Contamination with Biological Material | Skin/Mucous membranes exposure to biological contaminants | Fluid leak due to defective device (Eg Crack, holes, cuts) | Bacteremia; Moderate Cellulitis; Organ Inflammation (reversible) | 3 | 1 | Devices is designed not to leak when connected; Device and packaging is designed to withstand shipping and distribution; Device and packaging are designed to withstand the sterilization process | Labeling includes a caution to discard if device packaging is damaged or open; Labeling includes a general warning on the use of a damaged device or improper use; Labeling includes a general warning to inspect device and read instructions for use prior to usage; Labeling includes a precaution statement to clamp line before disconnecting from Tego, and between dialysis sessions per established clinical and catheter manufacturer’s practices; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Biological contaminants | Device Contamination with Biological Material | Skin/Mucous membranes exposure to biological contaminants | Fluid leak due to defective device (Eg Crack, holes, cuts) | Mild Cellulitis | 2 | 2 | Devices is designed not to leak when connected; Device and packaging is designed to withstand shipping and distribution; Device and packaging are designed to withstand the sterilization process | Labeling includes a caution to discard if device packaging is damaged or open; Labeling includes a general warning on the use of a damaged device or improper use; Labeling includes a general warning to inspect device and read instructions for use prior to usage; Labeling includes a precaution statement to clamp line before disconnecting from Tego, and between dialysis sessions per established clinical and catheter manufacturer’s practices; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Biological contaminants | Device Contamination with Biological Material | Skin/Mucous membranes exposure to biological contaminants | Fluid leak due to defective device (Eg Crack, holes, cuts) | Localized skin swelling, redness | 2 | 1 | Devices is designed not to leak when connected; Device and packaging is designed to withstand shipping and distribution; Device and packaging are designed to withstand the sterilization process | Labeling includes a caution to discard if device packaging is damaged or open; Labeling includes a general warning on the use of a damaged device or improper use; Labeling includes a general warning to inspect device and read instructions for use prior to usage; Labeling includes a precaution statement to clamp line before disconnecting from Tego, and between dialysis sessions per established clinical and catheter manufacturer’s practices; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Biological contaminants | Device Contamination with Biological Material | Skin/Mucous membranes exposure to biological contaminants | Fluid leak due to defective device (Eg Crack, holes, cuts) | Temporary discomfort | 1 | 1 | Devices is designed not to leak when connected; Device and packaging is designed to withstand shipping and distribution; Device and packaging are designed to withstand the sterilization process | Labeling includes a caution to discard if device packaging is damaged or open; Labeling includes a general warning on the use of a damaged device or improper use; Labeling includes a general warning to inspect device and read instructions for use prior to usage; Labeling includes a precaution statement to clamp line before disconnecting from Tego, and between dialysis sessions per established clinical and catheter manufacturer’s practices; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Biological contaminants | Device Contamination with Biological Material | Skin/Mucous membranes exposure to biological contaminants | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Temporary discomfort | 1 | 1 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Blood clot | Occlusion within Device | Embolization into the vascular system | Fluid path occlusion (Eg. Kinked tube, particulate etc.) | Death; Cardiopulmonary Arrest | 5 | 1 | Device is designed to have a open fluid path when activated; Device must allow a minimum flow rate | Labeling includes a precaution statement to not use needles, blunt cannulas, or luer-type caps on Tego; Labeling provides instructions to prime device before use; Labeling includes general statement that straight fluid path permits a flow rate of greater than 600 mL/min; Labeling provides instructions to prime device before use
Blood clot | Occlusion within Device | Embolization into the vascular system | Fluid path occlusion (Eg. Kinked tube, particulate etc.) | Pulmonary embolism; Myocardial infarction; Stroke | 4 | 1 | Device is designed to have a open fluid path when activated; Device must allow a minimum flow rate | Labeling includes a precaution statement to not use needles, blunt cannulas, or luer-type caps on Tego; Labeling provides instructions to prime device before use; Labeling includes general statement that straight fluid path permits a flow rate of greater than 600 mL/min; Labeling provides instructions to prime device before use
Blood clot | Occlusion within Device | Embolization into the vascular system | Fluid path occlusion (Eg. Kinked tube, particulate etc.) | Extremity swelling; Moderate Pain; Paresthesias | 3 | 1 | Device is designed to have a open fluid path when activated; Device must allow a minimum flow rate | Labeling includes a precaution statement to not use needles, blunt cannulas, or luer-type caps on Tego; Labeling provides instructions to prime device before use; Labeling includes general statement that straight fluid path permits a flow rate of greater than 600 mL/min; Labeling provides instructions to prime device before use
Blood clot | Occlusion within Device | Embolization into the vascular system | Fluid path occlusion (Eg. Kinked tube, particulate etc.) | Thrombophlebitis or pain | 2 | 1 | Device is designed to have a open fluid path when activated; Device must allow a minimum flow rate | Labeling includes a precaution statement to not use needles, blunt cannulas, or luer-type caps on Tego; Labeling provides instructions to prime device before use; Labeling includes general statement that straight fluid path permits a flow rate of greater than 600 mL/min; Labeling provides instructions to prime device before use
Blood clot | Occlusion within Device | Embolization into the vascular system | Fluid path occlusion (Eg. Kinked tube, particulate etc.) | Temporary Discomfort | 1 | 1 | Device is designed to have a open fluid path when activated; Device must allow a minimum flow rate | Labeling includes a precaution statement to not use needles, blunt cannulas, or luer-type caps on Tego; Labeling provides instructions to prime device before use; Labeling includes general statement that straight fluid path permits a flow rate of greater than 600 mL/min; Labeling provides instructions to prime device before use
Chemical-Other toxin | Biocompatibility Issue | Intravenous exposure to non-hemocompatible materials | Non-hemocompatible material in the fluid path | Death; Cardiopulmonary Arrest | 5 | N/A | Device is designed with biocompatible material | Labeling provides information that the device does not contain materials with DEHP or natural rubber latex in the manufacturing of the device
Chemical-Other toxin | Biocompatibility Issue | Intravenous exposure to non-hemocompatible materials | Non-hemocompatible material in the fluid path | Severe Hemolytic Anemia; Severe Hyperkalemia; Hypotension; Congestive Heart Failure; Acute Kidney Injury | 4 | 1 | Device is designed with biocompatible material | Labeling provides information that the device does not contain materials with DEHP or natural rubber latex in the manufacturing of the device
Chemical-Other toxin | Biocompatibility Issue | Intravenous exposure to non-hemocompatible materials | Non-hemocompatible material in the fluid path | Mild to Moderate Hemolytic Anemia; Mild to Moderate Hyperkalemia; Fever | 3 | 1 | Device is designed with biocompatible material | Labeling provides information that the device does not contain materials with DEHP or natural rubber latex in the manufacturing of the device
Chemical-Other toxin | Biocompatibility Issue | Intravenous exposure to non-hemocompatible materials | Non-hemocompatible material in the fluid path | Temporary discomfort | 1 | 1 | Device is designed with biocompatible material | Labeling provides information that the device does not contain materials with DEHP or natural rubber latex in the manufacturing of the device
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Device appears unsafe to use and needs to be replaced (Eg. Open package, Missing / Loose / Damaged components, Discoloration, Visible particulate etc.) | Mild to Moderate Hypotension; Mild to Moderate Tachycardia | 3 | 1 | Device and packaging is designed to withstand shipping and distribution; Device and packaging are designed to withstand the sterilization process | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes precaution statement that device should be changed in accordance with current, recognized guidelines of IV therapy; Labeling includes general warning to inspect and read instructions for use prior to usage; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination; Labeling includes general warning to inspect and read instructions for use prior to usage
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Device appears unsafe to use and needs to be replaced (Eg. Open package, Missing / Loose / Damaged components, Discoloration, Visible particulate etc.) | Mild Pain | 2 | 1 | Device and packaging is designed to withstand shipping and distribution; Device and packaging are designed to withstand the sterilization process | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes precaution statement that device should be changed in accordance with current, recognized guidelines of IV therapy; Labeling includes general warning to inspect and read instructions for use prior to usage; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination; Labeling includes general warning to inspect and read instructions for use prior to usage
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Device appears unsafe to use and needs to be replaced (Eg. Open package, Missing / Loose / Damaged components, Discoloration, Visible particulate etc.) | Temporary discomfort | 1 | 1 | Device and packaging is designed to withstand shipping and distribution; Device and packaging are designed to withstand the sterilization process | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes precaution statement that device should be changed in accordance with current, recognized guidelines of IV therapy; Labeling includes general warning to inspect and read instructions for use prior to usage; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination; Labeling includes general warning to inspect and read instructions for use prior to usage
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Air is noticed in the fluid path causing the clinician to reprime or change the set | Mild to Moderate Hypotension; Mild to Moderate Tachycardia | 3 | 1 | Tego is primed according to facility protocol | Labeling provides instructions to prime device before use
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Air is noticed in the fluid path causing the clinician to reprime or change the set | Mild Pain | 2 | 1 | Tego is primed according to facility protocol | Labeling provides instructions to prime device before use
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Air is noticed in the fluid path causing the clinician to reprime or change the set | Temporary discomfort | 1 | 1 | Tego is primed according to facility protocol | Labeling provides instructions to prime device before use
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Fluid leak due to defective device (Eg Crack, holes, cuts) | Mild to Moderate Hypotension; Mild to Moderate Tachycardia | 3 | 1 | Devices is designed not to leak when connected; Device and packaging is designed to withstand shipping and distribution; Device and packaging are designed to withstand the sterilization process | Labeling includes a caution to discard if device packaging is damaged or open; Labeling includes a general warning on the use of a damaged device or improper use; Labeling includes a general warning to inspect device and read instructions for use prior to usage; Labeling includes a precaution statement to clamp line before disconnecting from Tego, and between dialysis sessions per established clinical and catheter manufacturer’s practices; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Fluid leak due to defective device (Eg Crack, holes, cuts) | Mild Pain | 2 | 1 | Devices is designed not to leak when connected; Device and packaging is designed to withstand shipping and distribution; Device and packaging are designed to withstand the sterilization process | Labeling includes a caution to discard if device packaging is damaged or open; Labeling includes a general warning on the use of a damaged device or improper use; Labeling includes a general warning to inspect device and read instructions for use prior to usage; Labeling includes a precaution statement to clamp line before disconnecting from Tego, and between dialysis sessions per established clinical and catheter manufacturer’s practices; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Fluid leak due to defective device (Eg Crack, holes, cuts) | Temporary discomfort | 1 | 1 | Devices is designed not to leak when connected; Device and packaging is designed to withstand shipping and distribution; Device and packaging are designed to withstand the sterilization process | Labeling includes a caution to discard if device packaging is damaged or open; Labeling includes a general warning on the use of a damaged device or improper use; Labeling includes a general warning to inspect device and read instructions for use prior to usage; Labeling includes a precaution statement to clamp line before disconnecting from Tego, and between dialysis sessions per established clinical and catheter manufacturer’s practices; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Mild to Moderate Hypotension; Mild to Moderate Tachycardia | 3 | 1 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Mild Pain | 2 | 1 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Temporary discomfort | 1 | 1 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Problem identifying / selecting the device | Mild to Moderate Hypotension; Mild to Moderate Tachycardia | 3 | 1 | N/A | Labeling includes a caution statement that: device is intended to be used and handled by trained clinical personnel only; Care should be exercised to prevent unintended handling by children; and Federal (USA) law restricts the use of this device to sale by or on the order of a physician (for USA only).
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Problem identifying / selecting the device | Mild Pain | 2 | 1 | N/A | Labeling includes a caution statement that: device is intended to be used and handled by trained clinical personnel only; Care should be exercised to prevent unintended handling by children; and Federal (USA) law restricts the use of this device to sale by or on the order of a physician (for USA only).
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Problem identifying / selecting the device | Temporary discomfort | 1 | 1 | N/A | Labeling includes a caution statement that: device is intended to be used and handled by trained clinical personnel only; Care should be exercised to prevent unintended handling by children; and Federal (USA) law restricts the use of this device to sale by or on the order of a physician (for USA only).
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Unable to remove cap from the device | Mild to Moderate Hypotension; Mild to Moderate Tachycardia | 3 | 1 | Cap on device is designed for easy removal | Labeling provides instructions to remove cap when attaching Tego to compatible device
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Unable to remove cap from the device | Mild Pain | 2 | 1 | Cap on device is designed for easy removal | Labeling provides instructions to remove cap when attaching Tego to compatible device
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Unable to remove cap from the device | Temporary discomfort | 1 | 1 | Cap on device is designed for easy removal | Labeling provides instructions to remove cap when attaching Tego to compatible device
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Difficult / Unable to connect / disconnect the device with other mating devices | Mild to Moderate Hypotension; Mild to Moderate Tachycardia | 3 | 1 | Device is designed for use with luer; Product is designed with ribs to allow easier handling for connection | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a caution statement that: device is intended to be used and handled by trained clinical personnel only; Care should be exercised to prevent unintended handling by children; and Federal (USA) law restricts the use of this device to sale by or on the order of a physician (for USA only).
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Difficult / Unable to connect / disconnect the device with other mating devices | Mild Pain | 2 | 1 | Device is designed for use with luer; Product is designed with ribs to allow easier handling for connection | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a caution statement that: device is intended to be used and handled by trained clinical personnel only; Care should be exercised to prevent unintended handling by children; and Federal (USA) law restricts the use of this device to sale by or on the order of a physician (for USA only).
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Difficult / Unable to connect / disconnect the device with other mating devices | Temporary discomfort | 1 | 1 | Device is designed for use with luer; Product is designed with ribs to allow easier handling for connection | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a caution statement that: device is intended to be used and handled by trained clinical personnel only; Care should be exercised to prevent unintended handling by children; and Federal (USA) law restricts the use of this device to sale by or on the order of a physician (for USA only).
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Difficult / Unable to open the packaging | Mild to Moderate Hypotension; Mild to Moderate Tachycardia | 3 | 1 | Device is designed to have a maximum seal strength | Labeling includes a caution to discard if device packaging is damaged or open;
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Difficult / Unable to open the packaging | Mild Pain | 2 | 1 | Device is designed to have a maximum seal strength | Labeling includes a caution to discard if device packaging is damaged or open;
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Difficult / Unable to open the packaging | Temporary discomfort | 1 | 1 | Device is designed to have a maximum seal strength | Labeling includes a caution to discard if device packaging is damaged or open;
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Mild to Moderate Hypotension; Mild to Moderate Tachycardia | 3 | 1 | Device is designed for use with luer | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Mild Pain | 2 | 1 | Device is designed for use with luer | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion of blood products (blood product does not need to be replaced) | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Temporary discomfort | 1 | 1 | Device is designed for use with luer | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Fluid path occlusion (Eg. Kinked tube, particulate etc.) | Death; Cardiopulmonary Arrest | 5 | 1 | Device is designed to have a open fluid path when activated | Labeling includes a precaution statement to not use needles, blunt cannulas, or luer-type caps on Tego; Labeling provides instructions to prime device before use
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Fluid path occlusion (Eg. Kinked tube, particulate etc.) | Severe Arrhythmias; Severe Hypertension; Severe Hypotension; Status Epilepticus | 4 | 1 | Device is designed to have a open fluid path when activated | Labeling includes a precaution statement to not use needles, blunt cannulas, or luer-type caps on Tego; Labeling provides instructions to prime device before use
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Fluid path occlusion (Eg. Kinked tube, particulate etc.) | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 1 | Device is designed to have a open fluid path when activated | Labeling includes a precaution statement to not use needles, blunt cannulas, or luer-type caps on Tego; Labeling provides instructions to prime device before use
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Fluid path occlusion (Eg. Kinked tube, particulate etc.) | Mild to Moderate Pain | 2 | 1 | Device is designed to have a open fluid path when activated | Labeling includes a precaution statement to not use needles, blunt cannulas, or luer-type caps on Tego; Labeling provides instructions to prime device before use
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Fluid path occlusion (Eg. Kinked tube, particulate etc.) | Temporary discomfort | 1 | 2 | Device is designed to have a open fluid path when activated | Labeling includes a precaution statement to not use needles, blunt cannulas, or luer-type caps on Tego; Labeling provides instructions to prime device before use
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Air is noticed in the fluid path causing the clinician to reprime or change the set | Death; Cardiopulmonary Arrest | 5 | 1 | N/A | Labeling includes a caution statement that: device is intended to be used and handled by trained clinical personnel only; Care should be exercised to prevent unintended handling by children; and Federal (USA) law restricts the use of this device to sale by or on the order of a physician (for USA only).
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Air is noticed in the fluid path causing the clinician to reprime or change the set | Severe Arrhythmias; Severe Hypertension; Severe Hypotension; Status Epilepticus | 4 | 1 | N/A | Labeling includes a caution statement that: device is intended to be used and handled by trained clinical personnel only; Care should be exercised to prevent unintended handling by children; and Federal (USA) law restricts the use of this device to sale by or on the order of a physician (for USA only).
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Air is noticed in the fluid path causing the clinician to reprime or change the set | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 1 | N/A | Labeling includes a caution statement that: device is intended to be used and handled by trained clinical personnel only; Care should be exercised to prevent unintended handling by children; and Federal (USA) law restricts the use of this device to sale by or on the order of a physician (for USA only).
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Air is noticed in the fluid path causing the clinician to reprime or change the set | Mild to Moderate Pain | 2 | 1 | N/A | Labeling includes a caution statement that: device is intended to be used and handled by trained clinical personnel only; Care should be exercised to prevent unintended handling by children; and Federal (USA) law restricts the use of this device to sale by or on the order of a physician (for USA only).
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Air is noticed in the fluid path causing the clinician to reprime or change the set | Temporary discomfort | 1 | 2 | N/A | Labeling includes a caution statement that: device is intended to be used and handled by trained clinical personnel only; Care should be exercised to prevent unintended handling by children; and Federal (USA) law restricts the use of this device to sale by or on the order of a physician (for USA only).
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Fluid leak due to defective device (Eg Crack, holes, cuts) | Death; Cardiopulmonary Arrest | 5 | 1 | Devices is designed not to leak when connected; Device and packaging is designed to withstand shipping and distribution; Device and packaging are designed to withstand the sterilization process | Labeling includes a caution to discard if device packaging is damaged or open; Labeling includes a general warning on the use of a damaged device or improper use; Labeling includes a general warning to inspect device and read instructions for use prior to usage; Labeling includes a precaution statement to clamp line before disconnecting from Tego, and between dialysis sessions per established clinical and catheter manufacturer’s practices; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Fluid leak due to defective device (Eg Crack, holes, cuts) | Severe Arrhythmias; Severe Hypertension; Severe Hypotension; Status Epilepticus | 4 | 1 | Devices is designed not to leak when connected; Device and packaging is designed to withstand shipping and distribution; Device and packaging are designed to withstand the sterilization process | Labeling includes a caution to discard if device packaging is damaged or open; Labeling includes a general warning on the use of a damaged device or improper use; Labeling includes a general warning to inspect device and read instructions for use prior to usage; Labeling includes a precaution statement to clamp line before disconnecting from Tego, and between dialysis sessions per established clinical and catheter manufacturer’s practices; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Fluid leak due to defective device (Eg Crack, holes, cuts) | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 1 | Devices is designed not to leak when connected; Device and packaging is designed to withstand shipping and distribution; Device and packaging are designed to withstand the sterilization process | Labeling includes a caution to discard if device packaging is damaged or open; Labeling includes a general warning on the use of a damaged device or improper use; Labeling includes a general warning to inspect device and read instructions for use prior to usage; Labeling includes a precaution statement to clamp line before disconnecting from Tego, and between dialysis sessions per established clinical and catheter manufacturer’s practices; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Fluid leak due to defective device (Eg Crack, holes, cuts) | Mild to Moderate Pain | 2 | 1 | Devices is designed not to leak when connected; Device and packaging is designed to withstand shipping and distribution; Device and packaging are designed to withstand the sterilization process | Labeling includes a caution to discard if device packaging is damaged or open; Labeling includes a general warning on the use of a damaged device or improper use; Labeling includes a general warning to inspect device and read instructions for use prior to usage; Labeling includes a precaution statement to clamp line before disconnecting from Tego, and between dialysis sessions per established clinical and catheter manufacturer’s practices; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Fluid leak due to defective device (Eg Crack, holes, cuts) | Temporary discomfort | 1 | 2 | Devices is designed not to leak when connected; Device and packaging is designed to withstand shipping and distribution; Device and packaging are designed to withstand the sterilization process | Labeling includes a caution to discard if device packaging is damaged or open; Labeling includes a general warning on the use of a damaged device or improper use; Labeling includes a general warning to inspect device and read instructions for use prior to usage; Labeling includes a precaution statement to clamp line before disconnecting from Tego, and between dialysis sessions per established clinical and catheter manufacturer’s practices; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Death; Cardiopulmonary Arrest | 5 | 1 | Device is designed for use with luer | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Severe Arrhythmias; Severe Hypertension; Severe Hypotension; Status Epilepticus | 4 | 1 | Device is designed for use with luer | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 1 | Device is designed for use with luer | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Mild to Moderate Pain | 2 | 1 | Device is designed for use with luer | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Temporary discomfort | 1 | 2 | Device is designed for use with luer | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Death; Cardiopulmonary Arrest | 5 | 1 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Severe Arrhythmias; Severe Hypertension; Severe Hypotension; Status Epilepticus | 4 | 1 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 1 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Mild to Moderate Pain | 2 | 1 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Temporary discomfort | 1 | 2 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Difficult / Unable to connect / disconnect the device with other mating devices | Death; Cardiopulmonary Arrest | 5 | 1 | Device is designed for use with luer; Product is designed with ribs to allow easier handling for connection | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Difficult / Unable to connect / disconnect the device with other mating devices | Severe Arrhythmias; Severe Hypertension; Severe Hypotension; Status Epilepticus | 4 | 1 | Device is designed for use with luer; Product is designed with ribs to allow easier handling for connection | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Difficult / Unable to connect / disconnect the device with other mating devices | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 1 | Device is designed for use with luer; Product is designed with ribs to allow easier handling for connection | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Difficult / Unable to connect / disconnect the device with other mating devices | Mild to Moderate Pain | 2 | 1 | Device is designed for use with luer; Product is designed with ribs to allow easier handling for connection | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Difficult / Unable to connect / disconnect the device with other mating devices | Temporary discomfort | 1 | 2 | Device is designed for use with luer; Product is designed with ribs to allow easier handling for connection | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Patient movement causes set to disconnect at a connection point | Death; Cardiopulmonary Arrest | 5 | 1 | Device is designed to have ISO 80369-7 male luer; Device is designed to comply with ISO 80369-7 male luers | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Patient movement causes set to disconnect at a connection point | Severe Arrhythmias; Severe Hypertension; Severe Hypotension; Status Epilepticus | 4 | 1 | Device is designed to have ISO 80369-7 male luer; Device is designed to comply with ISO 80369-7 male luers | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Patient movement causes set to disconnect at a connection point | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 1 | Device is designed to have ISO 80369-7 male luer; Device is designed to comply with ISO 80369-7 male luers | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Patient movement causes set to disconnect at a connection point | Mild to Moderate Pain | 2 | 1 | Device is designed to have ISO 80369-7 male luer; Device is designed to comply with ISO 80369-7 male luers | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Patient movement causes set to disconnect at a connection point | Temporary discomfort | 1 | 2 | Device is designed to have ISO 80369-7 male luer; Device is designed to comply with ISO 80369-7 male luers | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Drug with altered activity | Medication Error | Infusion of drug with altered activity | Precipitate from drug / device interaction is introduced into the fluid path | Death; Cardiopulmonary Arrest | 5 | 1 | Device material is designed not to interact with drugs used with this application | Labeling includes a caution statement that: device is intended to be used and handled by trained clinical personnel only; Care should be exercised to prevent unintended handling by children; and Federal (USA) law restricts the use of this device to sale by or on the order of a physician (for USA only).
Drug with altered activity | Medication Error | Infusion of drug with altered activity | Precipitate from drug / device interaction is introduced into the fluid path | Pulmonary embolism; Septic shock; Severe Hyperglycemia; Severe Hypertension; Severe Hypoglycemia; Severe Hypotension; Severe Arrhythmia; Severe Pain | 4 | 1 | Device material is designed not to interact with drugs used with this application | Labeling includes a caution statement that: device is intended to be used and handled by trained clinical personnel only; Care should be exercised to prevent unintended handling by children; and Federal (USA) law restricts the use of this device to sale by or on the order of a physician (for USA only).
Drug with altered activity | Medication Error | Infusion of drug with altered activity | Precipitate from drug / device interaction is introduced into the fluid path | Mild to Moderate Arrhythmia; Moderate Hyperglycemia; Mild to Moderate Hypertension; Moderate Hypoglycemia; Mild to Moderate Hypotension; Moderate Pain | 3 | 1 | Device material is designed not to interact with drugs used with this application | Labeling includes a caution statement that: device is intended to be used and handled by trained clinical personnel only; Care should be exercised to prevent unintended handling by children; and Federal (USA) law restricts the use of this device to sale by or on the order of a physician (for USA only).
Drug with altered activity | Medication Error | Infusion of drug with altered activity | Precipitate from drug / device interaction is introduced into the fluid path | Mild Pain; Mild Hyperglycemia; Mild Hypoglycemia | 2 | 1 | Device material is designed not to interact with drugs used with this application | Labeling includes a caution statement that: device is intended to be used and handled by trained clinical personnel only; Care should be exercised to prevent unintended handling by children; and Federal (USA) law restricts the use of this device to sale by or on the order of a physician (for USA only).
Drug with altered activity | Medication Error | Infusion of drug with altered activity | Precipitate from drug / device interaction is introduced into the fluid path | Temporary discomfort | 1 | 1 | Device material is designed not to interact with drugs used with this application | Labeling includes a caution statement that: device is intended to be used and handled by trained clinical personnel only; Care should be exercised to prevent unintended handling by children; and Federal (USA) law restricts the use of this device to sale by or on the order of a physician (for USA only).
Ergonomic stress | Inadequate User Interface | Exposure to ergonomic stress | Difficult / Unable to connect / disconnect the device with other mating devices | Temporary discomfort | 1 | 1 | Device is designed for use with luer; Product is designed with ribs to allow easier handling for connection | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Ergonomic stress | Inadequate User Interface | Exposure to ergonomic stress | Difficult / Unable to open the packaging | Temporary discomfort | 1 | 1 | Device is designed to have a maximum seal strength | Labeling includes a caution to discard or do not use if device packaging is damaged or open;
Fluid leakage | Fluid Leak | Slips/Falls | Fluid leak due to defective device (Eg Crack, holes, cuts) | Death; Cardiopulmonary Arrest | 5 | 1 | Devices is designed not to leak when connected; Device and packaging is designed to withstand shipping and distribution; Device and packaging are designed to withstand the sterilization process | Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes a general warning on the use of a damaged device or improper use; Labeling includes a general warning to inspect device and read instructions for use prior to usage; Labeling includes a precaution statement to clamp line before disconnecting from Tego, and between dialysis sessions per established clinical and catheter manufacturer’s practices; Labeling includes a precaution to tighten and secure connections; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Fluid leakage | Fluid Leak | Slips/Falls | Fluid leak due to defective device (Eg Crack, holes, cuts) | Head Trauma with CNS injury | 4 | 1 | Devices is designed not to leak when connected; Device and packaging is designed to withstand shipping and distribution; Device and packaging are designed to withstand the sterilization process | Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes a general warning on the use of a damaged device or improper use; Labeling includes a general warning to inspect device and read instructions for use prior to usage; Labeling includes a precaution statement to clamp line before disconnecting from Tego, and between dialysis sessions per established clinical and catheter manufacturer’s practices; Labeling includes a precaution to tighten and secure connections; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Fluid leakage | Fluid Leak | Slips/Falls | Fluid leak due to defective device (Eg Crack, holes, cuts) | Fracture, Sprain, Strain, Laceration; Moderate Pain | 3 | 1 | Devices is designed not to leak when connected; Device and packaging is designed to withstand shipping and distribution; Device and packaging are designed to withstand the sterilization process | Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes a general warning on the use of a damaged device or improper use; Labeling includes a general warning to inspect device and read instructions for use prior to usage; Labeling includes a precaution statement to clamp line before disconnecting from Tego, and between dialysis sessions per established clinical and catheter manufacturer’s practices; Labeling includes a precaution to tighten and secure connections; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Fluid leakage | Fluid Leak | Slips/Falls | Fluid leak due to defective device (Eg Crack, holes, cuts) | Bruise; Skin Abrasion; Mild Pain | 2 | 2 | Devices is designed not to leak when connected; Device and packaging is designed to withstand shipping and distribution; Device and packaging are designed to withstand the sterilization process | Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes a general warning on the use of a damaged device or improper use; Labeling includes a general warning to inspect device and read instructions for use prior to usage; Labeling includes a precaution statement to clamp line before disconnecting from Tego, and between dialysis sessions per established clinical and catheter manufacturer’s practices; Labeling includes a precaution to tighten and secure connections; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Fluid leakage | Fluid Leak | Slips/Falls | Fluid leak due to defective device (Eg Crack, holes, cuts) | Temporary discomfort | 1 | 1 | Devices is designed not to leak when connected; Device and packaging is designed to withstand shipping and distribution; Device and packaging are designed to withstand the sterilization process | Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes a general warning on the use of a damaged device or improper use; Labeling includes a general warning to inspect device and read instructions for use prior to usage; Labeling includes a precaution statement to clamp line before disconnecting from Tego, and between dialysis sessions per established clinical and catheter manufacturer’s practices; Labeling includes a precaution to tighten and secure connections; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Fluid leakage | Fluid Leak | Slips/Falls | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Death; Cardiopulmonary Arrest | 5 | 1 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution to tighten and secure connections
Fluid leakage | Fluid Leak | Slips/Falls | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Head Trauma with CNS injury | 4 | 1 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution to tighten and secure connections
Fluid leakage | Fluid Leak | Slips/Falls | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Fracture, Sprain, Strain Laceration; Moderate Pain | 3 | 1 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution to tighten and secure connections
Fluid leakage | Fluid Leak | Slips/Falls | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Bruise; Skin Abrasion; Mild Pain | 2 | 2 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution to tighten and secure connections
Fluid leakage | Fluid Leak | Slips/Falls | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Temporary discomfort | 1 | 1 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models); Labeling includes a precaution to tighten and secure connections
Inappropriate disposal | Environmental Control or Utility Issue | Environmental (Landfill/Water Supply) exposure to toxic substances from device material | Proper disposal procedure not followed | Moderate toxic effects such as seizures, abdominal pain | 3 | 1 | N/A | Labeling includes instructions to dispose device in accordance with facility protocol; Labeling includes precaution statement that device should be changed in accordance with current, recognized guidelines of IV therapy.
Inappropriate disposal | Environmental Control or Utility Issue | Environmental (Landfill/Water Supply) exposure to toxic substances from device material | Proper disposal procedure not followed | Mild toxic effects such as fatigue, sleepiness, headaches, eye irritation, skin rashes | 2 | 1 | N/A | Labeling includes instructions to dispose device in accordance with facility protocol; Labeling includes precaution statement that device should be changed in accordance with current, recognized guidelines of IV therapy.
Inappropriate disposal | Environmental Control or Utility Issue | Environmental (Landfill/Water Supply) exposure to toxic substances from device material | Proper disposal procedure not followed | Temporary discomfort | 1 | 1 | N/A | Labeling includes instructions to dispose device in accordance with facility protocol; Labeling includes precaution statement that device should be changed in accordance with current, recognized guidelines of IV therapy.
Inappropriate disposal | Environmental Control or Utility Issue | Environmental (Landfill/Water Supply) exposure to toxic substances from device material | Device contains toxic material | Moderate toxic effects such as seizures, abdominal pain | 3 | 1 | Device is designed with non-toxic material | Labeling provides information that the device does not contain materials with DEHP or natural rubber latex in the manufacturing of the device
Inappropriate disposal | Environmental Control or Utility Issue | Environmental (Landfill/Water Supply) exposure to toxic substances from device material | Device contains toxic material | Mild toxic effects such as fatigue, sleepiness, headaches, eye irritation, skin rashes | 2 | 1 | Device is designed with non-toxic material | Labeling provides information that the device does not contain materials with DEHP or natural rubber latex in the manufacturing of the device
Inappropriate disposal | Environmental Control or Utility Issue | Environmental (Landfill/Water Supply) exposure to toxic substances from device material | Device contains toxic material | Temporary discomfort | 1 | 1 | Device is designed with non-toxic material | Labeling provides information that the device does not contain materials with DEHP or natural rubber latex in the manufacturing of the device
Open fluid path | Fluid Leak | Patient blood loss | Fluid leak due to defective device (Eg Crack, holes, cuts) | Death | 5 | 1 | Devices is designed not to leak when connected; Device and packaging is designed to withstand shipping and distribution; Device and packaging are designed to withstand the sterilization process | Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes a general warning on the use of a damaged device or improper use; Labeling includes a general warning to inspect device and read instructions for use prior to usage; Labeling includes a precaution statement to clamp line before disconnecting from Tego, and between dialysis sessions per established clinical and catheter manufacturer’s practices; Labeling includes a precaution to tighten and secure connections; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Open fluid path | Fluid Leak | Patient blood loss | Fluid leak due to defective device (Eg Crack, holes, cuts) | Hypovolemic shock | 4 | 1 | Devices is designed not to leak when connected; Device and packaging is designed to withstand shipping and distribution; Device and packaging are designed to withstand the sterilization process | Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes a general warning on the use of a damaged device or improper use; Labeling includes a general warning to inspect device and read instructions for use prior to usage; Labeling includes a precaution statement to clamp line before disconnecting from Tego, and between dialysis sessions per established clinical and catheter manufacturer’s practices; Labeling includes a precaution to tighten and secure connections; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Open fluid path | Fluid Leak | Patient blood loss | Fluid leak due to defective device (Eg Crack, holes, cuts) | Mild to Moderate Hypotension; Mild to Moderate Hypovolemia | 3 | 1 | Devices is designed not to leak when connected; Device and packaging is designed to withstand shipping and distribution; Device and packaging are designed to withstand the sterilization process | Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes a general warning on the use of a damaged device or improper use; Labeling includes a general warning to inspect device and read instructions for use prior to usage; Labeling includes a precaution statement to clamp line before disconnecting from Tego, and between dialysis sessions per established clinical and catheter manufacturer’s practices; Labeling includes a precaution to tighten and secure connections; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Open fluid path | Fluid Leak | Patient blood loss | Fluid leak due to defective device (Eg Crack, holes, cuts) | Temporary Discomfort | 1 | 2 | Devices is designed not to leak when connected; Device and packaging is designed to withstand shipping and distribution; Device and packaging are designed to withstand the sterilization process | Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes a general warning on the use of a damaged device or improper use; Labeling includes a general warning to inspect device and read instructions for use prior to usage; Labeling includes a precaution statement to clamp line before disconnecting from Tego, and between dialysis sessions per established clinical and catheter manufacturer’s practices; Labeling includes a precaution to tighten and secure connections; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Open fluid path | Fluid Leak | Patient blood loss | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Death | 5 | 1 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Open fluid path | Fluid Leak | Patient blood loss | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Hypovolemic shock | 4 | 1 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Open fluid path | Fluid Leak | Patient blood loss | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Mild to Moderate Hypotension; Mild to Moderate Hypovolemia | 3 | 1 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Open fluid path | Fluid Leak | Patient blood loss | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Temporary Discomfort | 1 | 2 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Open fluid path | Fluid Leak | Patient blood loss | Patient movement causes set to disconnect at a connection point | Death | 5 | 1 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Open fluid path | Fluid Leak | Patient blood loss | Patient movement causes set to disconnect at a connection point | Hypovolemic shock | 4 | 1 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Open fluid path | Fluid Leak | Patient blood loss | Patient movement causes set to disconnect at a connection point | Mild to Moderate Hypotension; Mild to Moderate Hypovolemia | 3 | 1 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Open fluid path | Fluid Leak | Patient blood loss | Patient movement causes set to disconnect at a connection point | Temporary Discomfort | 1 | 2 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Open fluid path | Fluid Leak | Patient blood loss | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Death | 5 | 1 | Device is designed for use with luer | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Open fluid path | Fluid Leak | Patient blood loss | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Hypovolemic shock | 4 | 1 | Device is designed for use with luer | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Open fluid path | Fluid Leak | Patient blood loss | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Mild to Moderate Hypotension; Mild to Moderate Hypovolemia | 3 | 1 | Device is designed for use with luer | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Open fluid path | Fluid Leak | Patient blood loss | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Temporary Discomfort | 1 | 2 | Device is designed for use with luer | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
particulate presented | Particulates | Embolization into the vascular system | Fluid path occlusion (Eg. Kinked tube, particulate etc.) | Death; Cardiopulmonary Arrest | 5 | 1 | Device is designed to have a open fluid path when activated | Labeling includes a precaution statement to not use needles, blunt cannulas, or luer-type caps on Tego; Labeling provides instructions to prime device before use
particulate presented | Particulates | Embolization into the vascular system | Fluid path occlusion (Eg. Kinked tube, particulate etc.) | Pulmonary embolism; Myocardial infarction; Stroke | 4 | 1 | Device is designed to have a open fluid path when activated | Labeling includes a precaution statement to not use needles, blunt cannulas, or luer-type caps on Tego; Labeling provides instructions to prime device before use
particulate presented | Particulates | Embolization into the vascular system | Fluid path occlusion (Eg. Kinked tube, particulate etc.) | Extremity swelling; Moderate Pain; Paresthesias | 3 | 1 | Device is designed to have a open fluid path when activated | Labeling includes a precaution statement to not use needles, blunt cannulas, or luer-type caps on Tego; Labeling provides instructions to prime device before use
particulate presented | Particulates | Embolization into the vascular system | Fluid path occlusion (Eg. Kinked tube, particulate etc.) | Thrombophlebitis; Mild Pain | 2 | 1 | Device is designed to have a open fluid path when activated | Labeling includes a precaution statement to not use needles, blunt cannulas, or luer-type caps on Tego; Labeling provides instructions to prime device before use
particulate presented | Particulates | Embolization into the vascular system | Fluid path occlusion (Eg. Kinked tube, particulate etc.) | Temporary Discomfort | 1 | 2 | Device is designed to have a open fluid path when activated | Labeling includes a precaution statement to not use needles, blunt cannulas, or luer-type caps on Tego; Labeling provides instructions to prime device before use
particulate presented | Particulates | Embolization into the vascular system | Particulate in the fluid path | Death; Cardiopulmonary Arrest | 5 | 1 | Fluid path particulate contamination index shall not exceed 90 | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling provides instructions to remove Tego from package using aseptic technique; Labeling includes instructions/precaution to disinfect Tego prior to each use; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
particulate presented | Particulates | Embolization into the vascular system | Particulate in the fluid path | Pulmonary embolism; Myocardial infarction; Stroke | 4 | 1 | Fluid path particulate contamination index shall not exceed 91 | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling provides instructions to remove Tego from package using aseptic technique; Labeling includes instructions/precaution to disinfect Tego prior to each use; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
particulate presented | Particulates | Embolization into the vascular system | Particulate in the fluid path | Extremity swelling; Moderate Pain; Paresthesias | 3 | 1 | Fluid path particulate contamination index shall not exceed 92 | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling provides instructions to remove Tego from package using aseptic technique; Labeling includes instructions/precaution to disinfect Tego prior to each use; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
particulate presented | Particulates | Embolization into the vascular system | Particulate in the fluid path | Thrombophlebitis; Mild Pain | 2 | 1 | Fluid path particulate contamination index shall not exceed 93 | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling provides instructions to remove Tego from package using aseptic technique; Labeling includes instructions/precaution to disinfect Tego prior to each use; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
particulate presented | Particulates | Embolization into the vascular system | Particulate in the fluid path | Temporary Discomfort | 1 | 2 | Fluid path particulate contamination index shall not exceed 94 | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling provides instructions to remove Tego from package using aseptic technique; Labeling includes instructions/precaution to disinfect Tego prior to each use; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
particulate presented | Particulates | Embolization into the vascular system | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Death; Cardiopulmonary Arrest | 5 | 1 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
particulate presented | Particulates | Embolization into the vascular system | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Pulmonary embolism; Myocardial infarction; Stroke | 4 | 1 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
particulate presented | Particulates | Embolization into the vascular system | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Extremity swelling; Moderate Pain; Paresthesias | 3 | 1 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
particulate presented | Particulates | Embolization into the vascular system | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Thrombophlebitis; Mild Pain | 2 | 1 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
particulate presented | Particulates | Embolization into the vascular system | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Temporary Discomfort | 1 | 2 | Luer connections are designed to prevent disconnection during normal use (per applicable ISO standards) | Labeling includes a precaution statement to access Tego straight on (without angled or sliding entry) with administration device; Labeling provides instructions when using a luer slip connection, to only insert half-way while turning slightly clockwise. When removing apply the same clockwise turning motion. Do not overtighten; Labeling includes a precaution statement for ISO compatibility (specific to D1005 models)
Removable component | Detachment of Device Component | User ingests removable component | Removable components left in patient bed or placed into mouth | Death; Cardiopulmonary Arrest | 5 | 1 | N/A | Labeling includes a caution statement that: device is intended to be used and handled by trained clinical personnel only; Care should be exercised to prevent unintended handling by children; and Federal (USA) law restricts the use of this device to sale by or on the order of a physician (for USA only).
Removable component | Detachment of Device Component | User ingests removable component | Removable components left in patient bed or placed into mouth | Hypoxic organ injury Respiratory Distress-Severe | 4 | 1 | N/A | Labeling includes a caution statement that: device is intended to be used and handled by trained clinical personnel only; Care should be exercised to prevent unintended handling by children; and Federal (USA) law restricts the use of this device to sale by or on the order of a physician (for USA only).
Removable component | Detachment of Device Component | User ingests removable component | Removable components left in patient bed or placed into mouth | Esophageal or respiratory airway obstruction Gastric foreign body | 3 | 1 | N/A | Labeling includes a caution statement that: device is intended to be used and handled by trained clinical personnel only; Care should be exercised to prevent unintended handling by children; and Federal (USA) law restricts the use of this device to sale by or on the order of a physician (for USA only).
Removable component | Detachment of Device Component | User ingests removable component | Removable components left in patient bed or placed into mouth | Coughing or gagging | 2 | 1 | N/A | Labeling includes a caution statement that: device is intended to be used and handled by trained clinical personnel only; Care should be exercised to prevent unintended handling by children; and Federal (USA) law restricts the use of this device to sale by or on the order of a physician (for USA only).
Removable component | Detachment of Device Component | User ingests removable component | Removable components left in patient bed or placed into mouth | Temporary discomfort | 1 | 1 | N/A | Labeling includes a caution statement that: device is intended to be used and handled by trained clinical personnel only; Care should be exercised to prevent unintended handling by children; and Federal (USA) law restricts the use of this device to sale by or on the order of a physician (for USA only).
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Fluid path contains toxins (E.g. systemic, mutagens, carcinogens) | Death; Cardiopulmonary Arrest | 5 | 1 | Device is designed with biocompatible material | Labeling provides information that the device does not contain materials with DEHP or natural rubber latex in the manufacturing of the device
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Fluid path contains toxins (E.g. systemic, mutagens, carcinogens) | Paralysis; Organ failure/injury; Malignancy; Severe Hypotension | 4 | 1 | Device is designed with biocompatible material | Labeling provides information that the device does not contain materials with DEHP or natural rubber latex in the manufacturing of the device
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Fluid path contains toxins (E.g. systemic, mutagens, carcinogens) | Mild to Moderate Hypotension; Moderate Inflammation (e.g. Phlebitis, Organ); Moderate Pain; Fever | 3 | 1 | Device is designed with biocompatible material | Labeling provides information that the device does not contain materials with DEHP or natural rubber latex in the manufacturing of the device
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Fluid path contains toxins (E.g. systemic, mutagens, carcinogens) | Mild Inflammation (e.g. Phlebitis, Organ); Mild Pain | 2 | 1 | Device is designed with biocompatible material | Labeling provides information that the device does not contain materials with DEHP or natural rubber latex in the manufacturing of the device
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Fluid path contains toxins (E.g. systemic, mutagens, carcinogens) | Temporary discomfort | 1 | 1 | Device is designed with biocompatible material | Labeling provides information that the device does not contain materials with DEHP or natural rubber latex in the manufacturing of the device
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Pyrogens or endotoxins in the fluid path | Death; Cardiopulmonary Arrest | 5 | 1 | Device is non-pyrogenic | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling provides instructions to remove Tego from package using aseptic technique; Labeling includes instructions/precaution to disinfect Tego prior to each use; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Pyrogens or endotoxins in the fluid path | Paralysis; Organ failure/injury; Malignancy; Severe Hypotension | 4 | 1 | Device is non-pyrogenic | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling provides instructions to remove Tego from package using aseptic technique; Labeling includes instructions/precaution to disinfect Tego prior to each use; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Pyrogens or endotoxins in the fluid path | Mild to Moderate Hypotension; Moderate Inflammation (e.g. Phlebitis, Organ); Moderate Pain; Fever | 3 | 1 | Device is non-pyrogenic | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling provides instructions to remove Tego from package using aseptic technique; Labeling includes instructions/precaution to disinfect Tego prior to each use; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Pyrogens or endotoxins in the fluid path | Mild Inflammation (e.g. Phlebitis, Organ); Mild Pain | 2 | 1 | Device is non-pyrogenic | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling provides instructions to remove Tego from package using aseptic technique; Labeling includes instructions/precaution to disinfect Tego prior to each use; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Pyrogens or endotoxins in the fluid path | Temporary discomfort | 1 | 1 | Device is non-pyrogenic | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling provides instructions to remove Tego from package using aseptic technique; Labeling includes instructions/precaution to disinfect Tego prior to each use; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Loss of packaging integrity and caps off allows foreign material to enter into the sterile fluid path | Death; Cardiopulmonary Arrest | 5 | 1 | Device is designed to maintain packaging integrity and device sterility | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Loss of packaging integrity and caps off allows foreign material to enter into the sterile fluid path | Paralysis; Organ failure/injury; Malignancy; Severe Hypotension | 4 | 1 | Device is designed to maintain packaging integrity and device sterility | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Loss of packaging integrity and caps off allows foreign material to enter into the sterile fluid path | Mild to Moderate Hypotension; Moderate Inflammation (e.g. Phlebitis, Organ); Moderate Pain; Fever | 3 | 1 | Device is designed to maintain packaging integrity and device sterility | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Loss of packaging integrity and caps off allows foreign material to enter into the sterile fluid path | Mild Inflammation (e.g. Phlebitis, Organ); Mild Pain | 2 | 1 | Device is designed to maintain packaging integrity and device sterility | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Loss of packaging integrity and caps off allows foreign material to enter into the sterile fluid path | Temporary discomfort | 1 | 1 | Device is designed to maintain packaging integrity and device sterility | Labeling includes information that the device is sterile and non-pyrogenic fluid path in unopened and undamaged package; Labeling includes a caution to discard or do not use if device packaging is damaged or open; Labeling includes caution statement to use by expiry date on label; Labeling includes caution statement not to reuse or re-sterilize. Reuse negatively impacts performance / sterility potentially resulting in product failure / contamination.
Air Embolism | Obstruction within Device | Air infused into the body | Air ingress into syringe during use | Death | 5 | N/A | Device is designed with filter to equalize pressure | N/A
Air Embolism | Obstruction within Device | Air infused into the body | Air ingress into syringe during use | Pulmonary Embolus; Stroke | 4 | N/A | Device is designed with filter to equalize pressure | N/A
Air Embolism | Obstruction within Device | Air infused into the body | Air ingress into syringe during use | Respiratory Distress, Mild to Moderate | 3 | N/A | Device is designed with filter to equalize pressure | N/A
Air Embolism | Obstruction within Device | Air infused into the body | Air ingress into syringe during use | Temporary discomfort | 1 | N/A | Device is designed with filter to equalize pressure | N/A
Allergenic Substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Allergenic substances on packaging/device material | Contact dermatitis with localized allergic reaction | 2 | N/A | Device/packaging material are biocompatible, and non-pyrogenic | N/A
Allergenic Substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Allergenic substances on packaging/device material | Temporary discomfort | 1 | N/A | Device/packaging material are biocompatible, and non-pyrogenic | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Aseptic techniques not followed | Death; Cardiopulmonary Arrest | 5 | N/A | Device is provided with caps to avoid touch contamination | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Aseptic techniques not followed | Septic shock; Organ injury | 4 | N/A | Device is provided with caps to avoid touch contamination | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Aseptic techniques not followed | Bacteremia; Pain, Moderate | 3 | N/A | Device is provided with caps to avoid touch contamination | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Aseptic techniques not followed | Mild Phlebitis; Local reaction/infection; Mild Pain | 2 | N/A | Device is provided with caps to avoid touch contamination | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Aseptic techniques not followed | Temporary discomfort | 1 | N/A | Device is provided with caps to avoid touch contamination | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Fluid path sterility compromised due to broken, loose or unsealed fluid path | Death; Cardiopulmonary Arrest | 5 | N/A | Device is designed to withstand handling and exposure to different environments without compromising the fluid path | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Fluid path sterility compromised due to broken, loose or unsealed fluid path | Septic shock; Organ injury | 4 | N/A | Device is designed to withstand handling and exposure to different environments without compromising the fluid path | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Fluid path sterility compromised due to broken, loose or unsealed fluid path | Bacteremia; Moderate Pain | 3 | N/A | Device is designed to withstand handling and exposure to different environments without compromising the fluid path | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Fluid path sterility compromised due to broken, loose or unsealed fluid path | Mild Phlebitis; Local reaction/infection; Mild Pain | 2 | N/A | Device is designed to withstand handling and exposure to different environments without compromising the fluid path | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Fluid path sterility compromised due to broken, loose or unsealed fluid path | Temporary discomfort | 1 | N/A | Device is designed to withstand handling and exposure to different environments without compromising the fluid path | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Inadequate sterilization of the device | Death; Cardiopulmonary Arrest | 5 | N/A | Device is sterilized under validated cycle | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Inadequate sterilization of the device | Septic shock; Organ injury | 4 | N/A | Device is sterilized under validated cycle | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Inadequate sterilization of the device | Bacteremia; Moderate Pain | 3 | N/A | Device is sterilized under validated cycle | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Inadequate sterilization of the device | Mild Phlebitis; Local reaction/infection; Mild Pain | 2 | N/A | Device is sterilized under validated cycle | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via a parenteral route into the body | Inadequate sterilization of the device | Temporary discomfort | 1 | N/A | Device is sterilized under validated cycle | N/A
Biological contaminants | Device Contamination with Biological Material | Skin/Mucous membranes exposure to biological contaminants | Aseptic techniques not followed | Death; Cardiopulmonary Arrest | 5 | N/A | Device is provided with caps to avoid touch contamination | N/A
Biological contaminants | Device Contamination with Biological Material | Skin/Mucous membranes exposure to biological contaminants | Aseptic techniques not followed | Septic shock; Organ injury | 4 | N/A | Device is provided with caps to avoid touch contamination | N/A
Biological contaminants | Device Contamination with Biological Material | Skin/Mucous membranes exposure to biological contaminants | Aseptic techniques not followed | Bacteremia; Moderate Cellulitis; Organ Inflammation (reversible) | 3 | N/A | Device is provided with caps to avoid touch contamination | N/A
Biological contaminants | Device Contamination with Biological Material | Skin/Mucous membranes exposure to biological contaminants | Aseptic techniques not followed | Mild Cellulitis | 2 | N/A | Device is provided with caps to avoid touch contamination | N/A
Biological contaminants | Device Contamination with Biological Material | Skin/Mucous membranes exposure to biological contaminants | Aseptic techniques not followed | Localized skin swelling, redness | 2 | N/A | Device is provided with caps to avoid touch contamination | N/A
Biological contaminants | Device Contamination with Biological Material | Skin/Mucous membranes exposure to biological contaminants | Aseptic techniques not followed | Temporary discomfort | 1 | N/A | Device is provided with caps to avoid touch contamination | N/A
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Fluid leak due to defective device (Eg Crack, holes, cuts) | Severe Burn; Sepsis; Severe Pain | 4 | N/A | Device is designed to uphold integrity | N/A
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Fluid leak due to defective device (Eg Crack, holes, cuts) | Moderate Burn; Moderate Pain | 3 | N/A | Device is designed to uphold integrity | N/A
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Fluid leak due to defective device (Eg Crack, holes, cuts) | Mild Burn; Mild Pain | 2 | N/A | Device is designed to uphold integrity | N/A
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Fluid leak due to defective device (Eg Crack, holes, cuts) | Temporary discomfort | 1 | N/A | Device is designed to uphold integrity | N/A
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Fluid leak due to improper connection/disconnection/use of the vial adapter | Severe Burn; Sepsis; Severe Pain | 4 | N/A | Device ports comply to ISO 80369-7 and other propriety connector product requirements | N/A
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Fluid leak due to improper connection/disconnection/use of the vial adapter | Moderate Burn; Moderate Pain | 3 | N/A | Device ports comply to ISO 80369-7 and other propriety connector product requirements | N/A
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Fluid leak due to improper connection/disconnection/use of the vial adapter | Mild Burn; Mild Pain | 2 | N/A | Device ports comply to ISO 80369-7 and other propriety connector product requirements | N/A
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Fluid leak due to improper connection/disconnection/use of the vial adapter | Temporary discomfort | 1 | N/A | Device ports comply to ISO 80369-7 and other propriety connector product requirements | N/A
Chemical-Other toxin | Biocompatibility Issue | Intravenous exposure to non-hemocompatible materials | Non-hemocompatible material in the fluid path | Severe Hemolytic Anemia; Severe Hyperkalemia; Hypotension; Congestive Heart Failure; Acute Kidney Injury | 4 | N/A | Device is designed with hemocompatible materials | N/A
Chemical-Other toxin | Biocompatibility Issue | Intravenous exposure to non-hemocompatible materials | Non-hemocompatible material in the fluid path | Mild to Moderate Hemolytic Anemia; Mild to Moderate Hyperkalemia; Fever | 3 | N/A | Device is designed with hemocompatible materials | N/A
Chemical-Other toxin | Biocompatibility Issue | Intravenous exposure to non-hemocompatible materials | Non-hemocompatible material in the fluid path | Temporary discomfort | 1 | N/A | Device is designed with hemocompatible materials | N/A
DEHP/BPA leached | Biocompatibility Issue | Intravenous exposure to DEHP or BPA | DEHP from the device materials leached into the fluid path | Organ damage | 4 | N/A | Device is designed to use non-DEHP materials | N/A
DEHP/BPA leached | Biocompatibility Issue | Intravenous exposure to DEHP or BPA | DEHP from the device materials leached into the fluid path | Reproductive system dysfunction | 3 | N/A | Device is designed to use non-DEHP materials | N/A
DEHP/BPA leached | Biocompatibility Issue | Intravenous exposure to DEHP or BPA | DEHP from the device materials leached into the fluid path | Temporary discomfort | 1 | N/A | Device is designed to use non-DEHP materials | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion solution | Device appears unsafe to use and needs to be replaced (Eg. Open package, Missing / Loose / Damaged components, Discoloration, Visible particulate etc.) | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | N/A | Product is designed to not be cracked, broken, or severely discolored and maintain integrity after sterilization | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion solution | Device appears unsafe to use and needs to be replaced (Eg. Open package, Missing / Loose / Damaged components, Discoloration, Visible particulate etc.) | Mild to Moderate Pain | 2 | N/A | Product is designed to not be cracked, broken, or severely discolored and maintain integrity after sterilization | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion solution | Device appears unsafe to use and needs to be replaced (Eg. Open package, Missing / Loose / Damaged components, Discoloration, Visible particulate etc.) | Temporary discomfort | 1 | N/A | Product is designed to not be cracked, broken, or severely discolored and maintain integrity after sterilization | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion solution | Problem identifying / selecting the device | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | N/A | Device is labeled with pertinent information | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion solution | Problem identifying / selecting the device | Mild to Moderate Pain | 2 | N/A | Device is labeled with pertinent information | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion solution | Problem identifying / selecting the device | Temporary discomfort | 1 | N/A | Device is labeled with pertinent information | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion solution | Incomplete fluid volume transfer | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | N/A | Device is designed to allow full fluid administration control | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion solution | Incomplete fluid volume transfer | Mild to Moderate Pain | 2 | N/A | Device is designed to allow full fluid administration control | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion solution | Incomplete fluid volume transfer | Temporary discomfort | 1 | N/A | Device is designed to allow full fluid administration control | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion solution | Fluid path occlusion / Filter Occlusion | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | N/A | Device is designed to allow a controlled fluid transfer | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion solution | Fluid path occlusion / Filter Occlusion | Mild to Moderate Pain | 2 | N/A | Device is designed to allow a controlled fluid transfer | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of infusion solution | Fluid path occlusion / Filter Occlusion | Temporary discomfort | 1 | N/A | Device is designed to allow a controlled fluid transfer | N/A
Ergonomic stress | Human-Device Interface Issue | Exposure to ergonomic stress | Difficult / Unable to connect / disconnect the device with other mating devices | Temporary discomfort | 1 | N/A | Device is designed to allow easy connection/disconnection with mating device | N/A
Ergonomic stress | Human-Device Interface Issue | Exposure to ergonomic stress | Unable to remove Cap / sheath from the device | Temporary discomfort | 1 | N/A | Cap is designed to be easily removable with one hand | N/A
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/fumes | Bad connection /disconnection between device and other mating devices | Mild to Moderate Respiratory distress; Moderate toxic effects such as seizures, abdominal pain | 3 | N/A | Device is design to maintain connection with female luer port | N/A
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/fumes | Bad connection /disconnection between device and other mating devices | Mild Pain; Mild toxic effects such as fatigue, sleepiness, headaches, eye irritation, skin rashes | 2 | N/A | Device is design to maintain connection with female luer port | N/A
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/fumes | Bad connection /disconnection between device and other mating devices | Temporary Discomfort | 1 | N/A | Device is design to maintain connection with female luer port | N/A
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/fumes | Fluid leak due to defective device (Eg Crack, holes, cuts) | Mild to Moderate Respiratory distress; Moderate toxic effects such as seizures, abdominal pain | 3 | N/A | Device is designed to uphold integrity | N/A
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/fumes | Fluid leak due to defective device (Eg Crack, holes, cuts) | Mild Pain; Mild toxic effects such as fatigue, sleepiness, headaches, eye irritation, skin rashes | 2 | N/A | Device is designed to uphold integrity | N/A
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/fumes | Fluid leak due to defective device (Eg Crack, holes, cuts) | Temporary Discomfort | 1 | N/A | Device is designed to uphold integrity | N/A
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/fumes | Vial breakage during insertion | Mild to Moderate Respiratory distress; Moderate toxic effects such as seizures, abdominal pain | 3 | N/A | Device is designed to allow easy insertion into vial | N/A
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/fumes | Vial breakage during insertion | Mild Pain; Mild toxic effects such as fatigue, sleepiness, headaches, eye irritation, skin rashes | 2 | N/A | Device is designed to allow easy insertion into vial | N/A
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/fumes | Vial breakage during insertion | Temporary Discomfort | 1 | N/A | Device is designed to allow easy insertion into vial | N/A
particulate presented | Particulates | Embolization into the vascular system | Particulate in the fluid path | Death; Cardiopulmonary Arrest | 5 | N/A | Device is designed to minimize particulate generation | N/A
particulate presented | Particulates | Embolization into the vascular system | Particulate in the fluid path | Pulmonary embolism; Myocardial infarction; Stroke | 4 | N/A | Device is designed to minimize particulate generation | N/A
particulate presented | Particulates | Embolization into the vascular system | Particulate in the fluid path | Extremity swelling; Moderate Pain; Paresthesias | 3 | N/A | Device is designed to minimize particulate generation | N/A
particulate presented | Particulates | Embolization into the vascular system | Particulate in the fluid path | Thrombophlebitis; Mild Pain | 2 | N/A | Device is designed to minimize particulate generation | N/A
particulate presented | Particulates | Embolization into the vascular system | Particulate in the fluid path | Temporary Discomfort | 1 | N/A | Device is designed to minimize particulate generation | N/A
particulate presented | Particulates | Embolization into the vascular system | Loss of packaging integrity and caps falling off allows foreign material to enter into the sterile fluid path | Death; Cardiopulmonary Arrest | 5 | N/A | Device is packaged as a closed fluid path | N/A
particulate presented | Particulates | Embolization into the vascular system | Loss of packaging integrity and caps falling off allows foreign material to enter into the sterile fluid path | Pulmonary embolism; Myocardial infarction; Stroke | 4 | N/A | Device is packaged as a closed fluid path | N/A
particulate presented | Particulates | Embolization into the vascular system | Loss of packaging integrity and caps falling off allows foreign material to enter into the sterile fluid path | Extremity swelling; Moderate Pain; Paresthesias | 3 | N/A | Device is packaged as a closed fluid path | N/A
particulate presented | Particulates | Embolization into the vascular system | Loss of packaging integrity and caps falling off allows foreign material to enter into the sterile fluid path | Thrombophlebitis; Mild Pain | 2 | N/A | Device is packaged as a closed fluid path | N/A
particulate presented | Particulates | Embolization into the vascular system | Loss of packaging integrity and caps falling off allows foreign material to enter into the sterile fluid path | Temporary Discomfort | 1 | N/A | Device is packaged as a closed fluid path | N/A
Sharp edge | Material Puncture | Exposure to sharp edge/pinch point/abrasive surface | Device contains sharp corners or edges | Skin Abrasion; Bruising; Mild Pain; Skin integrity compromised | 2 | N/A | Device is provided with protective cap to minimize exposure to sharp edges | N/A
Sharp edge | Material Puncture | Exposure to sharp edge/pinch point/abrasive surface | Device contains sharp corners or edges | Temporary discomfort | 1 | N/A | Device is provided with protective cap to minimize exposure to sharp edges | N/A
Sharp edge | Material Puncture | Exposure to sharp edge/pinch point/abrasive surface | Proper disposal procedure not followed | Skin Abrasion; Bruising; Mild Pain; Skin integrity compromised | 2 | N/A | N/A | N/A
Sharp edge | Material Puncture | Exposure to sharp edge/pinch point/abrasive surface | Proper disposal procedure not followed | Temporary discomfort | 1 | N/A | N/A | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Fluid path contains toxins (E.g. systemic, mutagens, carcinogens) | Death; Cardiopulmonary Arrest | 5 | N/A | Device is designed to be biocompatible, non-pyrogenic | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Fluid path contains toxins (E.g. systemic, mutagens, carcinogens) | Paralysis; Organ failure/injury; Malignancy; Severe Hypotension | 4 | N/A | Device is designed to be biocompatible, non-pyrogenic | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Fluid path contains toxins (E.g. systemic, mutagens, carcinogens) | Mild to Moderate Hypotension; Moderate Inflammation (e.g. Phlebitis, Organ); Moderate Pain; Fever | 3 | N/A | Device is designed to be biocompatible, non-pyrogenic | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Fluid path contains toxins (E.g. systemic, mutagens, carcinogens) | Mild Inflammation (e.g. Phlebitis, Organ); Mild Pain | 2 | N/A | Device is designed to be biocompatible, non-pyrogenic | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Fluid path contains toxins (E.g. systemic, mutagens, carcinogens) | Temporary discomfort | 1 | N/A | Device is designed to be biocompatible, non-pyrogenic | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Fluid path contains toxins (E.g. systemic, mutagens, carcinogens) | Temporary discomfort | 2 | N/A | Device is designed to be biocompatible, non-pyrogenic | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Pyrogens or endotoxins in the fluid path | Death; Cardiopulmonary Arrest | 5 | N/A | Device is designed to be biocompatible, non-pyrogenic | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Pyrogens or endotoxins in the fluid path | Paralysis; Organ failure/injury; Malignancy; Severe Hypotension | 4 | N/A | Device is designed to be biocompatible, non-pyrogenic | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Pyrogens or endotoxins in the fluid path | Mild to Moderate Hypotension; Moderate Inflammation (e.g. Phlebitis, Organ); Moderate Pain; Fever | 3 | N/A | Device is designed to be biocompatible, non-pyrogenic | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Pyrogens or endotoxins in the fluid path | Mild Inflammation (e.g. Phlebitis, Organ); Mild Pain | 2 | N/A | Device is designed to be biocompatible, non-pyrogenic | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Pyrogens or endotoxins in the fluid path | Temporary discomfort | 1 | N/A | Device is designed to be biocompatible, non-pyrogenic | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Loss of packaging integrity and caps falling off allows foreign material to enter into the sterile fluid path | Death; Cardiopulmonary Arrest | 5 | N/A | The seal of the packaging is designed to maintain its integrity after sterilization | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Loss of packaging integrity and caps falling off allows foreign material to enter into the sterile fluid path | Paralysis; Organ failure/injury; Malignancy; Severe Hypotension | 4 | N/A | The seal of the packaging is designed to maintain its integrity after sterilization | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Loss of packaging integrity and caps falling off allows foreign material to enter into the sterile fluid path | Mild to Moderate Hypotension; Moderate Inflammation (e.g. Phlebitis, Organ); Moderate Pain; Fever | 3 | N/A | The seal of the packaging is designed to maintain its integrity after sterilization | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Loss of packaging integrity and caps falling off allows foreign material to enter into the sterile fluid path | Mild Inflammation (e.g. Phlebitis, Organ); Mild Pain | 2 | N/A | The seal of the packaging is designed to maintain its integrity after sterilization | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Exposure to toxins | Loss of packaging integrity and caps falling off allows foreign material to enter into the sterile fluid path | Temporary discomfort | 1 | N/A | The seal of the packaging is designed to maintain its integrity after sterilization | N/A
Air in line | Improper Flow or Infusion | Air infused into the body | Air in line due to broken, loose, unsealed fluid path | Death | 5 | 1 | Device designed to function after pressure infusion requirements; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Device should meet assembly integrity test requirements; Components must not be cracked, broken, distorted, or severely discolored; Luer connectors are designed to be compliant with ISO standards for connection / disconnection; Device designed to withstand internal back pressure requirements | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm); Labeling includes instructions for using the device; Labeling provide instructions not to use needles, blunt cannulas or luer caps on luer activated injection access sites (Eg., Needlefree connectors)
Air in line | Improper Flow or Infusion | Air infused into the body | Air in line due to broken, loose, unsealed fluid path | Pulmonary Embolus Stroke | 4 | 1 | Device designed to function after pressure infusion requirements; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Device should meet assembly integrity test requirements; Components must not be cracked, broken, distorted, or severely discolored; Luer connectors are designed to be compliant with ISO standards for connection / disconnection; Device designed to withstand internal back pressure requirements | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm); Labeling includes instructions for using the device; Labeling provide instructions not to use needles, blunt cannulas or luer caps on luer activated injection access sites (Eg., Needlefree connectors)
Air in line | Improper Flow or Infusion | Air infused into the body | Air in line due to broken, loose, unsealed fluid path | Respiratory Distress, Mild to Moderate | 3 | 2 | Device designed to function after pressure infusion requirements; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Device should meet assembly integrity test requirements; Components must not be cracked, broken, distorted, or severely discolored; Luer connectors are designed to be compliant with ISO standards for connection / disconnection; Device designed to withstand internal back pressure requirements | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm); Labeling includes instructions for using the device; Labeling provide instructions not to use needles, blunt cannulas or luer caps on luer activated injection access sites (Eg., Needlefree connectors)
Air in line | Improper Flow or Infusion | Air infused into the body | Air in line due to broken, loose, unsealed fluid path | Temporary discomfort | 1 | 1 | Device designed to function after pressure infusion requirements; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Device should meet assembly integrity test requirements; Components must not be cracked, broken, distorted, or severely discolored; Luer connectors are designed to be compliant with ISO standards for connection / disconnection; Device designed to withstand internal back pressure requirements | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm); Labeling includes instructions for using the device; Labeling provide instructions not to use needles, blunt cannulas or luer caps on luer activated injection access sites (Eg., Needlefree connectors)
Air in line | Improper Flow or Infusion | Air infused into the body | Air in line due to improper priming of device | Death | 5 | 1 | Components designed to turn easily; Device designed easily attach/detach to connecting components; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Labeling provides instructions for priming a device (Eg., Air eliminating filter, Y site, Back check valve, Burette float valve, Manual blood pump, Hand pump, Clave/Microclave/Nano clave/Neutron, Flow regulators)
Air in line | Improper Flow or Infusion | Air infused into the body | Air in line due to improper priming of device | Pulmonary Embolus Stroke | 4 | 1 | Components designed to turn easily; Device designed easily attach/detach to connecting components; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Labeling provides instructions for priming a device (Eg., Air eliminating filter, Y site, Back check valve, Burette float valve, Manual blood pump, Hand pump, Clave/Microclave/Nano clave/Neutron, Flow regulators)
Air in line | Improper Flow or Infusion | Air infused into the body | Air in line due to improper priming of device | Respiratory Distress, Mild to Moderate | 3 | 2 | Components designed to turn easily; Device designed easily attach/detach to connecting components; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Labeling provides instructions for priming a device (Eg., Air eliminating filter, Y site, Back check valve, Burette float valve, Manual blood pump, Hand pump, Clave/Microclave/Nano clave/Neutron, Flow regulators)
Air in line | Improper Flow or Infusion | Air infused into the body | Air in line due to improper priming of device | Temporary discomfort | 1 | 1 | Components designed to turn easily; Device designed easily attach/detach to connecting components; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Labeling provides instructions for priming a device (Eg., Air eliminating filter, Y site, Back check valve, Burette float valve, Manual blood pump, Hand pump, Clave/Microclave/Nano clave/Neutron, Flow regulators)
Allergenic substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Contact dermatitis (skin inflammation) with systemic allergic reaction Eye inflammation | 3 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Allergenic substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Contact dermatitis with localized allergic reaction | 2 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Allergenic substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Temporary discomfort | 1 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Allergenic substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Contact dermatitis (skin inflammation) with systemic allergic reaction Eye inflammation | 3 | 2 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Allergenic substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Contact dermatitis with localized allergic reaction | 2 | 2 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Allergenic substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Temporary discomfort | 1 | 2 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Allergenic substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Contact dermatitis (skin inflammation) with systemic allergic reaction Eye inflammation | 3 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Allergenic substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Contact dermatitis with localized allergic reaction | 2 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Allergenic substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Temporary discomfort | 1 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Allergenic substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Patient movement causes set/device to disconnect at a connection point | Contact dermatitis (skin inflammation) with systemic allergic reaction Eye inflammation | 3 | 2 | Device designed easily attach/detach to connecting components; Device with rotating collars must meet rotating collar luer retention force per ISO 80369-7 requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | N/A
Allergenic substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Patient movement causes set/device to disconnect at a connection point | Contact dermatitis with localized allergic reaction | 2 | 2 | Device designed easily attach/detach to connecting components; Device with rotating collars must meet rotating collar luer retention force per ISO 80369-7 requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | N/A
Allergenic substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Patient movement causes set/device to disconnect at a connection point | Temporary discomfort | 1 | 2 | Device designed easily attach/detach to connecting components; Device with rotating collars must meet rotating collar luer retention force per ISO 80369-7 requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | N/A
Allergenic substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Device materials are not biocompatible | Contact dermatitis (skin inflammation) with systemic allergic reaction Eye inflammation | 3 | 2 | Devices are biocompatible (Device material are complaint to ISO 10993); Device must meet environmental requirements | N/A
Allergenic substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Device materials are not biocompatible | Contact dermatitis with localized allergic reaction | 2 | 2 | Devices are biocompatible (Device material are complaint to ISO 10993); Device must meet environmental requirements | N/A
Allergenic substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Device materials are not biocompatible | Temporary discomfort | 1 | 2 | Devices are biocompatible (Device material are complaint to ISO 10993); Device must meet environmental requirements | N/A
Allergenic substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Silicone sleeve stuck down after disconnecting mating component | Contact dermatitis (skin inflammation) with systemic allergic reaction Eye inflammation | 3 | 2 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
Allergenic substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Silicone sleeve stuck down after disconnecting mating component | Contact dermatitis with localized allergic reaction | 2 | 2 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
Allergenic substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Silicone sleeve stuck down after disconnecting mating component | Temporary discomfort | 1 | 2 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
Allergenic substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Tubing rupture due to high pressure | Contact dermatitis (skin inflammation) with systemic allergic reaction Eye inflammation | 3 | 2 | Device designed to function after pressure infusion requirements | Neutron/Neutron Ext Set package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 350psig or 10mL per second of contrast media; Administration set primary package labeling or instructions for use for extension sets or administration sets with distal injection access sites, excluding sets intended to be used with high pressure infusion equipment shall include a caution statement that set is not for high pressure infusion; Administration set primary package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 400 psig; Administration set with a MicroClave primary package labeling or instructions for use for high pressure extension sets contain a warning statement that high pressure infusion is conducted through the MicroClave in the middle of the assembly and the upstream clamp is engaged to isolate other upstream components from high pressures
Allergenic substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Tubing rupture due to high pressure | Contact dermatitis with localized allergic reaction | 2 | 2 | Device designed to function after pressure infusion requirements | Neutron/Neutron Ext Set package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 350psig or 10mL per second of contrast media; Administration set primary package labeling or instructions for use for extension sets or administration sets with distal injection access sites, excluding sets intended to be used with high pressure infusion equipment shall include a caution statement that set is not for high pressure infusion; Administration set primary package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 400 psig; Administration set with a MicroClave primary package labeling or instructions for use for high pressure extension sets contain a warning statement that high pressure infusion is conducted through the MicroClave in the middle of the assembly and the upstream clamp is engaged to isolate other upstream components from high pressures
Allergenic substances | Biocompatibility Issue | Skin/mucous membranes exposure to allergenic substances | Tubing rupture due to high pressure | Temporary discomfort | 1 | 2 | Device designed to function after pressure infusion requirements | Neutron/Neutron Ext Set package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 350psig or 10mL per second of contrast media; Administration set primary package labeling or instructions for use for extension sets or administration sets with distal injection access sites, excluding sets intended to be used with high pressure infusion equipment shall include a caution statement that set is not for high pressure infusion; Administration set primary package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 400 psig; Administration set with a MicroClave primary package labeling or instructions for use for high pressure extension sets contain a warning statement that high pressure infusion is conducted through the MicroClave in the middle of the assembly and the upstream clamp is engaged to isolate other upstream components from high pressures
Allergenic Substances | Biocompatibility Issue | Intravenous exposure to allergenic substances | Allergenic substances in the fluid path | Anaphylaxis | 5 | 2 | Devices are biocompatible (Device material are complaint to ISO 10993); Components must not be cracked, broken, distorted, or severely discolored; Devices should meet the storage and transport requirements; Package must be free of holes or openings that would allow contamination to enter; Package must be free of holes or openings that would allow contamination to enter; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Device must meet environmental requirements; Packaging shall maintain a sterile barrier until expiry date | N/A
Allergenic Substances | Biocompatibility Issue | Intravenous exposure to allergenic substances | Allergenic substances in the fluid path | Respiratory distress, Severe Stevens Johnson Syndrome | 4 | 2 | Devices are biocompatible (Device material are complaint to ISO 10993); Components must not be cracked, broken, distorted, or severely discolored; Devices should meet the storage and transport requirements; Package must be free of holes or openings that would allow contamination to enter; Package must be free of holes or openings that would allow contamination to enter; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Device must meet environmental requirements; Packaging shall maintain a sterile barrier until expiry date | N/A
Allergenic Substances | Biocompatibility Issue | Intravenous exposure to allergenic substances | Allergenic substances in the fluid path | Local allergic reaction/Dermatitis Organ inflammation | 3 | 2 | Devices are biocompatible (Device material are complaint to ISO 10993); Components must not be cracked, broken, distorted, or severely discolored; Devices should meet the storage and transport requirements; Package must be free of holes or openings that would allow contamination to enter; Package must be free of holes or openings that would allow contamination to enter; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Device must meet environmental requirements; Packaging shall maintain a sterile barrier until expiry date | N/A
Allergenic Substances | Biocompatibility Issue | Intravenous exposure to allergenic substances | Allergenic substances in the fluid path | Urticaria Pain, Mild | 2 | 2 | Devices are biocompatible (Device material are complaint to ISO 10993); Components must not be cracked, broken, distorted, or severely discolored; Devices should meet the storage and transport requirements; Package must be free of holes or openings that would allow contamination to enter; Package must be free of holes or openings that would allow contamination to enter; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Device must meet environmental requirements; Packaging shall maintain a sterile barrier until expiry date | N/A
Allergenic Substances | Biocompatibility Issue | Intravenous exposure to allergenic substances | Allergenic substances in the fluid path | Temporary discomfort | 1 | 2 | Devices are biocompatible (Device material are complaint to ISO 10993); Components must not be cracked, broken, distorted, or severely discolored; Devices should meet the storage and transport requirements; Package must be free of holes or openings that would allow contamination to enter; Package must be free of holes or openings that would allow contamination to enter; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Device must meet environmental requirements; Packaging shall maintain a sterile barrier until expiry date | N/A
Allergenic Substances | Biocompatibility Issue | Intravenous exposure to allergenic substances | Loss of packaging integrity and caps off allows foreign material to enter into the sterile fluid path | Anaphylaxis | 5 | 2 | Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Devices should meet the storage and transport requirements; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Seal strength shall be sufficient to prevent the package from opening prior to point of use; Protective caps should be secure but easily removable | N/A
Allergenic Substances | Biocompatibility Issue | Intravenous exposure to allergenic substances | Loss of packaging integrity and caps off allows foreign material to enter into the sterile fluid path | Severe Respiratory distress; Stevens Johnson Syndrome | 4 | 2 | Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Devices should meet the storage and transport requirements; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Seal strength shall be sufficient to prevent the package from opening prior to point of use; Protective caps should be secure but easily removable | N/A
Allergenic Substances | Biocompatibility Issue | Intravenous exposure to allergenic substances | Loss of packaging integrity and caps off allows foreign material to enter into the sterile fluid path | Local allergic reaction/Dermatitis; Organ inflammation | 3 | 2 | Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Devices should meet the storage and transport requirements; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Seal strength shall be sufficient to prevent the package from opening prior to point of use; Protective caps should be secure but easily removable | N/A
Allergenic Substances | Biocompatibility Issue | Intravenous exposure to allergenic substances | Loss of packaging integrity and caps off allows foreign material to enter into the sterile fluid path | Urticaria; Mild Pain | 2 | 2 | Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Devices should meet the storage and transport requirements; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Seal strength shall be sufficient to prevent the package from opening prior to point of use; Protective caps should be secure but easily removable | N/A
Allergenic Substances | Biocompatibility Issue | Intravenous exposure to allergenic substances | Loss of packaging integrity and caps off allows foreign material to enter into the sterile fluid path | Temporary discomfort | 1 | 2 | Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Devices should meet the storage and transport requirements; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Seal strength shall be sufficient to prevent the package from opening prior to point of use; Protective caps should be secure but easily removable | N/A
Allergenic Substances | Biocompatibility Issue | Intravenous exposure to allergenic substances | Device materials are incompatible to prolonged exposure to disinfectant caps | Anaphylaxis | 5 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
Allergenic Substances | Biocompatibility Issue | Intravenous exposure to allergenic substances | Device materials are incompatible to prolonged exposure to disinfectant caps | Severe Respiratory distress; Stevens Johnson Syndrome | 4 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
Allergenic Substances | Biocompatibility Issue | Intravenous exposure to allergenic substances | Device materials are incompatible to prolonged exposure to disinfectant caps | Local allergic reaction/Dermatitis; Organ inflammation | 3 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
Allergenic Substances | Biocompatibility Issue | Intravenous exposure to allergenic substances | Device materials are incompatible to prolonged exposure to disinfectant caps | Urticaria; Mild Pain | 2 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
Allergenic Substances | Biocompatibility Issue | Intravenous exposure to allergenic substances | Device materials are incompatible to prolonged exposure to disinfectant caps | Temporary discomfort | 1 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
Allergenic Substances | Biocompatibility Issue | Intravenous exposure to allergenic substances | Device materials are not biocompatible | Anaphylaxis | 5 | 2 | Device must meet environmental requirements; Devices are biocompatible (Device material are complaint to ISO 10993) | N/A
Allergenic Substances | Biocompatibility Issue | Intravenous exposure to allergenic substances | Device materials are not biocompatible | Severe Respiratory distress; Stevens Johnson Syndrome | 4 | 2 | Device must meet environmental requirements; Devices are biocompatible (Device material are complaint to ISO 10993) | N/A
Allergenic Substances | Biocompatibility Issue | Intravenous exposure to allergenic substances | Device materials are not biocompatible | Local allergic reaction/Dermatitis; Organ inflammation | 3 | 2 | Device must meet environmental requirements; Devices are biocompatible (Device material are complaint to ISO 10993) | N/A
Allergenic Substances | Biocompatibility Issue | Intravenous exposure to allergenic substances | Device materials are not biocompatible | Urticaria; Mild Pain | 2 | 2 | Device must meet environmental requirements; Devices are biocompatible (Device material are complaint to ISO 10993) | N/A
Allergenic Substances | Biocompatibility Issue | Intravenous exposure to allergenic substances | Device materials are not biocompatible | Temporary discomfort | 1 | 2 | Device must meet environmental requirements; Devices are biocompatible (Device material are complaint to ISO 10993) | N/A
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Severe Burn; Sepsis; Severe Pain | 4 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Moderate Burn Moderate Pain | 3 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Mild Burn; Mild Pain | 2 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Temporary discomfort | 1 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Severe Burn; Sepsis; Severe Pain | 4 | 2 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Moderate Burn Moderate Pain | 3 | 2 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Mild Burn; Mild Pain | 2 | 2 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Temporary discomfort | 1 | 2 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Severe Burn; Sepsis; Severe Pain | 4 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Moderate Burn Moderate Pain | 3 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Mild Burn; Mild Pain | 2 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Temporary discomfort | 1 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Patient movement causes set/device to disconnect at a connection point | Severe Burn; Sepsis; Severe Pain | 4 | 2 | Device designed easily attach/detach to connecting components; Device with rotating collars must meet rotating collar luer retention force per ISO 80369-7 requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | N/A
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Patient movement causes set/device to disconnect at a connection point | Moderate Burn Moderate Pain | 3 | 2 | Device designed easily attach/detach to connecting components; Device with rotating collars must meet rotating collar luer retention force per ISO 80369-7 requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | N/A
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Patient movement causes set/device to disconnect at a connection point | Mild Burn; Mild Pain | 2 | 2 | Device designed easily attach/detach to connecting components; Device with rotating collars must meet rotating collar luer retention force per ISO 80369-7 requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | N/A
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Patient movement causes set/device to disconnect at a connection point | Temporary discomfort | 1 | 2 | Device designed easily attach/detach to connecting components; Device with rotating collars must meet rotating collar luer retention force per ISO 80369-7 requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | N/A
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Silicone sleeve stuck down after disconnecting mating component | Severe Burn; Sepsis; Severe Pain | 4 | 2 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Silicone sleeve stuck down after disconnecting mating component | Moderate Burn Moderate Pain | 3 | 2 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Silicone sleeve stuck down after disconnecting mating component | Mild Burn; Mild Pain | 2 | 2 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
Caustic Substances | Material Erosion | Skin/mucous membranes exposure to caustic substances | Silicone sleeve stuck down after disconnecting mating component | Temporary discomfort | 1 | 2 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
Chemical-Other toxin | Biocompatibility Issue | Intravenous exposure to non-hemocompatible materials | Non-hemocompatible material in the fluid path | Death Cardiopulmonary Arrest | 5 | 1 | Device fluid path is sized accordingly for the administration and aspiration of blood without significant hemolysis; Device must meet the chemical compatibility requirements; Devices are biocompatible (Device material are complaint to ISO 10993) | N/A
Chemical-Other toxin | Biocompatibility Issue | Intravenous exposure to non-hemocompatible materials | Non-hemocompatible material in the fluid path | Hemolytic Anemia, Severe Hyperkalemia, Severe Hypotension Congestive Heart Failure Acute Kidney Injury | 4 | 1 | Device fluid path is sized accordingly for the administration and aspiration of blood without significant hemolysis; Device must meet the chemical compatibility requirements; Devices are biocompatible (Device material are complaint to ISO 10993) | N/A
Chemical-Other toxin | Biocompatibility Issue | Intravenous exposure to non-hemocompatible materials | Non-hemocompatible material in the fluid path | Hemolytic Anemia, Mild to Moderate Hyperkalemia, Mild to Moderate Fever | 3 | 1 | Device fluid path is sized accordingly for the administration and aspiration of blood without significant hemolysis; Device must meet the chemical compatibility requirements; Devices are biocompatible (Device material are complaint to ISO 10993) | N/A
Chemical-Other toxin | Biocompatibility Issue | Intravenous exposure to non-hemocompatible materials | Non-hemocompatible material in the fluid path | Temporary discomfort | 1 | 1 | Device fluid path is sized accordingly for the administration and aspiration of blood without significant hemolysis; Device must meet the chemical compatibility requirements; Devices are biocompatible (Device material are complaint to ISO 10993) | N/A
DEHP/BPA leached | Biocompatibility Issue | Intravenous exposure to DEHP/BPA | DEHP from the set materials leached into the fluid path | Organ damage | 4 | 1 | Administration sets specified as a Non-DEHP plasticizer PVC finished devices shall contain no more than 0.1% DEHP materials by mass(weight). Any finished set using DEHP plasticized PVC compounds in any tubing component will not qualify for Non-DEHP; Materials of construction shall comply to requirements per IS.DC.304 | Labeling provides information about the set containing DEHP labeled as containing DEHP
DEHP/BPA leached | Biocompatibility Issue | Intravenous exposure to DEHP/BPA | DEHP from the set materials leached into the fluid path | Reproductive system dysfunction | 3 | 1 | Administration sets specified as a Non-DEHP plasticizer PVC finished devices shall contain no more than 0.1% DEHP materials by mass(weight). Any finished set using DEHP plasticized PVC compounds in any tubing component will not qualify for Non-DEHP; Materials of construction shall comply to requirements per IS.DC.305 | Labeling provides information about the set containing DEHP labeled as containing DEHP
DEHP/BPA leached | Biocompatibility Issue | Intravenous exposure to DEHP/BPA | DEHP from the set materials leached into the fluid path | Temporary discomfort | 1 | 1 | Administration sets specified as a Non-DEHP plasticizer PVC finished devices shall contain no more than 0.1% DEHP materials by mass(weight). Any finished set using DEHP plasticized PVC compounds in any tubing component will not qualify for Non-DEHP; Materials of construction shall comply to requirements per IS.DC.306 | Labeling provides information about the set containing DEHP labeled as containing DEHP
DEHP/BPA leached | Biocompatibility Issue | Intravenous exposure to DEHP/BPA | BPA in the device materials leached into the fluid path | Organ damage | 4 | 1 | Materials of construction shall comply to requirements per IS.DC.304 | Labeling provides information about the set/device containing PVC labeled as contain/not contain PVC (Eg., Microclave/Nanoclave contains polycarbonate)
DEHP/BPA leached | Biocompatibility Issue | Intravenous exposure to DEHP/BPA | BPA in the device materials leached into the fluid path | Reproductive system dysfunction | 3 | 1 | Materials of construction shall comply to requirements per IS.DC.305 | Labeling provides information about the set/device containing PVC labeled as contain/not contain PVC (Eg., Microclave/Nanoclave contains polycarbonate)
DEHP/BPA leached | Biocompatibility Issue | Intravenous exposure to DEHP/BPA | BPA in the device materials leached into the fluid path | Temporary discomfort | 1 | 1 | Materials of construction shall comply to requirements per IS.DC.306 | Labeling provides information about the set/device containing PVC labeled as contain/not contain PVC (Eg., Microclave/Nanoclave contains polycarbonate)
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Intravenous exposure to toxins | Fluid path contains toxins (E.g. systemic, mutagens, carcinogens) or Pyrogens or endotoxins in the fluid path | Death; Cardiopulmonary Arrest | 5 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device must meet environmental requirements; Devices are biocompatible (Device material are complaint to ISO 10993); Packaging shall maintain a sterile barrier until expiry date | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Intravenous exposure to toxins | Fluid path contains toxins (E.g. systemic, mutagens, carcinogens) or Pyrogens or endotoxins in the fluid path | Paralysis; Organ failure/injury; Malignancy; Severe Hypertension; Severe Hypotension | 4 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device must meet environmental requirements; Devices are biocompatible (Device material are complaint to ISO 10993); Packaging shall maintain a sterile barrier until expiry date | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Intravenous exposure to toxins | Fluid path contains toxins (E.g. systemic, mutagens, carcinogens) or Pyrogens or endotoxins in the fluid path | Mild to Moderate Hypertension; Mild to Moderate Hypotension; Moderate Burn; Moderate Phlebitis; Moderate Pain; Fever; Reproductive system dysfunction | 3 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device must meet environmental requirements; Devices are biocompatible (Device material are complaint to ISO 10993); Packaging shall maintain a sterile barrier until expiry date | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Intravenous exposure to toxins | Fluid path contains toxins (E.g. systemic, mutagens, carcinogens) or Pyrogens or endotoxins in the fluid path | Mild Burn; Mild Phlebitis; Mild Pain | 2 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device must meet environmental requirements; Devices are biocompatible (Device material are complaint to ISO 10993); Packaging shall maintain a sterile barrier until expiry date | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Intravenous exposure to toxins | Fluid path contains toxins (E.g. systemic, mutagens, carcinogens) or Pyrogens or endotoxins in the fluid path | Temporary discomfort | 1 | 1 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device must meet environmental requirements; Devices are biocompatible (Device material are complaint to ISO 10993); Packaging shall maintain a sterile barrier until expiry date | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Intravenous exposure to toxins | Incompatibility of device materials and solutions | Death; Cardiopulmonary Arrest | 5 | 2 | Devices are biocompatible (Device material are complaint to ISO 10993) | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Intravenous exposure to toxins | Incompatibility of device materials and solutions | Paralysis; Organ failure/injury; Malignancy; Severe Hypertension; Severe Hypotension | 4 | 2 | Devices are biocompatible (Device material are complaint to ISO 10993) | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Intravenous exposure to toxins | Incompatibility of device materials and solutions | Mild to Moderate Hypertension; Mild to Moderate Hypotension; Moderate Burn; Moderate Phlebitis; Moderate Pain; Fever; Reproductive system dysfunction | 3 | 2 | Devices are biocompatible (Device material are complaint to ISO 10993) | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Intravenous exposure to toxins | Incompatibility of device materials and solutions | Mild Burn; Mild Phlebitis; Mild Pain | 2 | 2 | Devices are biocompatible (Device material are complaint to ISO 10993) | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Intravenous exposure to toxins | Incompatibility of device materials and solutions | Temporary discomfort | 1 | 1 | Devices are biocompatible (Device material are complaint to ISO 10993) | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Intravenous exposure to toxins | Loss of packaging integrity and caps off allows foreign material to enter into the sterile fluid path | Death; Cardiopulmonary Arrest | 5 | 2 | Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Devices should meet the storage and transport requirements; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Seal strength shall be sufficient to prevent the package from opening prior to point of use; Protective caps should be secure but easily removable | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Intravenous exposure to toxins | Loss of packaging integrity and caps off allows foreign material to enter into the sterile fluid path | Paralysis; Organ failure/injury; Malignancy; Severe Hypertension; Severe Hypotension | 4 | 2 | Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Devices should meet the storage and transport requirements; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Seal strength shall be sufficient to prevent the package from opening prior to point of use; Protective caps should be secure but easily removable | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Intravenous exposure to toxins | Loss of packaging integrity and caps off allows foreign material to enter into the sterile fluid path | Mild to Moderate Hypertension; Mild to Moderate Hypotension; Moderate Burn; Moderate Phlebitis; Moderate Pain; Fever; Reproductive system dysfunction | 3 | 2 | Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Devices should meet the storage and transport requirements; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Seal strength shall be sufficient to prevent the package from opening prior to point of use; Protective caps should be secure but easily removable | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Intravenous exposure to toxins | Loss of packaging integrity and caps off allows foreign material to enter into the sterile fluid path | Mild Burn; Mild Phlebitis; Mild Pain | 2 | 2 | Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Devices should meet the storage and transport requirements; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Seal strength shall be sufficient to prevent the package from opening prior to point of use; Protective caps should be secure but easily removable | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Intravenous exposure to toxins | Loss of packaging integrity and caps off allows foreign material to enter into the sterile fluid path | Temporary discomfort | 1 | 1 | Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Devices should meet the storage and transport requirements; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Seal strength shall be sufficient to prevent the package from opening prior to point of use; Protective caps should be secure but easily removable | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Intravenous exposure to toxins | Device materials are incompatible to prolonged exposure to disinfectant caps | Death; Cardiopulmonary Arrest | 5 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Intravenous exposure to toxins | Device materials are incompatible to prolonged exposure to disinfectant caps | Paralysis; Organ failure/injury; Malignancy; Severe Hypertension; Severe Hypotension | 4 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Intravenous exposure to toxins | Device materials are incompatible to prolonged exposure to disinfectant caps | Mild to Moderate Hypertension; Mild to Moderate Hypotension; Moderate Burn; Moderate Phlebitis; Moderate Pain; Fever; Reproductive system dysfunction | 3 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Intravenous exposure to toxins | Device materials are incompatible to prolonged exposure to disinfectant caps | Mild Burn; Mild Phlebitis; Mild Pain | 2 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
exposure to toxins (E.g. Pyrogens, cytokines, endotoxins) | Fumes or Vapors | Intravenous exposure to toxins | Device materials are incompatible to prolonged exposure to disinfectant caps | Temporary discomfort | 1 | 1 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
Open fluid path | Fluid Leak | Patient blood loss | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Death | 5 | 1 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Open fluid path | Fluid Leak | Patient blood loss | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Hypovolemic shock | 4 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Open fluid path | Fluid Leak | Patient blood loss | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Mild to Moderate Hypotension; Mild to Moderate Hypovolemia | 3 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Open fluid path | Fluid Leak | Patient blood loss | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Temporary Discomfort | 1 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Open fluid path | Fluid Leak | Patient blood loss | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Death | 5 | 1 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Open fluid path | Fluid Leak | Patient blood loss | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Hypovolemic shock | 4 | 2 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Open fluid path | Fluid Leak | Patient blood loss | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Mild to Moderate Hypotension; Mild to Moderate Hypovolemia | 3 | 2 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Open fluid path | Fluid Leak | Patient blood loss | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Temporary Discomfort | 1 | 2 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Open fluid path | Fluid Leak | Patient blood loss | Patient movement causes set/device to disconnect at a connection point | Death | 5 | 1 | Device designed easily attach/detach to connecting components; Device with rotating collars must meet rotating collar luer retention force per ISO 80369-7 requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | N/A
Open fluid path | Fluid Leak | Patient blood loss | Patient movement causes set/device to disconnect at a connection point | Hypovolemic shock | 4 | 2 | Device designed easily attach/detach to connecting components; Device with rotating collars must meet rotating collar luer retention force per ISO 80369-7 requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | N/A
Open fluid path | Fluid Leak | Patient blood loss | Patient movement causes set/device to disconnect at a connection point | Mild to Moderate Hypotension; Mild to Moderate Hypovolemia | 3 | 2 | Device designed easily attach/detach to connecting components; Device with rotating collars must meet rotating collar luer retention force per ISO 80369-7 requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | N/A
Open fluid path | Fluid Leak | Patient blood loss | Patient movement causes set/device to disconnect at a connection point | Temporary Discomfort | 1 | 2 | Device designed easily attach/detach to connecting components; Device with rotating collars must meet rotating collar luer retention force per ISO 80369-7 requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | N/A
Open fluid path | Fluid Leak | Patient blood loss | Silicone sleeve stuck down after disconnecting mating component | Death | 5 | 1 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
Open fluid path | Fluid Leak | Patient blood loss | Silicone sleeve stuck down after disconnecting mating component | Hypovolemic shock | 4 | 2 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
Open fluid path | Fluid Leak | Patient blood loss | Silicone sleeve stuck down after disconnecting mating component | Mild to Moderate Hypotension; Mild to Moderate Hypovolemia | 3 | 2 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
Open fluid path | Fluid Leak | Patient blood loss | Silicone sleeve stuck down after disconnecting mating component | Temporary Discomfort | 1 | 2 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
Open fluid path | Fluid Leak | Patient blood loss | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Death | 5 | 1 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Open fluid path | Fluid Leak | Patient blood loss | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Hypovolemic shock | 4 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Open fluid path | Fluid Leak | Patient blood loss | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Mild to Moderate Hypotension; Mild to Moderate Hypovolemia | 3 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Open fluid path | Fluid Leak | Patient blood loss | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Temporary Discomfort | 1 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Open fluid path | Fluid Leak | Patient blood loss | Device materials are incompatible to prolonged exposure to disinfectant caps | Death | 5 | 1 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements; Device designed to function after pressure infusion requirements | N/A
Open fluid path | Fluid Leak | Patient blood loss | Device materials are incompatible to prolonged exposure to disinfectant caps | Hypovolemic shock | 4 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements; Device designed to function after pressure infusion requirements | N/A
Open fluid path | Fluid Leak | Patient blood loss | Device materials are incompatible to prolonged exposure to disinfectant caps | Mild to Moderate Hypotension; Mild to Moderate Hypovolemia | 3 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements; Device designed to function after pressure infusion requirements | N/A
Open fluid path | Fluid Leak | Patient blood loss | Device materials are incompatible to prolonged exposure to disinfectant caps | Temporary Discomfort | 1 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements; Device designed to function after pressure infusion requirements | N/A
Open fluid path | Fluid Leak | Patient blood loss | Tubing rupture due to high pressure | Death | 5 | 1 | N/A | Administration set primary package labeling or instructions for use for extension sets or administration sets with distal injection access sites, excluding sets intended to be used with high pressure infusion equipment shall include a caution statement that set is not for high pressure infusion; Administration set primary package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 400 psig; Administration set with a MicroClave primary package labeling or instructions for use for high pressure extension sets contain a warning statement that high pressure infusion is conducted through the MicroClave in the middle of the assembly and the upstream clamp is engaged to isolate other upstream components from high pressures
Open fluid path | Fluid Leak | Patient blood loss | Tubing rupture due to high pressure | Hypovolemic shock | 4 | 2 | N/A | N/A
Open fluid path | Fluid Leak | Patient blood loss | Tubing rupture due to high pressure | Mild to Moderate Hypotension; Mild to Moderate Hypovolemia | 3 | 2 | N/A | N/A
Open fluid path | Fluid Leak | Patient blood loss | Tubing rupture due to high pressure | Temporary Discomfort | 1 | 2 | N/A | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Mild to Moderate Pain | 2 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Temporary discomfort/ User discomfort | 1 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Difficult / Unable to open the packaging | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 2 | Seal strength shall be sufficient to prevent the package from opening prior to point of use | Labeling includes instructions for using the device
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Difficult / Unable to open the packaging | Mild to Moderate Pain | 2 | 2 | Seal strength shall be sufficient to prevent the package from opening prior to point of use | Labeling includes instructions for using the device
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Difficult / Unable to open the packaging | Temporary discomfort/ User discomfort | 1 | 2 | Seal strength shall be sufficient to prevent the package from opening prior to point of use | Labeling includes instructions for using the device
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Difficult / Unable to remove Cap / cover / hood from the set/device | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 2 | Protective caps should be secure but easily removable | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Difficult / Unable to remove Cap / cover / hood from the set/device | Mild to Moderate Pain | 2 | 2 | Protective caps should be secure but easily removable | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Difficult / Unable to remove Cap / cover / hood from the set/device | Temporary discomfort/ User discomfort | 1 | 2 | Protective caps should be secure but easily removable | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 2 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Mild to Moderate Pain | 2 | 2 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Temporary discomfort/ User discomfort | 1 | 2 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Fluid path occlusion | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 2 | Device designed for 96 hours of continuous use; Device designed to maintain required flow rates and pressure leak requirements; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand internal back pressure requirements; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Fluid path occlusion | Mild to Moderate Pain | 2 | 2 | Device designed for 96 hours of continuous use; Device designed to maintain required flow rates and pressure leak requirements; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand internal back pressure requirements; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Fluid path occlusion | Temporary discomfort/ User discomfort | 1 | 2 | Device designed for 96 hours of continuous use; Device designed to maintain required flow rates and pressure leak requirements; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand internal back pressure requirements; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Mild to Moderate Pain | 2 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Temporary discomfort/ User discomfort | 1 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Problem identifying / selecting device | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 2 | Package labelling and printing must be intact and legible to point of use; Device must meet Labeling requirements (Individual device, the unit package, carton, case and shipment labeling) | Labeling provides information about the precautions (Eg., Usage of set/device, packaging conditions)
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Problem identifying / selecting device | Mild to Moderate Pain | 2 | 2 | Package labelling and printing must be intact and legible to point of use; Device must meet Labeling requirements (Individual device, the unit package, carton, case and shipment labeling) | Labeling provides information about the precautions (Eg., Usage of set/device, packaging conditions)
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Problem identifying / selecting device | Temporary discomfort/ User discomfort | 1 | 2 | Package labelling and printing must be intact and legible to point of use; Device must meet Labeling requirements (Individual device, the unit package, carton, case and shipment labeling) | Labeling provides information about the precautions (Eg., Usage of set/device, packaging conditions)
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Problem setting up / priming the device | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 2 | Components designed to turn easily; Device designed easily attach/detach to connecting components; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Devices should be free from cracks, splits, fractures; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Problem setting up / priming the device | Mild to Moderate Pain | 2 | 2 | Components designed to turn easily; Device designed easily attach/detach to connecting components; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Devices should be free from cracks, splits, fractures; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Problem setting up / priming the device | Temporary discomfort/ User discomfort | 1 | 2 | Components designed to turn easily; Device designed easily attach/detach to connecting components; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Devices should be free from cracks, splits, fractures; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Device appears unsafe to use and is replaced (Eg. Open package, Missing / Loose / Damaged components, Discoloration, Visible particulate etc.) | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 2 | Devices should be free from cracks, splits, fractures; Devices should meet the storage and transport requirements; Package labelling and printing must be intact and legible to point of use; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Seal strength shall be sufficient to prevent the package from opening prior to point of use | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Device appears unsafe to use and is replaced (Eg. Open package, Missing / Loose / Damaged components, Discoloration, Visible particulate etc.) | Mild to Moderate Pain | 2 | 2 | Devices should be free from cracks, splits, fractures; Devices should meet the storage and transport requirements; Package labelling and printing must be intact and legible to point of use; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Seal strength shall be sufficient to prevent the package from opening prior to point of use | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Device appears unsafe to use and is replaced (Eg. Open package, Missing / Loose / Damaged components, Discoloration, Visible particulate etc.) | Temporary discomfort/ User discomfort | 1 | 2 | Devices should be free from cracks, splits, fractures; Devices should meet the storage and transport requirements; Package labelling and printing must be intact and legible to point of use; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Seal strength shall be sufficient to prevent the package from opening prior to point of use | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | No Flow (Eg Vaccum / air lock) | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 2 | Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand stickdown and aspirate blood after 96 activation | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | No Flow (Eg Vaccum / air lock) | Mild to Moderate Pain | 2 | 2 | Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand stickdown and aspirate blood after 96 activation | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | No Flow (Eg Vaccum / air lock) | Temporary discomfort/ User discomfort | 1 | 2 | Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand stickdown and aspirate blood after 96 activation | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Device materials are incompatible to prolonged exposure to disinfectant caps | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Device materials are incompatible to prolonged exposure to disinfectant caps | Mild to Moderate Pain | 2 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Device materials are incompatible to prolonged exposure to disinfectant caps | Temporary discomfort/ User discomfort | 1 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Silicone sleeve stuck down after disconnecting mating component | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 2 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Silicone sleeve stuck down after disconnecting mating component | Mild to Moderate Pain | 2 | 2 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
setup delayed | Device Difficult to Setup or Prepare | Delayed delivery of Infusion solution | Silicone sleeve stuck down after disconnecting mating component | Temporary discomfort/ User discomfort | 1 | 2 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Death; Cardiopulmonary Arrest | 5 | 1 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Severe Arrhythmias; Severe Hypertension; Severe Hypotension; Status Epilepticus | 4 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Mild to Moderate Pain | 2 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Temporary discomfort | 1 | 3 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Difficult / Unable to connect / disconnect the set/device with other mating devices (eg Container, access site) | Death; Cardiopulmonary Arrest | 5 | 1 | Components designed to turn easily; Device designed easily attach/detach to connecting components; Device designed not to leak during normal product usage; Device should meet assembly integrity test requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Difficult / Unable to connect / disconnect the set/device with other mating devices (eg Container, access site) | Severe Arrhythmias; Severe Hypertension; Severe Hypotension; Status Epilepticus | 4 | 2 | Components designed to turn easily; Device designed easily attach/detach to connecting components; Device designed not to leak during normal product usage; Device should meet assembly integrity test requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Difficult / Unable to connect / disconnect the set/device with other mating devices (eg Container, access site) | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 2 | Components designed to turn easily; Device designed easily attach/detach to connecting components; Device designed not to leak during normal product usage; Device should meet assembly integrity test requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Difficult / Unable to connect / disconnect the set/device with other mating devices (eg Container, access site) | Mild to Moderate Pain | 2 | 2 | Components designed to turn easily; Device designed easily attach/detach to connecting components; Device designed not to leak during normal product usage; Device should meet assembly integrity test requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Difficult / Unable to connect / disconnect the set/device with other mating devices (eg Container, access site) | Temporary discomfort | 1 | 3 | Components designed to turn easily; Device designed easily attach/detach to connecting components; Device designed not to leak during normal product usage; Device should meet assembly integrity test requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Death; Cardiopulmonary Arrest | 5 | 1 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Severe Arrhythmias; Severe Hypertension; Severe Hypotension; Status Epilepticus | 4 | 2 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 2 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Mild to Moderate Pain | 2 | 2 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Temporary discomfort | 1 | 3 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Fluid path occlusion | Death; Cardiopulmonary Arrest | 5 | 1 | Device designed for 96 hours of continuous use; Device designed to maintain required flow rates and pressure leak requirements; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand internal back pressure requirements; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | N/A
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Fluid path occlusion | Severe Arrhythmias; Severe Hypertension; Severe Hypotension; Status Epilepticus | 4 | 2 | Device designed for 96 hours of continuous use; Device designed to maintain required flow rates and pressure leak requirements; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand internal back pressure requirements; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | N/A
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Fluid path occlusion | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 2 | Device designed for 96 hours of continuous use; Device designed to maintain required flow rates and pressure leak requirements; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand internal back pressure requirements; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | N/A
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Fluid path occlusion | Mild to Moderate Pain | 2 | 2 | Device designed for 96 hours of continuous use; Device designed to maintain required flow rates and pressure leak requirements; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand internal back pressure requirements; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | N/A
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Fluid path occlusion | Temporary discomfort | 1 | 3 | Device designed for 96 hours of continuous use; Device designed to maintain required flow rates and pressure leak requirements; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand internal back pressure requirements; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | N/A
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Death; Cardiopulmonary Arrest | 5 | 1 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Severe Arrhythmias; Severe Hypertension; Severe Hypotension; Status Epilepticus | 4 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Mild to Moderate Pain | 2 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Temporary discomfort | 1 | 3 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Patient movement causes set/device to disconnect at a connection point | Death; Cardiopulmonary Arrest | 5 | 1 | Device designed easily attach/detach to connecting components; Device with rotating collars must meet rotating collar luer retention force per ISO 80369-7 requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | N/A
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Patient movement causes set/device to disconnect at a connection point | Severe Arrhythmias; Severe Hypertension; Severe Hypotension; Status Epilepticus | 4 | 2 | Device designed easily attach/detach to connecting components; Device with rotating collars must meet rotating collar luer retention force per ISO 80369-7 requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | N/A
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Patient movement causes set/device to disconnect at a connection point | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 2 | Device designed easily attach/detach to connecting components; Device with rotating collars must meet rotating collar luer retention force per ISO 80369-7 requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | N/A
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Patient movement causes set/device to disconnect at a connection point | Mild to Moderate Pain | 2 | 2 | Device designed easily attach/detach to connecting components; Device with rotating collars must meet rotating collar luer retention force per ISO 80369-7 requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | N/A
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Patient movement causes set/device to disconnect at a connection point | Temporary discomfort | 1 | 3 | Device designed easily attach/detach to connecting components; Device with rotating collars must meet rotating collar luer retention force per ISO 80369-7 requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | N/A
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Device materials are incompatible to prolonged exposure to disinfectant caps | Death; Cardiopulmonary Arrest | 5 | 1 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Device materials are incompatible to prolonged exposure to disinfectant caps | Severe Arrhythmias; Severe Hypertension; Severe Hypotension; Status Epilepticus | 4 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Device materials are incompatible to prolonged exposure to disinfectant caps | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Device materials are incompatible to prolonged exposure to disinfectant caps | Mild to Moderate Pain | 2 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Device materials are incompatible to prolonged exposure to disinfectant caps | Temporary discomfort | 1 | 3 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Silicone sleeve stuck down after disconnecting mating component | Death; Cardiopulmonary Arrest | 5 | 1 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Silicone sleeve stuck down after disconnecting mating component | Severe Arrhythmias; Severe Hypertension; Severe Hypotension; Status Epilepticus | 4 | 2 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Silicone sleeve stuck down after disconnecting mating component | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 2 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Silicone sleeve stuck down after disconnecting mating component | Mild to Moderate Pain | 2 | 2 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Silicone sleeve stuck down after disconnecting mating component | Temporary discomfort | 1 | 3 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Tubing rupture due to high pressure | Death; Cardiopulmonary Arrest | 5 | 1 | Device designed to function after pressure infusion requirements | Neutron/Neutron Ext Set package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 350psig or 10mL per second of contrast media; Administration set primary package labeling or instructions for use for extension sets or administration sets with distal injection access sites, excluding sets intended to be used with high pressure infusion equipment shall include a caution statement that set is not for high pressure infusion; Administration set primary package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 400 psig; Administration set with a MicroClave primary package labeling or instructions for use for high pressure extension sets contain a warning statement that high pressure infusion is conducted through the MicroClave in the middle of the assembly and the upstream clamp is engaged to isolate other upstream components from high pressures
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Tubing rupture due to high pressure | Severe Arrhythmias; Severe Hypertension; Severe Hypotension; Status Epilepticus | 4 | 2 | Device designed to function after pressure infusion requirements | Neutron/Neutron Ext Set package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 350psig or 10mL per second of contrast media; Administration set primary package labeling or instructions for use for extension sets or administration sets with distal injection access sites, excluding sets intended to be used with high pressure infusion equipment shall include a caution statement that set is not for high pressure infusion; Administration set primary package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 400 psig; Administration set with a MicroClave primary package labeling or instructions for use for high pressure extension sets contain a warning statement that high pressure infusion is conducted through the MicroClave in the middle of the assembly and the upstream clamp is engaged to isolate other upstream components from high pressures
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Tubing rupture due to high pressure | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 2 | Device designed to function after pressure infusion requirements | Neutron/Neutron Ext Set package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 350psig or 10mL per second of contrast media; Administration set primary package labeling or instructions for use for extension sets or administration sets with distal injection access sites, excluding sets intended to be used with high pressure infusion equipment shall include a caution statement that set is not for high pressure infusion; Administration set primary package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 400 psig; Administration set with a MicroClave primary package labeling or instructions for use for high pressure extension sets contain a warning statement that high pressure infusion is conducted through the MicroClave in the middle of the assembly and the upstream clamp is engaged to isolate other upstream components from high pressures
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Tubing rupture due to high pressure | Mild to Moderate Pain | 2 | 2 | Device designed to function after pressure infusion requirements | Neutron/Neutron Ext Set package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 350psig or 10mL per second of contrast media; Administration set primary package labeling or instructions for use for extension sets or administration sets with distal injection access sites, excluding sets intended to be used with high pressure infusion equipment shall include a caution statement that set is not for high pressure infusion; Administration set primary package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 400 psig; Administration set with a MicroClave primary package labeling or instructions for use for high pressure extension sets contain a warning statement that high pressure infusion is conducted through the MicroClave in the middle of the assembly and the upstream clamp is engaged to isolate other upstream components from high pressures
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Tubing rupture due to high pressure | Temporary discomfort | 1 | 3 | Device designed to function after pressure infusion requirements | Neutron/Neutron Ext Set package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 350psig or 10mL per second of contrast media; Administration set primary package labeling or instructions for use for extension sets or administration sets with distal injection access sites, excluding sets intended to be used with high pressure infusion equipment shall include a caution statement that set is not for high pressure infusion; Administration set primary package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 400 psig; Administration set with a MicroClave primary package labeling or instructions for use for high pressure extension sets contain a warning statement that high pressure infusion is conducted through the MicroClave in the middle of the assembly and the upstream clamp is engaged to isolate other upstream components from high pressures
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Device appears unsafe to use and is replaced (Eg. Open package, Missing / Loose / Damaged components, Discoloration, Visible particulate etc.) | Death; Cardiopulmonary Arrest | 5 | 1 | Devices should be free from cracks, splits, fractures; Devices should meet the storage and transport requirements; Package labelling and printing must be intact and legible to point of use; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Seal strength shall be sufficient to prevent the package from opening prior to point of use | N/A
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Device appears unsafe to use and is replaced (Eg. Open package, Missing / Loose / Damaged components, Discoloration, Visible particulate etc.) | Severe Arrhythmias; Severe Hypertension; Severe Hypotension; Status Epilepticus | 4 | 2 | Devices should be free from cracks, splits, fractures; Devices should meet the storage and transport requirements; Package labelling and printing must be intact and legible to point of use; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Seal strength shall be sufficient to prevent the package from opening prior to point of use | N/A
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Device appears unsafe to use and is replaced (Eg. Open package, Missing / Loose / Damaged components, Discoloration, Visible particulate etc.) | Mild to Moderate Arrhythmia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Seizure; Severe Pain | 3 | 2 | Devices should be free from cracks, splits, fractures; Devices should meet the storage and transport requirements; Package labelling and printing must be intact and legible to point of use; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Seal strength shall be sufficient to prevent the package from opening prior to point of use | N/A
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Device appears unsafe to use and is replaced (Eg. Open package, Missing / Loose / Damaged components, Discoloration, Visible particulate etc.) | Mild to Moderate Pain | 2 | 2 | Devices should be free from cracks, splits, fractures; Devices should meet the storage and transport requirements; Package labelling and printing must be intact and legible to point of use; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Seal strength shall be sufficient to prevent the package from opening prior to point of use | N/A
therapy interrupted | Failure to Deliver | Delayed delivery of infusion solution | Device appears unsafe to use and is replaced (Eg. Open package, Missing / Loose / Damaged components, Discoloration, Visible particulate etc.) | Temporary discomfort | 1 | 3 | Devices should be free from cracks, splits, fractures; Devices should meet the storage and transport requirements; Package labelling and printing must be intact and legible to point of use; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Seal strength shall be sufficient to prevent the package from opening prior to point of use | N/A
Blood clot | Occlusion within Device | Embolization into the vascular system | Blood clot (from blood product, or access site) introduced into the device/set | Death; Cardiopulmonary Arrest | 5 | 1 | Devices should be free from cracks, splits, fractures; Device should not exceed residual volume requirements | Administration set primary package labeling or instructions for use for transfusion sets include a warning statement that other solutions should never be infused with blood or blood products except for those approved by your institution's policies and procedures
Blood clot | Occlusion within Device | Embolization into the vascular system | Blood clot (from blood product, or access site) introduced into the device/set | Pulmonary embolism; Myocardial infarction; Stroke | 4 | 1 | Devices should be free from cracks, splits, fractures; Device should not exceed residual volume requirements | Administration set primary package labeling or instructions for use for transfusion sets include a warning statement that other solutions should never be infused with blood or blood products except for those approved by your institution's policies and procedures
Blood clot | Occlusion within Device | Embolization into the vascular system | Blood clot (from blood product, or access site) introduced into the device/set | Extremity swelling Pain; Moderate Paresthesias | 3 | 1 | Devices should be free from cracks, splits, fractures; Device should not exceed residual volume requirements | Administration set primary package labeling or instructions for use for transfusion sets include a warning statement that other solutions should never be infused with blood or blood products except for those approved by your institution's policies and procedures
Blood clot | Occlusion within Device | Embolization into the vascular system | Blood clot (from blood product, or access site) introduced into the device/set | Thrombophlebitis or pain | 2 | 1 | Devices should be free from cracks, splits, fractures; Device should not exceed residual volume requirements | Administration set primary package labeling or instructions for use for transfusion sets include a warning statement that other solutions should never be infused with blood or blood products except for those approved by your institution's policies and procedures
Blood clot | Occlusion within Device | Embolization into the vascular system | Blood clot (from blood product, or access site) introduced into the device/set | Temporary Discomfort | 1 | 1 | Devices should be free from cracks, splits, fractures; Device should not exceed residual volume requirements | Administration set primary package labeling or instructions for use for transfusion sets include a warning statement that other solutions should never be infused with blood or blood products except for those approved by your institution's policies and procedures
particulate presented | Particulates | Embolization into the vascular system | Loss of packaging integrity and caps off allows foreign material to enter into the sterile fluid path | Death Cardiopulmonary Arrest | 5 | 1 | Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Devices should meet the storage and transport requirements; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date'; Seal strength shall be sufficient to prevent the package from opening prior to point of use; Protective caps should be secure but easily removable | N/A
particulate presented | Particulates | Embolization into the vascular system | Loss of packaging integrity and caps off allows foreign material to enter into the sterile fluid path | Pulmonary embolism Myocardial infarction Stroke | 4 | 2 | Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Devices should meet the storage and transport requirements; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date'; Seal strength shall be sufficient to prevent the package from opening prior to point of use; Protective caps should be secure but easily removable | N/A
particulate presented | Particulates | Embolization into the vascular system | Loss of packaging integrity and caps off allows foreign material to enter into the sterile fluid path | Extremity swelling Pain, Moderate Paresthesias | 3 | 2 | Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Devices should meet the storage and transport requirements; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date'; Seal strength shall be sufficient to prevent the package from opening prior to point of use; Protective caps should be secure but easily removable | N/A
particulate presented | Particulates | Embolization into the vascular system | Loss of packaging integrity and caps off allows foreign material to enter into the sterile fluid path | Thrombophlebitis Pain, Mild | 2 | 2 | Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Devices should meet the storage and transport requirements; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date'; Seal strength shall be sufficient to prevent the package from opening prior to point of use; Protective caps should be secure but easily removable | N/A
particulate presented | Particulates | Embolization into the vascular system | Loss of packaging integrity and caps off allows foreign material to enter into the sterile fluid path | Temporary Discomfort | 1 | 2 | Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Devices should meet the storage and transport requirements; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date'; Seal strength shall be sufficient to prevent the package from opening prior to point of use; Protective caps should be secure but easily removable | N/A
particulate presented | Particulates | Embolization into the vascular system | Particulate in the fluid path | Death Cardiopulmonary Arrest | 5 | 1 | Device designed to maintain sterility requirements; Device must meet the chemical compatibility requirements; Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Environmental contol in manufacturing facility | N/A
particulate presented | Particulates | Embolization into the vascular system | Particulate in the fluid path | Pulmonary embolism Myocardial infarction Stroke | 4 | 2 | Device designed to maintain sterility requirements; Device must meet the chemical compatibility requirements; Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Environmental contol in manufacturing facility | N/A
particulate presented | Particulates | Embolization into the vascular system | Particulate in the fluid path | Extremity swelling Pain, Moderate Paresthesias | 3 | 2 | Device designed to maintain sterility requirements; Device must meet the chemical compatibility requirements; Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Environmental contol in manufacturing facility | N/A
particulate presented | Particulates | Embolization into the vascular system | Particulate in the fluid path | Thrombophlebitis Pain, Mild | 2 | 2 | Device designed to maintain sterility requirements; Device must meet the chemical compatibility requirements; Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Environmental contol in manufacturing facility | N/A
particulate presented | Particulates | Embolization into the vascular system | Particulate in the fluid path | Temporary Discomfort | 1 | 2 | Device designed to maintain sterility requirements; Device must meet the chemical compatibility requirements; Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Environmental contol in manufacturing facility | N/A
particulate presented | Particulates | Embolization into the vascular system | Device materials are incompatible to prolonged exposure to disinfectant caps | Death Cardiopulmonary Arrest | 5 | 1 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
particulate presented | Particulates | Embolization into the vascular system | Device materials are incompatible to prolonged exposure to disinfectant caps | Pulmonary embolism Myocardial infarction Stroke | 4 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
particulate presented | Particulates | Embolization into the vascular system | Device materials are incompatible to prolonged exposure to disinfectant caps | Extremity swelling Pain, Moderate Paresthesias | 3 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
particulate presented | Particulates | Embolization into the vascular system | Device materials are incompatible to prolonged exposure to disinfectant caps | Thrombophlebitis Pain, Mild | 2 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
particulate presented | Particulates | Embolization into the vascular system | Device materials are incompatible to prolonged exposure to disinfectant caps | Temporary Discomfort | 1 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/fumes | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Mild to Moderate Respiratory distress; Moderate toxic effects such as seizures, abdominal pain | 3 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/fumes | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Mild Pain; Mild toxic effects such as fatigue, sleepiness, headaches, eye irritation, skin rashes | 2 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/fumes | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Temporary discomfort | 1 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/fumes | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Mild to Moderate Respiratory distress; Moderate toxic effects such as seizures, abdominal pain | 3 | 2 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/fumes | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Mild Pain; Mild toxic effects such as fatigue, sleepiness, headaches, eye irritation, skin rashes | 2 | 2 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/fumes | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Temporary discomfort | 1 | 2 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/fumes | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Mild to Moderate Respiratory distress; Moderate toxic effects such as seizures, abdominal pain | 3 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/fumes | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Mild Pain; Mild toxic effects such as fatigue, sleepiness, headaches, eye irritation, skin rashes | 2 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/fumes | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Temporary discomfort | 1 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/fumes | Silicone sleeve stuck down after disconnecting mating component | Mild to Moderate Respiratory distress; Moderate toxic effects such as seizures, abdominal pain | 3 | 2 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/fumes | Silicone sleeve stuck down after disconnecting mating component | Mild Pain; Mild toxic effects such as fatigue, sleepiness, headaches, eye irritation, skin rashes | 2 | 2 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/fumes | Silicone sleeve stuck down after disconnecting mating component | Temporary discomfort | 1 | 2 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/fumes | Tubing rupture due to high pressure | Mild to Moderate Respiratory distress; Moderate toxic effects such as seizures, abdominal pain | 3 | 2 | Device designed to function after pressure infusion requirements | Neutron/Neutron Ext Set package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 350psig or 10mL per second of contrast media; Administration set primary package labeling or instructions for use for extension sets or administration sets with distal injection access sites, excluding sets intended to be used with high pressure infusion equipment shall include a caution statement that set is not for high pressure infusion; Administration set primary package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 400 psig; Administration set with a MicroClave primary package labeling or instructions for use for high pressure extension sets contain a warning statement that high pressure infusion is conducted through the MicroClave in the middle of the assembly and the upstream clamp is engaged to isolate other upstream components from high pressures
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/fumes | Tubing rupture due to high pressure | Mild Pain; Mild toxic effects such as fatigue, sleepiness, headaches, eye irritation, skin rashes | 2 | 2 | Device designed to function after pressure infusion requirements | Neutron/Neutron Ext Set package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 350psig or 10mL per second of contrast media; Administration set primary package labeling or instructions for use for extension sets or administration sets with distal injection access sites, excluding sets intended to be used with high pressure infusion equipment shall include a caution statement that set is not for high pressure infusion; Administration set primary package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 400 psig; Administration set with a MicroClave primary package labeling or instructions for use for high pressure extension sets contain a warning statement that high pressure infusion is conducted through the MicroClave in the middle of the assembly and the upstream clamp is engaged to isolate other upstream components from high pressures
Hazardous fluids/fumes released into environment | Contamination During Use | Exposure to hazardous fluids/fumes | Tubing rupture due to high pressure | Temporary discomfort | 1 | 2 | Device designed to function after pressure infusion requirements | Neutron/Neutron Ext Set package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 350psig or 10mL per second of contrast media; Administration set primary package labeling or instructions for use for extension sets or administration sets with distal injection access sites, excluding sets intended to be used with high pressure infusion equipment shall include a caution statement that set is not for high pressure infusion; Administration set primary package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 400 psig; Administration set with a MicroClave primary package labeling or instructions for use for high pressure extension sets contain a warning statement that high pressure infusion is conducted through the MicroClave in the middle of the assembly and the upstream clamp is engaged to isolate other upstream components from high pressures
Inappropriate disposal | Environmental Control or Utility Issue | Environmental (Landfill/ Water Supply) exposure to toxic substances from set material | Device contains toxic materials | Moderate toxic effects such as seizures, abdominal pain | 3 | 1 | Device must meet environmental requirements; Devices are biocompatible (Device material are complaint to ISO 10993) | N/A
Inappropriate disposal | Environmental Control or Utility Issue | Environmental (Landfill/ Water Supply) exposure to toxic substances from set material | Device contains toxic materials | Mild toxic effects such as fatigue, sleepiness, headaches, eye irritation, skin rashes | 2 | 1 | Device must meet environmental requirements; Devices are biocompatible (Device material are complaint to ISO 10993) | N/A
Inappropriate disposal | Environmental Control or Utility Issue | Environmental (Landfill/ Water Supply) exposure to toxic substances from set material | Device contains toxic materials | Temporary discomfort | 1 | 1 | Device must meet environmental requirements; Devices are biocompatible (Device material are complaint to ISO 10993) | N/A
Inappropriate disposal | Environmental Control or Utility Issue | Environmental (Landfill/ Water Supply) exposure to toxic substances from set material | Proper disposal procedure not followed | Moderate toxic effects such as seizures, abdominal pain | 3 | 1 | Device must meet environmental requirements | N/A
Inappropriate disposal | Environmental Control or Utility Issue | Environmental (Landfill/ Water Supply) exposure to toxic substances from set material | Proper disposal procedure not followed | Mild toxic effects such as fatigue, sleepiness, headaches, eye irritation, skin rashes | 2 | 1 | Device must meet environmental requirements | N/A
Inappropriate disposal | Environmental Control or Utility Issue | Environmental (Landfill/ Water Supply) exposure to toxic substances from set material | Proper disposal procedure not followed | Temporary discomfort | 1 | 1 | Device must meet environmental requirements | N/A
Wrong route | Misconnection | Incorrect access connection | Fluid infused through wrong route (Misconnection) | Death; Cardiopulmonary Arrest(Cardiac arrest) | 5 | 1 | Device must meet Labeling requirements (Individual device, the unit package, carton, case and shipment labeling); Package labelling and printing must be intact and legible to point of use; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device
Wrong route | Misconnection | Incorrect access connection | Fluid infused through wrong route (Misconnection) | Stroke; Organ injury or failure; Pulmonary embolism; Severe Respiratory Distress; Severe Hypertension; Severe Hypotension; Severe Hyperglycemia; Hypoglycemia | 4 | 2 | Device must meet Labeling requirements (Individual device, the unit package, carton, case and shipment labeling); Package labelling and printing must be intact and legible to point of use; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device
Wrong route | Misconnection | Incorrect access connection | Fluid infused through wrong route (Misconnection) | Mild to Moderate Respiratory Distress; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Moderate Hyperglycemia; Moderate Hypoglycemia | 3 | 2 | Device must meet Labeling requirements (Individual device, the unit package, carton, case and shipment labeling); Package labelling and printing must be intact and legible to point of use; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device
Wrong route | Misconnection | Incorrect access connection | Fluid infused through wrong route (Misconnection) | Mild Hyperglycemia; Mild Hypoglycemia | 2 | 1 | Device must meet Labeling requirements (Individual device, the unit package, carton, case and shipment labeling); Package labelling and printing must be intact and legible to point of use; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device
Wrong route | Misconnection | Incorrect access connection | Fluid infused through wrong route (Misconnection) | Temporary discomfort | 1 | 1 | Device must meet Labeling requirements (Individual device, the unit package, carton, case and shipment labeling); Package labelling and printing must be intact and legible to point of use; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Aseptic techniques not followed | Death; Cardiopulmonary Arrest | 5 | 2 | N/A | Label instructs to use aseptic technique to prevent contamination
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Aseptic techniques not followed | Septic shock; Organ injury | 4 | 2 | N/A | Label instructs to use aseptic technique to prevent contamination
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Aseptic techniques not followed | Bacteremia; Moderate Pain | 3 | 2 | N/A | Label instructs to use aseptic technique to prevent contamination
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Aseptic techniques not followed | Mild Phlebitis Local reaction/ infection; Mild Pain | 2 | 2 | N/A | Label instructs to use aseptic technique to prevent contamination
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Aseptic techniques not followed | Temporary discomfort | 1 | 2 | N/A | Label instructs to use aseptic technique to prevent contamination
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Difficult / Unable to clean access sites | Death; Cardiopulmonary Arrest | 5 | 2 | Devices should be free from cracks, splits, fractures; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed for 96 hours of continuous use; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Difficult / Unable to clean access sites | Septic shock; Organ injury | 4 | 2 | Devices should be free from cracks, splits, fractures; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed for 96 hours of continuous use; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Difficult / Unable to clean access sites | Bacteremia; Moderate Pain | 3 | 2 | Devices should be free from cracks, splits, fractures; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed for 96 hours of continuous use; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Difficult / Unable to clean access sites | Mild Phlebitis Local reaction/ infection; Mild Pain | 2 | 2 | Devices should be free from cracks, splits, fractures; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed for 96 hours of continuous use; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Difficult / Unable to clean access sites | Temporary discomfort | 1 | 2 | Devices should be free from cracks, splits, fractures; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed for 96 hours of continuous use; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Fluid path sterility compromised due to broken, loose or unsealed fluid path | Death; Cardiopulmonary Arrest | 5 | 2 | Device designed easily attach/detach to connecting components; Device designed not to leak during normal product usage; Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements; Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Components must not be cracked, broken, distorted, or severely discolored; Devices should meet the storage and transport requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection; Device should meet assembly integrity test requirements | Labeling includes instructions for using the device
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Fluid path sterility compromised due to broken, loose or unsealed fluid path | Septic shock; Organ injury | 4 | 2 | Device designed easily attach/detach to connecting components; Device designed not to leak during normal product usage; Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements; Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Components must not be cracked, broken, distorted, or severely discolored; Devices should meet the storage and transport requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection; Device should meet assembly integrity test requirements | Labeling includes instructions for using the device
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Fluid path sterility compromised due to broken, loose or unsealed fluid path | Bacteremia; Moderate Pain | 3 | 2 | Device designed easily attach/detach to connecting components; Device designed not to leak during normal product usage; Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements; Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Components must not be cracked, broken, distorted, or severely discolored; Devices should meet the storage and transport requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection; Device should meet assembly integrity test requirements | Labeling includes instructions for using the device
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Fluid path sterility compromised due to broken, loose or unsealed fluid path | Mild Phlebitis Local reaction/ infection; Mild Pain | 2 | 2 | Device designed easily attach/detach to connecting components; Device designed not to leak during normal product usage; Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements; Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Components must not be cracked, broken, distorted, or severely discolored; Devices should meet the storage and transport requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection; Device should meet assembly integrity test requirements | Labeling includes instructions for using the device
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Fluid path sterility compromised due to broken, loose or unsealed fluid path | Temporary discomfort | 1 | 2 | Device designed easily attach/detach to connecting components; Device designed not to leak during normal product usage; Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements; Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Components must not be cracked, broken, distorted, or severely discolored; Devices should meet the storage and transport requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection; Device should meet assembly integrity test requirements | Labeling includes instructions for using the device
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Death; Cardiopulmonary Arrest | 5 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Septic shock; Organ injury | 4 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Bacteremia; Moderate Pain | 3 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Mild Phlebitis Local reaction/ infection; Mild Pain | 2 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Temporary discomfort | 1 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Organism ingress due to unsterile fluid in the interstitial space of the needleless valve components | Death; Cardiopulmonary Arrest | 5 | 2 | Device designed not to leak during normal product usage; Device designed to prevent the leak when subject to internal negative pressure; Device designed to maintain sterility requirements; Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements; Device should not exceed residual volume requirements | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Organism ingress due to unsterile fluid in the interstitial space of the needleless valve components | Septic shock; Organ injury | 4 | 2 | Device designed not to leak during normal product usage; Device designed to prevent the leak when subject to internal negative pressure; Device designed to maintain sterility requirements; Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements; Device should not exceed residual volume requirements | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Organism ingress due to unsterile fluid in the interstitial space of the needleless valve components | Bacteremia; Moderate Pain | 3 | 2 | Device designed not to leak during normal product usage; Device designed to prevent the leak when subject to internal negative pressure; Device designed to maintain sterility requirements; Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements; Device should not exceed residual volume requirements | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Organism ingress due to unsterile fluid in the interstitial space of the needleless valve components | Mild Phlebitis Local reaction/ infection; Mild Pain | 2 | 2 | Device designed not to leak during normal product usage; Device designed to prevent the leak when subject to internal negative pressure; Device designed to maintain sterility requirements; Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements; Device should not exceed residual volume requirements | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Organism ingress due to unsterile fluid in the interstitial space of the needleless valve components | Temporary discomfort | 1 | 2 | Device designed not to leak during normal product usage; Device designed to prevent the leak when subject to internal negative pressure; Device designed to maintain sterility requirements; Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements; Device should not exceed residual volume requirements | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Reuse of the same device/set on multiple patients | Death; Cardiopulmonary Arrest | 5 | 2 | N/A | Labeling includes instructions for using the device; Labeling provides information about the precautions (Eg., Usage of set/device, packaging conditions); Labeling provide instructions not reuse set/device which are labeled for single use only; Needleless Connector Piggyback Adapter labeling provide information NOT to use on pre-slit needleless access ports; Labeling provide instructions not to use needles, blunt cannulas or luer caps on luer activated injection access sites (Eg., Needlefree connectors); Labeling includes instructions for using the device
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Reuse of the same device/set on multiple patients | Septic shock; Organ injury | 4 | 2 | N/A | Labeling includes instructions for using the device; Labeling provides information about the precautions (Eg., Usage of set/device, packaging conditions); Labeling provide instructions not reuse set/device which are labeled for single use only; Needleless Connector Piggyback Adapter labeling provide information NOT to use on pre-slit needleless access ports; Labeling provide instructions not to use needles, blunt cannulas or luer caps on luer activated injection access sites (Eg., Needlefree connectors); Labeling includes instructions for using the device
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Reuse of the same device/set on multiple patients | Bacteremia; Moderate Pain | 3 | 2 | N/A | Labeling includes instructions for using the device; Labeling provides information about the precautions (Eg., Usage of set/device, packaging conditions); Labeling provide instructions not reuse set/device which are labeled for single use only; Needleless Connector Piggyback Adapter labeling provide information NOT to use on pre-slit needleless access ports; Labeling provide instructions not to use needles, blunt cannulas or luer caps on luer activated injection access sites (Eg., Needlefree connectors); Labeling includes instructions for using the device
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Reuse of the same device/set on multiple patients | Mild Phlebitis; Local reaction/ infection; Mild Pain | 2 | 2 | N/A | Labeling includes instructions for using the device; Labeling provides information about the precautions (Eg., Usage of set/device, packaging conditions); Labeling provide instructions not reuse set/device which are labeled for single use only; Needleless Connector Piggyback Adapter labeling provide information NOT to use on pre-slit needleless access ports; Labeling provide instructions not to use needles, blunt cannulas or luer caps on luer activated injection access sites (Eg., Needlefree connectors); Labeling includes instructions for using the device
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Reuse of the same device/set on multiple patients | Temporary discomfort | 1 | 2 | N/A | Labeling includes instructions for using the device; Labeling provides information about the precautions (Eg., Usage of set/device, packaging conditions); Labeling provide instructions not reuse set/device which are labeled for single use only; Needleless Connector Piggyback Adapter labeling provide information NOT to use on pre-slit needleless access ports; Labeling provide instructions not to use needles, blunt cannulas or luer caps on luer activated injection access sites (Eg., Needlefree connectors); Labeling includes instructions for using the device
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Needle sticks | Septic shock; Organ injury | 4 | 2 | N/A | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Needle sticks | Bacteremia; Moderate Pain | 3 | 2 | N/A | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Needle sticks | Mild Phlebitis; Local reaction/ infection; Mild Pain | 2 | 2 | N/A | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Needle sticks | Temporary discomfort | 1 | 2 | N/A | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Device materials are incompatible to prolonged exposure to disinfectant caps | Death; Cardiopulmonary Arrest | 5 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Device materials are incompatible to prolonged exposure to disinfectant caps | Septic shock; Organ injury | 4 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Device materials are incompatible to prolonged exposure to disinfectant caps | Bacteremia; Moderate Pain | 3 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Device materials are incompatible to prolonged exposure to disinfectant caps | Mild Phlebitis; Local reaction/ infection; Mild Pain | 2 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Device materials are incompatible to prolonged exposure to disinfectant caps | Temporary discomfort | 1 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Sterility compromised due to loss of packaging integrity | Death; Cardiopulmonary Arrest | 5 | 2 | Components must not be cracked, broken, distorted, or severely discolored; Devices should meet the storage and transport requirements; Package labelling and printing must be intact and legible to point of use; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Seal strength shall be sufficient to prevent the package from opening prior to point of use | Administration set primary package labeling for sterile content sets indicate that the set is sterile and nonpyrogenic in intact packaging and shall not be re-sterilized
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Sterility compromised due to loss of packaging integrity | Septic shock; Organ injury | 4 | 2 | Components must not be cracked, broken, distorted, or severely discolored; Devices should meet the storage and transport requirements; Package labelling and printing must be intact and legible to point of use; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Seal strength shall be sufficient to prevent the package from opening prior to point of use | Administration set primary package labeling for sterile content sets indicate that the set is sterile and nonpyrogenic in intact packaging and shall not be re-sterilized
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Sterility compromised due to loss of packaging integrity | Bacteremia; Moderate Pain | 3 | 2 | Components must not be cracked, broken, distorted, or severely discolored; Devices should meet the storage and transport requirements; Package labelling and printing must be intact and legible to point of use; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Seal strength shall be sufficient to prevent the package from opening prior to point of use | Administration set primary package labeling for sterile content sets indicate that the set is sterile and nonpyrogenic in intact packaging and shall not be re-sterilized
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Sterility compromised due to loss of packaging integrity | Mild Phlebitis; Local reaction/ infection; Mild Pain | 2 | 2 | Components must not be cracked, broken, distorted, or severely discolored; Devices should meet the storage and transport requirements; Package labelling and printing must be intact and legible to point of use; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Seal strength shall be sufficient to prevent the package from opening prior to point of use | Administration set primary package labeling for sterile content sets indicate that the set is sterile and nonpyrogenic in intact packaging and shall not be re-sterilized
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Sterility compromised due to loss of packaging integrity | Temporary discomfort | 1 | 2 | Components must not be cracked, broken, distorted, or severely discolored; Devices should meet the storage and transport requirements; Package labelling and printing must be intact and legible to point of use; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Seal strength shall be sufficient to prevent the package from opening prior to point of use | Administration set primary package labeling for sterile content sets indicate that the set is sterile and nonpyrogenic in intact packaging and shall not be re-sterilized
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | The device is used for a longer period than specified | Death; Cardiopulmonary Arrest | 5 | 2 | N/A | Labeling provides information about legal manufacturer and date of manufacture; Labeling provides information about the precautions (Eg., Usage of set/device, packaging conditions); Labeling includes an expiry date where required by country or regional regulation
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | The device is used for a longer period than specified | Septic shock; Organ injury | 4 | 2 | N/A | Labeling provides information about legal manufacturer and date of manufacture; Labeling provides information about the precautions (Eg., Usage of set/device, packaging conditions); Labeling includes an expiry date where required by country or regional regulation
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | The device is used for a longer period than specified | Bacteremia; Moderate Pain | 3 | 2 | N/A | Labeling provides information about legal manufacturer and date of manufacture; Labeling provides information about the precautions (Eg., Usage of set/device, packaging conditions); Labeling includes an expiry date where required by country or regional regulation
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | The device is used for a longer period than specified | Mild Phlebitis; Local reaction/ infection; Mild Pain | 2 | 2 | N/A | Labeling provides information about legal manufacturer and date of manufacture; Labeling provides information about the precautions (Eg., Usage of set/device, packaging conditions); Labeling includes an expiry date where required by country or regional regulation
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | The device is used for a longer period than specified | Temporary discomfort | 1 | 2 | N/A | Labeling provides information about legal manufacturer and date of manufacture; Labeling provides information about the precautions (Eg., Usage of set/device, packaging conditions); Labeling includes an expiry date where required by country or regional regulation
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Use of expired device | Death; Cardiopulmonary Arrest | 5 | 2 | N/A | Labeling provides information about legal manufacturer and date of manufacture; Labeling provides information about the precautions (Eg., Usage of set/device, packaging conditions); Labeling includes an expiry date where required by country or regional regulation
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Use of expired device | Septic shock; Organ injury | 4 | 2 | N/A | Labeling provides information about legal manufacturer and date of manufacture; Labeling provides information about the precautions (Eg., Usage of set/device, packaging conditions); Labeling includes an expiry date where required by country or regional regulation
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Use of expired device | Bacteremia; Moderate Pain | 3 | 2 | N/A | Labeling provides information about legal manufacturer and date of manufacture; Labeling provides information about the precautions (Eg., Usage of set/device, packaging conditions); Labeling includes an expiry date where required by country or regional regulation
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Use of expired device | Mild Phlebitis; Local reaction/ infection; Mild Pain | 2 | 2 | N/A | Labeling provides information about legal manufacturer and date of manufacture; Labeling provides information about the precautions (Eg., Usage of set/device, packaging conditions); Labeling includes an expiry date where required by country or regional regulation
Biological contaminants | Device Contamination with Biological Material | Exposure to biological contaminants infused via parenteral route into the body | Use of expired device | Temporary discomfort | 1 | 2 | N/A | Labeling provides information about legal manufacturer and date of manufacture; Labeling provides information about the precautions (Eg., Usage of set/device, packaging conditions); Labeling includes an expiry date where required by country or regional regulation
Biological contaminants | Device Contamination with Biological Material | Skin / Mucous membrane exposure to infectious agent | Patient movement causes set/device to disconnect at a connection point | Death Cardiopulmonary Arrest | 5 | 1 | Device designed easily attach/detach to connecting components; Device with rotating collars must meet rotating collar luer retention force per ISO 80369-7 requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | N/A
Biological contaminants | Device Contamination with Biological Material | Skin / Mucous membrane exposure to infectious agent | Patient movement causes set/device to disconnect at a connection point | Septic shock Organ injury | 4 | 1 | Device designed easily attach/detach to connecting components; Device with rotating collars must meet rotating collar luer retention force per ISO 80369-7 requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | N/A
Biological contaminants | Device Contamination with Biological Material | Skin / Mucous membrane exposure to infectious agent | Patient movement causes set/device to disconnect at a connection point | Bacteremia Cellulitis, Moderate Organ Inflammation (reversible) | 3 | 2 | Device designed easily attach/detach to connecting components; Device with rotating collars must meet rotating collar luer retention force per ISO 80369-7 requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | N/A
Biological contaminants | Device Contamination with Biological Material | Skin / Mucous membrane exposure to infectious agent | Patient movement causes set/device to disconnect at a connection point | Cellulitis, mild | 2 | 2 | Device designed easily attach/detach to connecting components; Device with rotating collars must meet rotating collar luer retention force per ISO 80369-7 requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | N/A
Biological contaminants | Device Contamination with Biological Material | Skin / Mucous membrane exposure to infectious agent | Patient movement causes set/device to disconnect at a connection point | Temporary discomfort | 1 | 1 | Device designed easily attach/detach to connecting components; Device with rotating collars must meet rotating collar luer retention force per ISO 80369-7 requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | N/A
Biological contaminants | Device Contamination with Biological Material | Skin / Mucous membrane exposure to infectious agent | Needle sticks | Death Cardiopulmonary Arrest | 5 | 1 | N/A | Needleless Connector Piggyback Adapter labeling provide information NOT to use on pre-slit needleless access ports; Labeling provide instructions not to use needles, blunt cannulas or luer caps on luer activated injection access sites (Eg., Needlefree connectors); Labeling includes instructions for using the device
Biological contaminants | Device Contamination with Biological Material | Skin / Mucous membrane exposure to infectious agent | Needle sticks | Septic shock Organ injury | 4 | 1 | N/A | Needleless Connector Piggyback Adapter labeling provide information NOT to use on pre-slit needleless access ports; Labeling provide instructions not to use needles, blunt cannulas or luer caps on luer activated injection access sites (Eg., Needlefree connectors); Labeling includes instructions for using the device
Biological contaminants | Device Contamination with Biological Material | Skin / Mucous membrane exposure to infectious agent | Needle sticks | Bacteremia Cellulitis, Moderate Organ Inflammation (reversible) | 3 | 2 | N/A | Needleless Connector Piggyback Adapter labeling provide information NOT to use on pre-slit needleless access ports; Labeling provide instructions not to use needles, blunt cannulas or luer caps on luer activated injection access sites (Eg., Needlefree connectors); Labeling includes instructions for using the device
Biological contaminants | Device Contamination with Biological Material | Skin / Mucous membrane exposure to infectious agent | Needle sticks | Cellulitis, mild | 2 | 2 | N/A | Needleless Connector Piggyback Adapter labeling provide information NOT to use on pre-slit needleless access ports; Labeling provide instructions not to use needles, blunt cannulas or luer caps on luer activated injection access sites (Eg., Needlefree connectors); Labeling includes instructions for using the device
Biological contaminants | Device Contamination with Biological Material | Skin / Mucous membrane exposure to infectious agent | Needle sticks | Temporary discomfort | 1 | 1 | N/A | Needleless Connector Piggyback Adapter labeling provide information NOT to use on pre-slit needleless access ports; Labeling provide instructions not to use needles, blunt cannulas or luer caps on luer activated injection access sites (Eg., Needlefree connectors); Labeling includes instructions for using the device
Biological contaminants | Device Contamination with Biological Material | Skin / Mucous membrane exposure to infectious agent | Sterile field is compromised | Death Cardiopulmonary Arrest | 5 | 1 | Components must not be cracked, broken, distorted, or severely discolored; Devices should meet the storage and transport requirements; Package labelling and printing must be intact and legible to point of use; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Seal strength shall be sufficient to prevent the package from opening prior to point of use | Administration set primary package labeling for sterile content sets indicate that the set is sterile and nonpyrogenic in intact packaging and shall not be re-sterilized
Biological contaminants | Device Contamination with Biological Material | Skin / Mucous membrane exposure to infectious agent | Sterile field is compromised | Septic shock Organ injury | 4 | 1 | Components must not be cracked, broken, distorted, or severely discolored; Devices should meet the storage and transport requirements; Package labelling and printing must be intact and legible to point of use; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Seal strength shall be sufficient to prevent the package from opening prior to point of use | Administration set primary package labeling for sterile content sets indicate that the set is sterile and nonpyrogenic in intact packaging and shall not be re-sterilized
Biological contaminants | Device Contamination with Biological Material | Skin / Mucous membrane exposure to infectious agent | Sterile field is compromised | Bacteremia Cellulitis, Moderate Organ Inflammation (reversible) | 3 | 2 | Components must not be cracked, broken, distorted, or severely discolored; Devices should meet the storage and transport requirements; Package labelling and printing must be intact and legible to point of use; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Seal strength shall be sufficient to prevent the package from opening prior to point of use | Administration set primary package labeling for sterile content sets indicate that the set is sterile and nonpyrogenic in intact packaging and shall not be re-sterilized
Biological contaminants | Device Contamination with Biological Material | Skin / Mucous membrane exposure to infectious agent | Sterile field is compromised | Cellulitis, mild | 2 | 2 | Components must not be cracked, broken, distorted, or severely discolored; Devices should meet the storage and transport requirements; Package labelling and printing must be intact and legible to point of use; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Seal strength shall be sufficient to prevent the package from opening prior to point of use | Administration set primary package labeling for sterile content sets indicate that the set is sterile and nonpyrogenic in intact packaging and shall not be re-sterilized
Biological contaminants | Device Contamination with Biological Material | Skin / Mucous membrane exposure to infectious agent | Sterile field is compromised | Temporary discomfort | 1 | 1 | Components must not be cracked, broken, distorted, or severely discolored; Devices should meet the storage and transport requirements; Package labelling and printing must be intact and legible to point of use; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Seal strength shall be sufficient to prevent the package from opening prior to point of use | Administration set primary package labeling for sterile content sets indicate that the set is sterile and nonpyrogenic in intact packaging and shall not be re-sterilized
Biological contaminants | Device Contamination with Biological Material | Skin / Mucous membrane exposure to infectious agent | Tubing rupture due to high pressure | Death Cardiopulmonary Arrest | 5 | 1 | Device designed to function after pressure infusion requirements | Neutron/Neutron Ext Set package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 350psig or 10mL per second of contrast media; Administration set primary package labeling or instructions for use for extension sets or administration sets with distal injection access sites, excluding sets intended to be used with high pressure infusion equipment shall include a caution statement that set is not for high pressure infusion; Administration set primary package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 400 psig; Administration set with a MicroClave primary package labeling or instructions for use for high pressure extension sets contain a warning statement that high pressure infusion is conducted through the MicroClave in the middle of the assembly and the upstream clamp is engaged to isolate other upstream components from high pressures
Biological contaminants | Device Contamination with Biological Material | Skin / Mucous membrane exposure to infectious agent | Tubing rupture due to high pressure | Septic shock Organ injury | 4 | 1 | Device designed to function after pressure infusion requirements | Neutron/Neutron Ext Set package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 350psig or 10mL per second of contrast media; Administration set primary package labeling or instructions for use for extension sets or administration sets with distal injection access sites, excluding sets intended to be used with high pressure infusion equipment shall include a caution statement that set is not for high pressure infusion; Administration set primary package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 400 psig; Administration set with a MicroClave primary package labeling or instructions for use for high pressure extension sets contain a warning statement that high pressure infusion is conducted through the MicroClave in the middle of the assembly and the upstream clamp is engaged to isolate other upstream components from high pressures
Biological contaminants | Device Contamination with Biological Material | Skin / Mucous membrane exposure to infectious agent | Tubing rupture due to high pressure | Bacteremia Cellulitis, Moderate Organ Inflammation (reversible) | 3 | 2 | Device designed to function after pressure infusion requirements | Neutron/Neutron Ext Set package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 350psig or 10mL per second of contrast media; Administration set primary package labeling or instructions for use for extension sets or administration sets with distal injection access sites, excluding sets intended to be used with high pressure infusion equipment shall include a caution statement that set is not for high pressure infusion; Administration set primary package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 400 psig; Administration set with a MicroClave primary package labeling or instructions for use for high pressure extension sets contain a warning statement that high pressure infusion is conducted through the MicroClave in the middle of the assembly and the upstream clamp is engaged to isolate other upstream components from high pressures
Biological contaminants | Device Contamination with Biological Material | Skin / Mucous membrane exposure to infectious agent | Tubing rupture due to high pressure | Cellulitis, mild | 2 | 2 | Device designed to function after pressure infusion requirements | Neutron/Neutron Ext Set package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 350psig or 10mL per second of contrast media; Administration set primary package labeling or instructions for use for extension sets or administration sets with distal injection access sites, excluding sets intended to be used with high pressure infusion equipment shall include a caution statement that set is not for high pressure infusion; Administration set primary package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 400 psig; Administration set with a MicroClave primary package labeling or instructions for use for high pressure extension sets contain a warning statement that high pressure infusion is conducted through the MicroClave in the middle of the assembly and the upstream clamp is engaged to isolate other upstream components from high pressures
Biological contaminants | Device Contamination with Biological Material | Skin / Mucous membrane exposure to infectious agent | Tubing rupture due to high pressure | Temporary discomfort | 1 | 1 | Device designed to function after pressure infusion requirements | Neutron/Neutron Ext Set package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 350psig or 10mL per second of contrast media; Administration set primary package labeling or instructions for use for extension sets or administration sets with distal injection access sites, excluding sets intended to be used with high pressure infusion equipment shall include a caution statement that set is not for high pressure infusion; Administration set primary package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 400 psig; Administration set with a MicroClave primary package labeling or instructions for use for high pressure extension sets contain a warning statement that high pressure infusion is conducted through the MicroClave in the middle of the assembly and the upstream clamp is engaged to isolate other upstream components from high pressures
Biological contaminants | Device Contamination with Biological Material | Skin / Mucous membrane exposure to infectious agent | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Death; Cardiopulmonary Arrest | 5 | 1 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Biological contaminants | Device Contamination with Biological Material | Skin / Mucous membrane exposure to infectious agent | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Septic; shock; Organ injury | 4 | 1 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Biological contaminants | Device Contamination with Biological Material | Skin / Mucous membrane exposure to infectious agent | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Bacteremia; Moderate Cellulitis; Organ Inflammation (reversible) | 3 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Biological contaminants | Device Contamination with Biological Material | Skin / Mucous membrane exposure to infectious agent | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | mild Cellulitis; | 2 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Biological contaminants | Device Contamination with Biological Material | Skin / Mucous membrane exposure to infectious agent | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Temporary discomfort | 1 | 1 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Biological contaminants | Device Contamination with Biological Material | Skin / Mucous membrane exposure to infectious agent | Proper disposal procedure not followed | Death; Cardiopulmonary Arrest | 5 | 1 | Device must meet environmental requirements | N/A
Biological contaminants | Device Contamination with Biological Material | Skin / Mucous membrane exposure to infectious agent | Proper disposal procedure not followed | Septic; shock; Organ injury | 4 | 1 | Device must meet environmental requirements | N/A
Biological contaminants | Device Contamination with Biological Material | Skin / Mucous membrane exposure to infectious agent | Proper disposal procedure not followed | Bacteremia; Moderate Cellulitis; Organ Inflammation (reversible) | 3 | 2 | Device must meet environmental requirements | N/A
Biological contaminants | Device Contamination with Biological Material | Skin / Mucous membrane exposure to infectious agent | Proper disposal procedure not followed | mild Cellulitis; | 2 | 2 | Device must meet environmental requirements | N/A
Biological contaminants | Device Contamination with Biological Material | Skin / Mucous membrane exposure to infectious agent | Proper disposal procedure not followed | Temporary discomfort | 1 | 1 | Device must meet environmental requirements | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Systemic infection from virus, bacteria, or other infectious agents | 3 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Mild Cellulitis | 2 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Temporary discomfort | 1 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Systemic infection from virus, bacteria, or other infectious agents | 3 | 2 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements;Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Mild Cellulitis | 2 | 2 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements;Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Temporary discomfort | 1 | 2 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements;Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Systemic infection from virus, bacteria, or other infectious agents | 3 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Mild Cellulitis | 2 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Temporary discomfort | 1 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Patient movement causes set/device to disconnect at a connection point | Systemic infection from virus, bacteria, or other infectious agents | 3 | 2 | Device designed easily attach/detach to connecting components; Device with rotating collars must meet rotating collar luer retention force per ISO 80369-7 requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Patient movement causes set/device to disconnect at a connection point | Mild Cellulitis | 2 | 2 | Device designed easily attach/detach to connecting components; Device with rotating collars must meet rotating collar luer retention force per ISO 80369-7 requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Patient movement causes set/device to disconnect at a connection point | Temporary discomfort | 1 | 2 | Device designed easily attach/detach to connecting components; Device with rotating collars must meet rotating collar luer retention force per ISO 80369-7 requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | N/A
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Silicone sleeve stuck down after disconnecting mating component | Systemic infection from virus, bacteria, or other infectious agents | 3 | 2 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Silicone sleeve stuck down after disconnecting mating component | Mild Cellulitis | 2 | 2 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Silicone sleeve stuck down after disconnecting mating component | Temporary discomfort | 1 | 2 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Tubing rupture due to high pressure | Systemic infection from virus, bacteria, or other infectious agents | 3 | 2 | Device designed to function after pressure infusion requirements | Neutron/Neutron Ext Set package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 350psig or 10mL per second of contrast media; Administration set primary package labeling or instructions for use for extension sets or administration sets with distal injection access sites, excluding sets intended to be used with high pressure infusion equipment shall include a caution statement that set is not for high pressure infusion; Administration set primary package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 400 psig; Administration set with a MicroClave primary package labeling or instructions for use for high pressure extension sets contain a warning statement that high pressure infusion is conducted through the MicroClave in the middle of the assembly and the upstream clamp is engaged to isolate other upstream components from high pressures
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Tubing rupture due to high pressure | Mild Cellulitis | 2 | 2 | Device designed to function after pressure infusion requirements | Neutron/Neutron Ext Set package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 350psig or 10mL per second of contrast media; Administration set primary package labeling or instructions for use for extension sets or administration sets with distal injection access sites, excluding sets intended to be used with high pressure infusion equipment shall include a caution statement that set is not for high pressure infusion; Administration set primary package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 400 psig; Administration set with a MicroClave primary package labeling or instructions for use for high pressure extension sets contain a warning statement that high pressure infusion is conducted through the MicroClave in the middle of the assembly and the upstream clamp is engaged to isolate other upstream components from high pressures
Biological contaminants | Device Contamination with Biological Material | Exposure to patient blood caused by fluid leaks | Tubing rupture due to high pressure | Temporary discomfort | 1 | 2 | Device designed to function after pressure infusion requirements | Neutron/Neutron Ext Set package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 350psig or 10mL per second of contrast media; Administration set primary package labeling or instructions for use for extension sets or administration sets with distal injection access sites, excluding sets intended to be used with high pressure infusion equipment shall include a caution statement that set is not for high pressure infusion; Administration set primary package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 400 psig; Administration set with a MicroClave primary package labeling or instructions for use for high pressure extension sets contain a warning statement that high pressure infusion is conducted through the MicroClave in the middle of the assembly and the upstream clamp is engaged to isolate other upstream components from high pressures
Biologic - Bacteria, virus, other agent | Device Contamination with Biological Material | Exposure to blood products caused by fluid leaks | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Infections from viral, bacteria, protozoan, and other infectious agents | 3 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Biologic - Bacteria, virus, other agent | Device Contamination with Biological Material | Exposure to blood products caused by fluid leaks | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Infections from viral, bacteria, protozoan, and other infectious agents | 3 | 2 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Biologic - Bacteria, virus, other agent | Device Contamination with Biological Material | Exposure to blood products caused by fluid leaks | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Infections from viral, bacteria, protozoan, and other infectious agents | 3 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Biologic - Bacteria, virus, other agent | Device Contamination with Biological Material | Exposure to blood products caused by fluid leaks | Silicone sleeve stuck down after disconnecting mating component | Infections from viral, bacteria, protozoan, and other infectious agents | 3 | 2 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
Ergonomic stress | Inadequate User Interface | Exposure to ergonomic stress | Difficult / Unable to connect / disconnect the set/device with other mating devices (eg Container, access site) | Temporary discomfort | 1 | 1 | Components designed to turn easily; Device designed easily attach/detach to connecting components; Device designed not to leak during normal product usage; Device should meet assembly integrity test requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device
Ergonomic stress | Inadequate User Interface | Exposure to ergonomic stress | Difficult / Unable to open the packaging | Temporary discomfort | 1 | 1 | Seal strength shall be sufficient to prevent the package from opening prior to point of use | Labeling includes instructions for using the device
Ergonomic stress | Inadequate User Interface | Exposure to ergonomic stress | Difficult / Unable to remove Cap / cover / hood from the set/device | Temporary discomfort | 1 | 1 | Protective caps should be secure but easily removable | N/A
Fluid leakage | Fluid Leak | Slips/Falls | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Death; Cardiopulmonary Arrest | 5 | 1 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Fluid leakage | Fluid Leak | Slips/Falls | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Head Trauma with CNS; injury | 4 | 1 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Fluid leakage | Fluid Leak | Slips/Falls | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Fracture, Sprain, Strain, Laceration; Moderate Pain | 3 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Fluid leakage | Fluid Leak | Slips/Falls | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Bruise; Scrape; skin Abrasion | 2 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Fluid leakage | Fluid Leak | Slips/Falls | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Temporary discomfort | 1 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Fluid leakage | Fluid Leak | Slips/Falls | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Death; Cardiopulmonary Arrest | 5 | 1 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Fluid leakage | Fluid Leak | Slips/Falls | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Head Trauma with CNS; injury | 4 | 1 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Fluid leakage | Fluid Leak | Slips/Falls | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Fracture, Sprain, Strain, Laceration; Moderate Pain | 3 | 2 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Fluid leakage | Fluid Leak | Slips/Falls | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Bruise; Scrape; skin Abrasion | 2 | 2 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Fluid leakage | Fluid Leak | Slips/Falls | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Temporary discomfort | 1 | 2 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Fluid leakage | Fluid Leak | Slips/Falls | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Death; Cardiopulmonary Arrest | 5 | 1 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Fluid leakage | Fluid Leak | Slips/Falls | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Head Trauma with CNS; injury | 4 | 1 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Fluid leakage | Fluid Leak | Slips/Falls | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Fracture, Sprain, Strain, Laceration; Moderate Pain | 3 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Fluid leakage | Fluid Leak | Slips/Falls | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Bruise; Scrape; skin Abrasion | 2 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Fluid leakage | Fluid Leak | Slips/Falls | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Temporary discomfort | 1 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Fluid leakage | Fluid Leak | Slips/Falls | Patient movement causes set/device to disconnect at a connection point | Death; Cardiopulmonary Arrest | 5 | 1 | Device designed easily attach/detach to connecting components; Device with rotating collars must meet rotating collar luer retention force per ISO 80369-7 requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | N/A
Fluid leakage | Fluid Leak | Slips/Falls | Patient movement causes set/device to disconnect at a connection point | Head Trauma with CNS; injury | 4 | 1 | Device designed easily attach/detach to connecting components; Device with rotating collars must meet rotating collar luer retention force per ISO 80369-7 requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | N/A
Fluid leakage | Fluid Leak | Slips/Falls | Patient movement causes set/device to disconnect at a connection point | Fracture, Sprain, Strain, Laceration; Moderate Pain | 3 | 2 | Device designed easily attach/detach to connecting components; Device with rotating collars must meet rotating collar luer retention force per ISO 80369-7 requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | N/A
Fluid leakage | Fluid Leak | Slips/Falls | Patient movement causes set/device to disconnect at a connection point | Bruise; Scrape; skin Abrasion | 2 | 2 | Device designed easily attach/detach to connecting components; Device with rotating collars must meet rotating collar luer retention force per ISO 80369-7 requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | N/A
Fluid leakage | Fluid Leak | Slips/Falls | Patient movement causes set/device to disconnect at a connection point | Temporary discomfort | 1 | 2 | Device designed easily attach/detach to connecting components; Device with rotating collars must meet rotating collar luer retention force per ISO 80369-7 requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | N/A
Fluid leakage | Fluid Leak | Slips/Falls | Reuse of the same device/set on multiple patients | Death; Cardiopulmonary Arrest | 5 | 1 | N/A | Labeling includes instructions for using the device; Labeling provides information about the precautions (Eg., Usage of set/device, packaging conditions); Labeling provide instructions not reuse set/device which are labeled for single use only
Fluid leakage | Fluid Leak | Slips/Falls | Reuse of the same device/set on multiple patients | Head Trauma with CNS; injury | 4 | 1 | N/A | Labeling includes instructions for using the device; Labeling provides information about the precautions (Eg., Usage of set/device, packaging conditions); Labeling provide instructions not reuse set/device which are labeled for single use only
Fluid leakage | Fluid Leak | Slips/Falls | Reuse of the same device/set on multiple patients | Fracture, Sprain, Strain, Laceration; Moderate Pain | 3 | 2 | N/A | Labeling includes instructions for using the device; Labeling provides information about the precautions (Eg., Usage of set/device, packaging conditions); Labeling provide instructions not reuse set/device which are labeled for single use only
Fluid leakage | Fluid Leak | Slips/Falls | Reuse of the same device/set on multiple patients | Bruise; Scrape; skin Abrasion | 2 | 2 | N/A | Labeling includes instructions for using the device; Labeling provides information about the precautions (Eg., Usage of set/device, packaging conditions); Labeling provide instructions not reuse set/device which are labeled for single use only
Fluid leakage | Fluid Leak | Slips/Falls | Reuse of the same device/set on multiple patients | Temporary discomfort | 1 | 2 | N/A | Labeling includes instructions for using the device; Labeling provides information about the precautions (Eg., Usage of set/device, packaging conditions); Labeling provide instructions not reuse set/device which are labeled for single use only
Fluid leakage | Fluid Leak | Slips/Falls | Silicone sleeve stuck down after disconnecting mating component | Death; Cardiopulmonary Arrest | 5 | 1 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
Fluid leakage | Fluid Leak | Slips/Falls | Silicone sleeve stuck down after disconnecting mating component | Head Trauma with CNS; injury | 4 | 1 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
Fluid leakage | Fluid Leak | Slips/Falls | Silicone sleeve stuck down after disconnecting mating component | Fracture, Sprain, Strain, Laceration; Moderate Pain | 3 | 2 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
Fluid leakage | Fluid Leak | Slips/Falls | Silicone sleeve stuck down after disconnecting mating component | Bruise; Scrape; skin Abrasion | 2 | 2 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
Fluid leakage | Fluid Leak | Slips/Falls | Silicone sleeve stuck down after disconnecting mating component | Temporary discomfort | 1 | 2 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
Fluid leakage | Fluid Leak | Slips/Falls | Tubing rupture due to high pressure | Death; Cardiopulmonary Arrest | 5 | 1 | Device designed to function after pressure infusion requirements | Neutron/Neutron Ext Set package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 350psig or 10mL per second of contrast media; Administration set primary package labeling or instructions for use for extension sets or administration sets with distal injection access sites, excluding sets intended to be used with high pressure infusion equipment shall include a caution statement that set is not for high pressure infusion; Administration set primary package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 400 psig; Administration set with a MicroClave primary package labeling or instructions for use for high pressure extension sets contain a warning statement that high pressure infusion is conducted through the MicroClave in the middle of the assembly and the upstream clamp is engaged to isolate other upstream components from high pressures
Fluid leakage | Fluid Leak | Slips/Falls | Tubing rupture due to high pressure | Head Trauma with CNS; injury | 4 | 1 | Device designed to function after pressure infusion requirements | Neutron/Neutron Ext Set package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 350psig or 10mL per second of contrast media; Administration set primary package labeling or instructions for use for extension sets or administration sets with distal injection access sites, excluding sets intended to be used with high pressure infusion equipment shall include a caution statement that set is not for high pressure infusion; Administration set primary package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 400 psig; Administration set with a MicroClave primary package labeling or instructions for use for high pressure extension sets contain a warning statement that high pressure infusion is conducted through the MicroClave in the middle of the assembly and the upstream clamp is engaged to isolate other upstream components from high pressures
Fluid leakage | Fluid Leak | Slips/Falls | Tubing rupture due to high pressure | Fracture, Sprain, Strain, Laceration; Moderate Pain | 3 | 2 | Device designed to function after pressure infusion requirements | Neutron/Neutron Ext Set package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 350psig or 10mL per second of contrast media; Administration set primary package labeling or instructions for use for extension sets or administration sets with distal injection access sites, excluding sets intended to be used with high pressure infusion equipment shall include a caution statement that set is not for high pressure infusion; Administration set primary package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 400 psig; Administration set with a MicroClave primary package labeling or instructions for use for high pressure extension sets contain a warning statement that high pressure infusion is conducted through the MicroClave in the middle of the assembly and the upstream clamp is engaged to isolate other upstream components from high pressures
Fluid leakage | Fluid Leak | Slips/Falls | Tubing rupture due to high pressure | Bruise; Scrape; skin Abrasion | 2 | 2 | Device designed to function after pressure infusion requirements | Neutron/Neutron Ext Set package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 350psig or 10mL per second of contrast media; Administration set primary package labeling or instructions for use for extension sets or administration sets with distal injection access sites, excluding sets intended to be used with high pressure infusion equipment shall include a caution statement that set is not for high pressure infusion; Administration set primary package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 400 psig; Administration set with a MicroClave primary package labeling or instructions for use for high pressure extension sets contain a warning statement that high pressure infusion is conducted through the MicroClave in the middle of the assembly and the upstream clamp is engaged to isolate other upstream components from high pressures
Fluid leakage | Fluid Leak | Slips/Falls | Tubing rupture due to high pressure | Temporary discomfort | 1 | 2 | Device designed to function after pressure infusion requirements | Neutron/Neutron Ext Set package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 350psig or 10mL per second of contrast media; Administration set primary package labeling or instructions for use for extension sets or administration sets with distal injection access sites, excluding sets intended to be used with high pressure infusion equipment shall include a caution statement that set is not for high pressure infusion; Administration set primary package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 400 psig; Administration set with a MicroClave primary package labeling or instructions for use for high pressure extension sets contain a warning statement that high pressure infusion is conducted through the MicroClave in the middle of the assembly and the upstream clamp is engaged to isolate other upstream components from high pressures
particulate presented | Particulates | Infusing particulates outside of the vessel | Particulate in the fluid path | Edema Local tissue necrosis Pain, Severe | 4 | 1 | Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device must meet the chemical compatibility requirements; Device designed to maintain sterility requirements; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Environmental contol in manufacturing facility | N/A
particulate presented | Particulates | Infusing particulates outside of the vessel | Particulate in the fluid path | Hematoma Edema Pain, Moderate | 3 | 2 | Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device must meet the chemical compatibility requirements; Device designed to maintain sterility requirements; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Environmental contol in manufacturing facility | N/A
particulate presented | Particulates | Infusing particulates outside of the vessel | Particulate in the fluid path | Burning sensation Edema Pain, Mild | 2 | 2 | Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device must meet the chemical compatibility requirements; Device designed to maintain sterility requirements; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Environmental contol in manufacturing facility | N/A
particulate presented | Particulates | Infusing particulates outside of the vessel | Particulate in the fluid path | Temporary discomfort | 1 | 2 | Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device must meet the chemical compatibility requirements; Device designed to maintain sterility requirements; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Environmental contol in manufacturing facility | N/A
particulate presented | Particulates | Infusing particulates outside of the vessel | Device materials are incompatible to prolonged exposure to disinfectant caps | Edema Local tissue necrosis Pain, Severe | 4 | 1 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
particulate presented | Particulates | Infusing particulates outside of the vessel | Device materials are incompatible to prolonged exposure to disinfectant caps | Hematoma Edema Pain, Moderate | 3 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
particulate presented | Particulates | Infusing particulates outside of the vessel | Device materials are incompatible to prolonged exposure to disinfectant caps | Burning sensation Edema Pain, Mild | 2 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
particulate presented | Particulates | Infusing particulates outside of the vessel | Device materials are incompatible to prolonged exposure to disinfectant caps | Temporary discomfort | 1 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
particulate presented | Particulates | Infusing particulates outside of the vessel | Loss of packaging integrity and caps off allows foreign material to enter into the sterile fluid path | Edema Local tissue necrosis Pain, Severe | 4 | 1 | Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Devices should meet the storage and transport requirements; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Seal strength shall be sufficient to prevent the package from opening prior to point of use; Protective caps should be secure but easily removable | N/A
particulate presented | Particulates | Infusing particulates outside of the vessel | Loss of packaging integrity and caps off allows foreign material to enter into the sterile fluid path | Hematoma Edema Pain, Moderate | 3 | 2 | Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Devices should meet the storage and transport requirements; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Seal strength shall be sufficient to prevent the package from opening prior to point of use; Protective caps should be secure but easily removable | N/A
particulate presented | Particulates | Infusing particulates outside of the vessel | Loss of packaging integrity and caps off allows foreign material to enter into the sterile fluid path | Burning sensation Edema Pain, Mild | 2 | 2 | Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Devices should meet the storage and transport requirements; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Seal strength shall be sufficient to prevent the package from opening prior to point of use; Protective caps should be secure but easily removable | N/A
particulate presented | Particulates | Infusing particulates outside of the vessel | Loss of packaging integrity and caps off allows foreign material to enter into the sterile fluid path | Temporary discomfort | 1 | 2 | Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Devices should meet the storage and transport requirements; Package must be free of holes or openings that would allow contamination to enter; Package shall provide protection to the product and maintain sterility (as applicable) throughout the handling distribution, and storage system; Packaging shall maintain a sterile barrier until expiry date; Seal strength shall be sufficient to prevent the package from opening prior to point of use; Protective caps should be secure but easily removable | N/A
Sharp edge | Material Puncture | Exposure to sharp edge/pinch point/abrasive surface | Device contains sharp corners/edges | Skin Abrasion; Mild Bruising Pain; compromised Skin integrity | 2 | 2 | Components must not be cracked, broken, distorted, or severely discolored | N/A
Sharp edge | Material Puncture | Exposure to sharp edge/pinch point/abrasive surface | Device contains sharp corners/edges | Temporary discomfort | 1 | 2 | Components must not be cracked, broken, distorted, or severely discolored | N/A
Sharp edge | Material Puncture | Exposure to sharp edge/pinch point/abrasive surface | Needle sticks | Skin Abrasion; Mild Bruising Pain; compromised Skin integrity | 2 | 2 | N/A | Needleless Connector Piggyback Adapter labeling provide information NOT to use on pre-slit needleless access ports; Labeling provide instructions not to use needles, blunt cannulas or luer caps on luer activated injection access sites (Eg., Needlefree connectors); Labeling includes instructions for using the device
Sharp edge | Material Puncture | Exposure to sharp edge/pinch point/abrasive surface | Needle sticks | Temporary discomfort | 1 | 2 | N/A | Needleless Connector Piggyback Adapter labeling provide information NOT to use on pre-slit needleless access ports; Labeling provide instructions not to use needles, blunt cannulas or luer caps on luer activated injection access sites (Eg., Needlefree connectors); Labeling includes instructions for using the device
Vein clotted off | Use of Device Issue | Exposure to physical injury at infusion site | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Hematoma; Moderate Pain; Skin irritation; bruising; Mild to Moderate Cellulitis | 3 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Vein clotted off | Use of Device Issue | Exposure to physical injury at infusion site | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Hematoma; Mild Pain | 2 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Vein clotted off | Use of Device Issue | Exposure to physical injury at infusion site | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Temporary discomfort | 1 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Vein clotted off | Use of Device Issue | Exposure to physical injury at infusion site | Fluid path sterility compromised due to broken, loose or unsealed fluid path | Hematoma; Moderate Pain; Skin irritation; bruising; Mild to Moderate Cellulitis | 3 | 2 | Device should meet assembly integrity test requirements; Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Luer connectors are designed to be compliant with ISO standards for connection / disconnection; Device designed easily attach/detach to connecting components; Device designed not to leak during normal product usage; Device designed to meet the lipid/alcohol reactivity requirements; Components must not be cracked, broken, distorted, or severely discolored; Devices should meet the storage and transport requirements | Labeling includes instructions for using the device
Vein clotted off | Use of Device Issue | Exposure to physical injury at infusion site | Fluid path sterility compromised due to broken, loose or unsealed fluid path | Hematoma; Mild Pain | 2 | 2 | Device should meet assembly integrity test requirements; Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Luer connectors are designed to be compliant with ISO standards for connection / disconnection; Device designed easily attach/detach to connecting components; Device designed not to leak during normal product usage; Device designed to meet the lipid/alcohol reactivity requirements; Components must not be cracked, broken, distorted, or severely discolored; Devices should meet the storage and transport requirements | Labeling includes instructions for using the device
Vein clotted off | Use of Device Issue | Exposure to physical injury at infusion site | Fluid path sterility compromised due to broken, loose or unsealed fluid path | Temporary discomfort | 1 | 2 | Device should meet assembly integrity test requirements; Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Luer connectors are designed to be compliant with ISO standards for connection / disconnection; Device designed easily attach/detach to connecting components; Device designed not to leak during normal product usage; Device designed to meet the lipid/alcohol reactivity requirements; Components must not be cracked, broken, distorted, or severely discolored; Devices should meet the storage and transport requirements | Labeling includes instructions for using the device
Damaged Blood Products | Improper Flow or Infusion | Hemolyzed blood delivered to the patient | Red blood cells damaged | Mild to Moderate Anemia; Mild to Moderate Hypotension; Congestive; Heart Failure; Mild to Moderate Hyperkalemia | 3 | 2 | Device fluid path is sized accordingly for the administration and aspiration of blood without significant hemolysis; Devices are biocompatible (Device material are complaint to ISO 10993) | Administration set primary package labeling or instructions for use for transfusion sets include a warning statement that other solutions should never be infused with blood or blood products except for those approved by your institution's policies and procedures
Damaged Blood Products | Improper Flow or Infusion | Hemolyzed blood delivered to the patient | Red blood cells damaged | transient Hematuria | 2 | 2 | Device fluid path is sized accordingly for the administration and aspiration of blood without significant hemolysis; Devices are biocompatible (Device material are complaint to ISO 10993) | Administration set primary package labeling or instructions for use for transfusion sets include a warning statement that other solutions should never be infused with blood or blood products except for those approved by your institution's policies and procedures
Damaged Blood Products | Improper Flow or Infusion | Hemolyzed blood delivered to the patient | Red blood cells damaged | Temporary discomfort | 1 | 1 | Device fluid path is sized accordingly for the administration and aspiration of blood without significant hemolysis; Devices are biocompatible (Device material are complaint to ISO 10993) | Administration set primary package labeling or instructions for use for transfusion sets include a warning statement that other solutions should never be infused with blood or blood products except for those approved by your institution's policies and procedures
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Death; Cardiopulmonary Arrest | 5 | 1 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Pulmonary embolism; Septic shock; Severe Hyperglycemia; Severe Hypoglycemia; Severe Hypertension; Severe Hypotension; Severe Arrhythmia; Severe Pain | 4 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Mild to Moderate Hypotension; Moderate Hyperglycemia; Moderate Hypoglycemia; Moderate Pain | 3 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Mild Pain; Mild Hyperglycemia; Mild Hypoglycemia | 2 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Bad connection / disconnection between container / set / Other mating devices (Eg Loose connection, improper mating) | Temporary discomfort | 1 | 2 | Device designed easily attach/detach to connecting components; Device designed for 96 hours of continuous use; Device designed to withstand stickdown and aspirate blood after 96 activation; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Death; Cardiopulmonary Arrest | 5 | 1 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Pulmonary embolism; Septic shock; Severe Hyperglycemia; Severe Hypoglycemia; Severe Hypertension; Severe Hypotension; Severe Arrhythmia; Severe Pain | 4 | 2 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Mild to Moderate Hypotension; Moderate Hyperglycemia; Moderate Hypoglycemia; Moderate Pain | 3 | 2 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Mild Pain; Mild Hyperglycemia; Mild Hypoglycemia | 2 | 2 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Fluid leak due to defective set/device (Eg Crack, holes, cuts, incorrect solvent bonding) | Temporary discomfort | 1 | 2 | Devices should be free from cracks, splits, fractures; Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should function after 5 years of storage at a temperature between 68°F to 77°F (aging); Device designed not to leak during normal product usage; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand internal back pressure requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Incompatibility of device materials and solutions | Death; Cardiopulmonary Arrest | 5 | 1 | Devices are biocompatible (Device material are complaint to ISO 10993) | N/A
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Incompatibility of device materials and solutions | Pulmonary embolism; Septic shock; Severe Hyperglycemia; Severe Hypoglycemia; Severe Hypertension; Severe Hypotension; Severe Arrhythmia; Severe Pain | 4 | 2 | Devices are biocompatible (Device material are complaint to ISO 10993) | N/A
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Incompatibility of device materials and solutions | Mild to Moderate Hypotension; Moderate Hyperglycemia; Moderate Hypoglycemia; Moderate Pain | 3 | 2 | Devices are biocompatible (Device material are complaint to ISO 10993) | N/A
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Incompatibility of device materials and solutions | Mild Pain; Mild Hyperglycemia; Mild Hypoglycemia | 2 | 2 | Devices are biocompatible (Device material are complaint to ISO 10993) | N/A
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Incompatibility of device materials and solutions | Temporary discomfort | 1 | 2 | Devices are biocompatible (Device material are complaint to ISO 10993) | N/A
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Death; Cardiopulmonary Arrest | 5 | 1 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Pulmonary embolism; Septic shock; Severe Hyperglycemia; Severe Hypoglycemia; Severe Hypertension; Severe Hypotension; Severe Arrhythmia; Severe Pain | 4 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Mild to Moderate Hypotension; Moderate Hyperglycemia; Moderate Hypoglycemia; Moderate Pain | 3 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Mild Pain; Mild Hyperglycemia; Mild Hypoglycemia | 2 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Incompatible device (needle or non-compatible luer) inserted into needleless connector | Temporary discomfort | 1 | 2 | Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | Labeling includes instructions for using the device; Primary package labeling states the CLAVE / MicroCLAVE ISO male luers internal diameter range (Eg., between 0.062” & 0.110” or between 1,55 mm and 2,8 mm)
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Patient movement causes set/device to disconnect at a connection point | Death; Cardiopulmonary Arrest | 5 | 1 | Device designed easily attach/detach to connecting components; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Device with rotating collars must meet rotating collar luer retention force per ISO 80369-7 requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | N/A
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Patient movement causes set/device to disconnect at a connection point | Pulmonary embolism; Septic shock; Severe Hyperglycemia; Severe Hypoglycemia; Severe Hypertension; Severe Hypotension; Severe Arrhythmia; Severe Pain | 4 | 2 | Device designed easily attach/detach to connecting components; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Device with rotating collars must meet rotating collar luer retention force per ISO 80369-7 requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | N/A
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Patient movement causes set/device to disconnect at a connection point | Mild to Moderate Hypotension; Moderate Hyperglycemia; Moderate Hypoglycemia; Moderate Pain | 3 | 2 | Device designed easily attach/detach to connecting components; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Device with rotating collars must meet rotating collar luer retention force per ISO 80369-7 requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | N/A
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Patient movement causes set/device to disconnect at a connection point | Mild Pain; Mild Hyperglycemia; Mild Hypoglycemia | 2 | 2 | Device designed easily attach/detach to connecting components; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Device with rotating collars must meet rotating collar luer retention force per ISO 80369-7 requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | N/A
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Patient movement causes set/device to disconnect at a connection point | Temporary discomfort | 1 | 2 | Device designed easily attach/detach to connecting components; Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Device with rotating collars must meet rotating collar luer retention force per ISO 80369-7 requirements; Luer connectors are designed to be compliant with ISO standards for connection / disconnection | N/A
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Device materials are incompatible to prolonged exposure to disinfectant caps | Death; Cardiopulmonary Arrest | 5 | 1 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Device materials are incompatible to prolonged exposure to disinfectant caps | Pulmonary embolism; Septic shock; Severe Hyperglycemia; Severe Hypoglycemia; Severe Hypertension; Severe Hypotension; Severe Arrhythmia; Severe Pain | 4 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Device materials are incompatible to prolonged exposure to disinfectant caps | Mild to Moderate Hypotension; Moderate Hyperglycemia; Moderate Hypoglycemia; Moderate Pain | 3 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Device materials are incompatible to prolonged exposure to disinfectant caps | Mild Pain; Mild Hyperglycemia; Mild Hypoglycemia | 2 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Device materials are incompatible to prolonged exposure to disinfectant caps | Temporary discomfort | 1 | 2 | Device designed to meet microbial barrier, disinfection compatibility and Antimicrobial Efficacy; Device designed to meet the lipid/alcohol reactivity requirements | N/A
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Silicone sleeve stuck down after disconnecting mating component | Death; Cardiopulmonary Arrest | 5 | 1 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Silicone sleeve stuck down after disconnecting mating component | Pulmonary embolism; Septic shock; Severe Hyperglycemia; Severe Hypoglycemia; Severe Hypertension; Severe Hypotension; Severe Arrhythmia; Severe Pain | 4 | 2 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Silicone sleeve stuck down after disconnecting mating component | Mild to Moderate Hypotension; Moderate Hyperglycemia; Moderate Hypoglycemia; Moderate Pain | 3 | 2 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Silicone sleeve stuck down after disconnecting mating component | Mild Pain; Mild Hyperglycemia; Mild Hypoglycemia | 2 | 2 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Silicone sleeve stuck down after disconnecting mating component | Temporary discomfort | 1 | 2 | Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | Labeling includes instructions for using the device
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Tubing rupture due to high pressure | Death; Cardiopulmonary Arrest | 5 | 1 | Device designed to function after pressure infusion requirements | Neutron/Neutron Ext Set package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 350psig or 10mL per second of contrast media; Neutron/Neutron Ext Set package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 350psig or 10mL per second of contrast media; Administration set primary package labeling or instructions for use for extension sets or administration sets with distal injection access sites, excluding sets intended to be used with high pressure infusion equipment shall include a caution statement that set is not for high pressure infusion; Administration set primary package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 400 psig; Administration set with a MicroClave primary package labeling or instructions for use for high pressure extension sets contain a warning statement that high pressure infusion is conducted through the MicroClave in the middle of the assembly and the upstream clamp is engaged to isolate other upstream components from high pressures
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Tubing rupture due to high pressure | Pulmonary embolism; Septic shock; Severe Hyperglycemia; Severe Hypoglycemia; Severe Hypertension; Severe Hypotension; Severe Arrhythmia; Severe Pain | 4 | 2 | Device designed to function after pressure infusion requirements | Neutron/Neutron Ext Set package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 350psig or 10mL per second of contrast media; Neutron/Neutron Ext Set package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 350psig or 10mL per second of contrast media; Administration set primary package labeling or instructions for use for extension sets or administration sets with distal injection access sites, excluding sets intended to be used with high pressure infusion equipment shall include a caution statement that set is not for high pressure infusion; Administration set primary package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 400 psig; Administration set with a MicroClave primary package labeling or instructions for use for high pressure extension sets contain a warning statement that high pressure infusion is conducted through the MicroClave in the middle of the assembly and the upstream clamp is engaged to isolate other upstream components from high pressures
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Tubing rupture due to high pressure | Mild to Moderate Hypotension; Moderate Hyperglycemia; Moderate Hypoglycemia; Moderate Pain | 3 | 2 | Device designed to function after pressure infusion requirements | Neutron/Neutron Ext Set package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 350psig or 10mL per second of contrast media; Neutron/Neutron Ext Set package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 350psig or 10mL per second of contrast media; Administration set primary package labeling or instructions for use for extension sets or administration sets with distal injection access sites, excluding sets intended to be used with high pressure infusion equipment shall include a caution statement that set is not for high pressure infusion; Administration set primary package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 400 psig; Administration set with a MicroClave primary package labeling or instructions for use for high pressure extension sets contain a warning statement that high pressure infusion is conducted through the MicroClave in the middle of the assembly and the upstream clamp is engaged to isolate other upstream components from high pressures
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Tubing rupture due to high pressure | Mild Pain; Mild Hyperglycemia; Mild Hypoglycemia | 2 | 2 | Device designed to function after pressure infusion requirements | Neutron/Neutron Ext Set package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 350psig or 10mL per second of contrast media; Neutron/Neutron Ext Set package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 350psig or 10mL per second of contrast media; Administration set primary package labeling or instructions for use for extension sets or administration sets with distal injection access sites, excluding sets intended to be used with high pressure infusion equipment shall include a caution statement that set is not for high pressure infusion; Administration set primary package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 400 psig; Administration set with a MicroClave primary package labeling or instructions for use for high pressure extension sets contain a warning statement that high pressure infusion is conducted through the MicroClave in the middle of the assembly and the upstream clamp is engaged to isolate other upstream components from high pressures
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Tubing rupture due to high pressure | Temporary discomfort | 1 | 2 | Device designed to function after pressure infusion requirements | Neutron/Neutron Ext Set package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 350psig or 10mL per second of contrast media; Neutron/Neutron Ext Set package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 350psig or 10mL per second of contrast media; Administration set primary package labeling or instructions for use for extension sets or administration sets with distal injection access sites, excluding sets intended to be used with high pressure infusion equipment shall include a caution statement that set is not for high pressure infusion; Administration set primary package labeling or instructions for use for high pressure extension sets contain a warning statement that the set is pressure rated to 400 psig; Administration set with a MicroClave primary package labeling or instructions for use for high pressure extension sets contain a warning statement that high pressure infusion is conducted through the MicroClave in the middle of the assembly and the upstream clamp is engaged to isolate other upstream components from high pressures
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Too much volume lost through the process of flushing or priming | Death; Cardiopulmonary Arrest | 5 | 1 | Device should not exceed residual volume requirements; Device designed to maintain required flow rates and pressure leak requirements | Administration set primary package labeling includes priming volume information; Labeling includes instructions for using the device
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Too much volume lost through the process of flushing or priming | Pulmonary embolism; Septic shock; Severe Hyperglycemia; Severe Hypoglycemia; Severe Hypertension; Severe Hypotension; Severe Arrhythmia; Severe Pain | 4 | 2 | Device should not exceed residual volume requirements; Device designed to maintain required flow rates and pressure leak requirements | Administration set primary package labeling includes priming volume information; Labeling includes instructions for using the device
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Too much volume lost through the process of flushing or priming | Mild to Moderate Hypotension; Moderate Hyperglycemia; Moderate Hypoglycemia; Moderate Pain | 3 | 2 | Device should not exceed residual volume requirements; Device designed to maintain required flow rates and pressure leak requirements | Administration set primary package labeling includes priming volume information; Labeling includes instructions for using the device
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Too much volume lost through the process of flushing or priming | Mild Pain; Mild Hyperglycemia; Mild Hypoglycemia | 2 | 2 | Device should not exceed residual volume requirements; Device designed to maintain required flow rates and pressure leak requirements | Administration set primary package labeling includes priming volume information; Labeling includes instructions for using the device
Insufficient solution volume | Short Fill | Patient receives less infusion solution/fluid than intended | Too much volume lost through the process of flushing or priming | Temporary discomfort | 1 | 2 | Device should not exceed residual volume requirements; Device designed to maintain required flow rates and pressure leak requirements | Administration set primary package labeling includes priming volume information; Labeling includes instructions for using the device
Undetected occlusion | Insufficient Flow or Underinfusion | Under-delivery to the patient | Blood clot (from blood product, or access site) introduced into the device/set | Death; Cardiopulmonary Arrest | 5 | 1 | Devices should be free from cracks, splits, fractures; Device should not exceed residual volume requirements | Administration set primary package labeling or instructions for use for transfusion sets include a warning statement that other solutions should never be infused with blood or blood products except for those approved by your institution's policies and procedures
Undetected occlusion | Insufficient Flow or Underinfusion | Under-delivery to the patient | Blood clot (from blood product, or access site) introduced into the device/set | Pulmonary embolism; Septic shock; Severe Hyperglycemia; Severe Hypoglycemia; Severe Hypertension; Severe Hypotension; Severe Arrhythmia; Severe Pain | 4 | 1 | Devices should be free from cracks, splits, fractures; Device should not exceed residual volume requirements | Administration set primary package labeling or instructions for use for transfusion sets include a warning statement that other solutions should never be infused with blood or blood products except for those approved by your institution's policies and procedures
Undetected occlusion | Insufficient Flow or Underinfusion | Under-delivery to the patient | Blood clot (from blood product, or access site) introduced into the device/set | Mild to Moderate Arrhythmia; Moderate Hyperglycemia; Moderate Hypoglycemia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Moderate Pain | 3 | 2 | Devices should be free from cracks, splits, fractures; Device should not exceed residual volume requirements | Administration set primary package labeling or instructions for use for transfusion sets include a warning statement that other solutions should never be infused with blood or blood products except for those approved by your institution's policies and procedures
Undetected occlusion | Insufficient Flow or Underinfusion | Under-delivery to the patient | Blood clot (from blood product, or access site) introduced into the device/set | Mild Pain; Mild Hyperglycemia; Mild Hypoglycemia | 2 | 2 | Devices should be free from cracks, splits, fractures; Device should not exceed residual volume requirements | Administration set primary package labeling or instructions for use for transfusion sets include a warning statement that other solutions should never be infused with blood or blood products except for those approved by your institution's policies and procedures
Undetected occlusion | Insufficient Flow or Underinfusion | Under-delivery to the patient | Blood clot (from blood product, or access site) introduced into the device/set | Temporary discomfort | 1 | 2 | Devices should be free from cracks, splits, fractures; Device should not exceed residual volume requirements | Administration set primary package labeling or instructions for use for transfusion sets include a warning statement that other solutions should never be infused with blood or blood products except for those approved by your institution's policies and procedures
Undetected occlusion | Insufficient Flow or Underinfusion | Under-delivery to the patient | Fluid path occlusion | Death; Cardiopulmonary Arrest | 5 | 1 | Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to maintain required flow rates and pressure leak requirements; Device designed to withstand internal back pressure requirements; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | N/A
Undetected occlusion | Insufficient Flow or Underinfusion | Under-delivery to the patient | Fluid path occlusion | Pulmonary embolism; Septic shock; Severe Hyperglycemia; Severe Hypoglycemia; Severe Hypertension; Severe Hypotension; Severe Arrhythmia; Severe Pain | 4 | 1 | Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to maintain required flow rates and pressure leak requirements; Device designed to withstand internal back pressure requirements; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | N/A
Undetected occlusion | Insufficient Flow or Underinfusion | Under-delivery to the patient | Fluid path occlusion | Mild to Moderate Arrhythmia; Moderate Hyperglycemia; Moderate Hypoglycemia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Moderate Pain | 3 | 2 | Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to maintain required flow rates and pressure leak requirements; Device designed to withstand internal back pressure requirements; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | N/A
Undetected occlusion | Insufficient Flow or Underinfusion | Under-delivery to the patient | Fluid path occlusion | Mild Pain; Mild Hyperglycemia; Mild Hypoglycemia | 2 | 2 | Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to maintain required flow rates and pressure leak requirements; Device designed to withstand internal back pressure requirements; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | N/A
Undetected occlusion | Insufficient Flow or Underinfusion | Under-delivery to the patient | Fluid path occlusion | Temporary discomfort | 1 | 2 | Device designed to withstand stickdown and aspirate blood after 96 activation; Device designed for 96 hours of continuous use; Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device designed to maintain required flow rates and pressure leak requirements; Device designed to withstand internal back pressure requirements; Device designed to withstand stickdown and aspirate blood/fluid following extended activation in an elevated temperature environment such as an incubator | N/A
Undetected occlusion | Insufficient Flow or Underinfusion | Under-delivery to the patient | No Flow (Eg Vaccum / air lock) | Death; Cardiopulmonary Arrest | 5 | 1 | Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand stickdown and aspirate blood after 96 activation | N/A
Undetected occlusion | Insufficient Flow or Underinfusion | Under-delivery to the patient | No Flow (Eg Vaccum / air lock) | Pulmonary embolism; Septic shock; Severe Hyperglycemia; Severe Hypoglycemia; Severe Hypertension; Severe Hypotension; Severe Arrhythmia; Severe Pain | 4 | 1 | Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand stickdown and aspirate blood after 96 activation | N/A
Undetected occlusion | Insufficient Flow or Underinfusion | Under-delivery to the patient | No Flow (Eg Vaccum / air lock) | Mild to Moderate Arrhythmia; Moderate Hyperglycemia; Moderate Hypoglycemia; Mild to Moderate Hypertension; Mild to Moderate Hypotension; Moderate Pain | 3 | 2 | Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand stickdown and aspirate blood after 96 activation | N/A
Undetected occlusion | Insufficient Flow or Underinfusion | Under-delivery to the patient | No Flow (Eg Vaccum / air lock) | Mild Pain; Mild Hyperglycemia; Mild Hypoglycemia | 2 | 2 | Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand stickdown and aspirate blood after 96 activation | N/A
Undetected occlusion | Insufficient Flow or Underinfusion | Under-delivery to the patient | No Flow (Eg Vaccum / air lock) | Temporary discomfort | 1 | 2 | Device designed to tolerate multiple activations (96 activations with a 3-cc flush between each activation); Device should compatible with Luer Internal diameters greater than 0.062” and less than 0.110''; Device designed to maintain required flow rates and pressure leak requirements; Device designed to prevent the leak when subject to 36'' head height back pressure; Device designed to prevent the leak when subject to internal negative pressure; Device designed to withstand stickdown and aspirate blood after 96 activation | N/A
Removable component | Detachment of Device Component | User ingests removable component | Removable components left in patient bed or placed into mouth | Death Cardiopulmonary Arrest | 5 | 1 | Protective caps should be secure but easily removable | Labeling includes instructions for using the device
Removable component | Detachment of Device Component | User ingests removable component | Removable components left in patient bed or placed into mouth | Hypoxic organ injury Respiratory Distress-Severe | 4 | 1 | Protective caps should be secure but easily removable | Labeling includes instructions for using the device
Removable component | Detachment of Device Component | User ingests removable component | Removable components left in patient bed or placed into mouth | Esophageal or respiratory airway obstruction Gastric foreign body | 3 | 1 | Protective caps should be secure but easily removable | Labeling includes instructions for using the device
Removable component | Detachment of Device Component | User ingests removable component | Removable components left in patient bed or placed into mouth | Coughing or gagging | 2 | 1 | Protective caps should be secure but easily removable | Labeling includes instructions for using the device
Removable component | Detachment of Device Component | User ingests removable component | Removable components left in patient bed or placed into mouth | Temporary discomfort | 1 | 1 | Protective caps should be secure but easily removable | Labeling includes instructions for using the device