## TOC
| No. | Title | rows | Implementation
Within 7-4-2021 Rev 04 | 1 | Fabian brochure | 51 | Y
| 2 | user need and market requirement | 85 | Y
| 3 | product requirement specification | 748 | Y
| 4 | IFU | 598 | N
| 5 | System architecture | 182 | Y
| 6 | Hardware subsystem(gas, pneumatic) | 85 | Y
| 7 | Hardware subsystem(others) | 200 | N
| 8 | Software subsystem (9+1) | 292+ | N | HFO GUI: 1030
| 9 | Auris (CIR, DIR, Market, clinical_workflow, clinical_HA) | 244 | Y
10-1-2021 Rev 05

## Combined
Index | category | Inputs | Failure_Modes | Hazard | Source
FAB-BRO-1 | brochure | EASE OF USE: The fabian HFO is designed to optimize clinicians‘ and nurses‘ work flow within a single device, reducing the time needed to manage ventilator controls, configurations or different devices, resulting in more time and focused care for the baby. | inappropriate interface design | installation-related problem | Brochure
FAB-BRO-2 | brochure | DEVELOPMENTAL CARE: Particularly silent, the fabian HFO, creates a low-stress and low-noise environment for the undisturbed handling, bonding and kangaroo care of the baby during its development. | noise emitted | audible noise | Brochure
FAB-BRO-3 | brochure | PATIENT SAFETY: The fabian HFO is designed to keep babies safe and protected whenever they are ventilated. This includes the reduced risk of volutrauma in HFO with VG, the reduced risk of contamination with the HFO pulse generated from the inspiratory side and reduced risks of user errors with a user-friendly interface. | inappropriate safety factor | protective measure issue | Brochure
FAB-BRO-4 | brochure | MODULARITY: Future-oriented: in addition to the pre-integrated conventional and triggered ventilation modes, software and hardware modules that may be required in the future are easily integrated and can be configured according to your needs. | fail to integrate | Component Incompatible | Brochure
FAB-BRO-5 | brochure | MOBILITY: Designed for mobility: thanks to its compact size, the fabian HFO allows for transport within the hospital from delivery room to NICU and easy access to the baby at all times at the bedside. | fail to transport | use of device issue | Brochure
FAB-BRO-6 | brochure | COMPATIBILITY & INTERFACE: This flexible system is compatible with all conventional breathing circuits currently available on the market. In addition, the fabian HFO is equipped with several communication interfaces (Acutronic AcuLink, Philips VueLink and CapsuleTech) for compatibility to hospital data management systems | unique compatibility violated | component or accessory incompatibility | Brochure
FAB-BRO-7 | brochure | Ease of use: A clear, bright touch-screen TFT with intuitive commands reduces human error and puts the focus back on the baby | user confused | use of device issue | Brochure
FAB-BRO-8 | brochure | GRAPHICS: See waveforms, numeric data, loops, ventilator status, and more. Clearly displayed and represented. | incorrect message display | incorrect display | Brochure
FAB-BRO-9 | brochure | ONE-TOUCH PRINCIPLE: All operation procedures are carried out using a direct, time-saving and easy-tounderstand one-touch principle. | inappropriate operation | use of device issue | Brochure
FAB-BRO-10 | brochure | Modularity: A true modular system with a variety of ventilation modes for the ventilator configuration you need – more can be added and enabled whenever desired | defective module | Component Incompatible | Brochure
FAB-BRO-11 | brochure | 4-IN-1: All these features and even more in just one device: All neonatal and pediatric conventional ventilation modes, (Volume Guarantee included), Non-Invasive Ventilation modes (NIV), O2 therapy with High Flow Nasal Cannula (HFNC) and unique High Frequency Oscillation (HFO) with Lung Recruitment Maneuver, Volume Guarantee | incompatible mode | Component Incompatible | Brochure
FAB-BRO-12 | brochure | FUTURE-ORIENTED: Neonatal and pediatric ventilation modes (IPPV, SIPPV, SIMV, PSV and SIMV+PSV) with independent adjustment of inspiratory and expiratory flow reduces the work of breathing and inadvertent PEEP. | fail to adjust independently | Imprecision | Brochure
FAB-BRO-13 | brochure | VENTILATION BEYOND LIMITS: nCPAP and DuoPAP modes, for nasal and mask ventilation, with automatic leak compensation, featuring two-level CPAP mode to improve CO2 removal. | inadequate leak compensation | Imprecision | Brochure
FAB-BRO-14 | brochure | Option for in-line etCO2 mainstream for pediatric and sidestream Microstream® for neonatal patients. | inappropriate patient type | Unexpected Therapeutic Results | Brochure
FAB-BRO-15 | brochure | Option for real-time SpO2 monitoring with adjustable Perfusion Index (PI) alarm. | fail to adjust alarm settings | device alarm system issue | Brochure
FAB-BRO-16 | brochure | PRICO - PRedictive Intelligent Control of Oxygenation | oxygen delivery compromised | Use of Incorrect Control Settings | Brochure
FAB-BRO-17 | brochure | Option for F.O.T. (Forced Oscillation Technique) for the assessment of the true lung recruitment | inappropriate technique applied | Improper or Incorrect Procedure or Method | Brochure
FAB-BRO-18 | brochure | New lung function monitoring parameters and graphics for decision-maker. Loops both in conventional and HFO modes. Rapid Shallow Breathing Index (RSBI), C20/C (overdistension index), resistance, compliance, time constant, minute volume and many others, available at a glance in the configurable numeric blocks | incorrect parameter | programming issue | Brochure
FAB-BRO-19 | brochure | Developmental care: Premature babies are fragile. One of their special needs is quiet time for normal growth and development. The fabian HFO is equipped with an electromagnetic expiratory valve, highly precise electronic flow meters and a uniquely designed HFO module, making the fabian HFOs one of the most silent ventilators in the world today – without using silencers | noise over specification | audible noise | Brochure
FAB-BRO-20 | brochure | Patient safety: On the receiving end of our ventilators are the tiniest and most fragile patients imaginable. Keeping them safe and protected whenever they are ventilated is our first priority of all our thoughts | ventilation support compromised | Failure to Cycle | Brochure
FAB-BRO-21 | brochure | Integrated lock function to prevent inadvertent settings | locking mechanism failed | Use of Incorrect Control Settings | Brochure
FAB-BRO-22 | brochure | Easy maintenance with few parts for disinfection and sterilization without any risk of infection inside the ventilator – even during HFO ventilation | failed to be disinfected | failure to disinfect | Brochure
FAB-BRO-23 | brochure | No bacterial filters and therefore no risk of increased resistance during humidified ventilatory therapies. | increased resistance | Physical Resistance | Brochure
FAB-BRO-24 | brochure | Optimal support with reduced Work Of Breathing (iWOB). | reduced breath | use of device issue | Brochure
FAB-BRO-25 | brochure | Reduced user error risk thanks to the self-training ACUTRONIC concept inside | insufficient training | Inadequate Training | Brochure
FAB-BRO-26 | brochure | Built-in backup battery pack for up to 2.5 hours of self-powered operation | low battery | battery issue | Brochure
FAB-BRO-27 | brochure | Designed for mobility: Medical technology should simplify the work flow of clinicians. Interfaces should be familiar and intuitive. They should be consistent across different workstations, thus reducing the potential for error. Data should be entered just once and be available across multiple platforms to support decision-making. | design lack of intuitively implemented | inadequate user interface | Brochure
FAB-BRO-28 | brochure | Ultimately, innovations in ventilator design and clinical information management can improve both patient safety and outcomes, while supporting maximum flexibility. | less flexible | improper device output | Brochure
FAB-BRO-29 | brochure | COMPREHENSIVE DEVICE: Allowing easy access to the baby at all times. No disturbances at the bedside in NICU/PICU from device reshuffles. | unable to access | human-device interface issue | Brochure
FAB-BRO-30 | brochure | LIGHTWEIGHT UNIT: Designed for safe, practical transportation with integrated electronic gas blender and built-in battery. | over weighing | patient-device incompatibility | Brochure
FAB-BRO-31 | brochure | SERVICE MODE: Service mode with advanced diagnostics via USB port, standard Ethernet connection | inadequate diagnostic features | Use of Device Issue | Brochure
FAB-BRO-32 | brochure | COMMUNICATION INTERFACE: Digital interface with hospital information systems, download settings and monitoring data. | inappropriate interface design | inadequate user interface | Brochure
FAB-BRO-33 | brochure | FLEXIBLE & SCALABLE: The system is flexible and can be used with all conventional breathing circuits currently available on the market. | fail to integrate | Accessory Incompatible | Brochure
FAB-BRO-34 | brochure | The new fabian family: The new fabian generation consists of three Gold Club members. Each one represents the benchmark in its class, either in non-invasive, conventional or sophisticated high frequency ventilation. They are configured for your specific needs and can be upgraded with further options. | fail to upgrade software | application program version or upgrade problem | Brochure
FAB-BRO-35 | brochure | fabian HFO: our 4-in-1 device, The latest edition of fabian HFO comes with the HFO fully integrated into the inspiratory port. No additional tubing for HFO needed. Furthermore, the fresh gas port allows for easier use of NO systems. The fresh gas port also function as the nCPAP port, making it even easier to switch modalities. | fail to integrate | detachment of device component | Brochure
FAB-BRO-36 | brochure | fabian +nCPAP evolution: our 3-in-1 device, The fabian +nCPAP evolution is a therapy system encased in a small yet powerful device. Ideal for NICU, PICU and transport applications. | incompatible mode | Component Incompatible | Brochure
FAB-BRO-37 | brochure | fabian Therapy evolution: our 2-in-1 device, The fabian Therapy evolution is the smallest but not the least powerful fabian family member. A true, complete and highly advanced non-invasive ventilator featuring all classic and new NIV modes | incompatible mode | Component Incompatible | Brochure
FAB-BRO-38 | brochure | Finding your device: Non-invasive, conventional or sophisticated high frequency ventilation. No matter which product you choose, you’re getting a device that features the latest technology and is ready to ventilate beyond limits. | lead to new risk | No Fail-Safe Mechanism | Brochure
FAB-BRO-39 | brochure | Designed in Switzerland: The fabian HFO is the result of more than 30 years of experience in design, development and craftsmanship. Powerful components such as long-lasting battery packs, innovative modules like a onepiece expiratory valve and a cleverly reduced assembling concept with easily accessible parts make this device user-friendly. | inappropriate interface design | Human-Device Interface Issue | Brochure
FAB-BRO-40 | brochure | Together with a customized service program, fabian helps to minimize the cost of ownership | inappropriate program selected | Application Program Issue | Brochure
FAB-BRO-41 | brochure | Functionality: The fabian HFO is intended for premature infants, newborns as well as children weighting up to 30 kg. | incorrect patient information | Patient Data Issue | Brochure
FAB-BRO-42 | brochure | The fabian HFO is intended for “in-patient use” in hospitals, medically-used rooms and intrahospital patient transport | inappropriate usage | user used incorrect product for intended use | Brochure
FAB-BRO-43 | brochure | The fabian HFO is an electronically, microprocessor controlled ventilator. | ventilation loss of control | Failure to Cycle | Brochure
FAB-BRO-44 | brochure | The fabian HFO ventilates with excess pressure based on the continuous-flow principle. (Time-cycled, pressure-/volume-limited) | inappropriate pressure setting | pressure issue | Brochure
FAB-BRO-45 | brochure | Oxygen is metered by the integrated air/O2 blender | incorrect readings | imprecision | Brochure
FAB-BRO-46 | brochure | The oxygen concentration is measured internally with a galvanic oxygen sensor. | inaccurate measurement | Incorrect Measurement | Brochure
FAB-BRO-47 | brochure | Ventilation modes: HFOV(high frequency oscillation), IPPV(intermittent positive pressure), SIMV(synchronized intermittent mandatory ventilation), SIPPV(synchronized inermittent positive pressure ventilation), CPAP(cintinuous positive airway pressure), PSV(pressure support ventilation), NIV(non-invasive ventilation, mask), HFNC(O2 therapy, option), PRICO(predictive intelligent controlof oxygenation). | inadequate ventilation mode | use of Device Issue | Brochure
FAB-BRO-48 | brochure | Enhancements: apnea ventilation, backup ventilation, Combined SIMV + PSV ventilation, Flow trigger, Volume trigger, Pressure trigger, Adjustable PSV termination criteria, NIV with leak compensation, NIV Trigger (Synchronized NIV), Volume Guarantee, Volume Limit, Demand flow in CPAP | inappropriate ventilation mode | use of Device Issue | Brochure
FAB-BRO-49 | brochure | Special procedures: Manual inspiration/hold (Sustained Lung Inflation), optional Lung recruitment (sigh) in HFO mode, optional O2 flush, optional PRICO, Forced Oscillation Technique | improper procedure | Improper or Incorrect Procedure or Method | Brochure
FAB-BRO-50 | brochure | Dimension: (W x H x D) 30 cm x 37 cm x 40 cm, Weight approx. 20 kg with HFO module, Diagonal screen size 10.4", TFT color touch-screen with LED backlight, Keypad with 10 buttons | out of dimension | Defective Component | Brochure
FAB-BRO-51 | brochure | Input and output ports: 1 USB port (for SW updates/Log files), 1 RS 232 (9 pin) port for PDMS/HIS, 1 RJ 45 Ethernet port for PDMS/HIS, 1 CO2 sensor connection, 1 SpO2 sensor connection, 1 Video out connector (HDMI), 1 nurse call connector | port mismatched | Component Incompatible | Brochure
FAB-BRO-52 | brochure | Intended Use: The fabian ventilators provide mechanical invasive ventilation, non-invasive ventilation, and high flow oxygen therapy. The fabian HFO and fabian +nCPAP evolution ventilators along with their accessories can assist or replace spontaneous breathing. The fabian Therapy evolution provides NIV support only (face mask). It is indicated for acute or chronic respiratory failure. | inappropriate usage | User used incorrect product for intended use | Brochure
FAB-BRO-53 | brochure | Indications for Use: The fabian ventilator is intended for premature infants, newborns neonates, infant and pediatric patients, as well as children weighing from 300 g to 30 kg | connection issue | User used incorrect product for intended use | Brochure
FAB-BRO-54 | brochure | Medical Purpose and Description of the Device: The fabian ventilator is intended to provide continuous respiratory support for neonatal, infant and pediatric patients via invasive ventilation(into throat), non-invasive ventilation(facial mask), and high flow oxygen therapy. The device is intended for use in the hospital and institutional healthcare environment | inappropriate usage | User used incorrect product for intended use | Brochure
UN001 | clinical | The clinician needs to be able to provide ventilatory support (invasive and non invasive) to infant and pediatric patients between 300 gram and 30 kg | inappropriate user | User used incorrect product for intended use | User_Needs
UN002 | clinical | The hospital needs to be able to use the device for approx. 50 patients per year and 5000 operating hours / year | inadequate capacity | Output Issue | User_Needs
UN003 | clinical | The patient needs ventilation for durations of up to 30 days. | reduced service life | Shelf Life Exceeded | User_Needs
UN004 | clinical | The clinician needs to be able to ventilate patients invasively (endotracheal tube intubation), non-invasively (nasal mask or prongs) or provide high flow oxygen therapy (nasal cannula set) | inadequate therapy modes | Unexpected Therapeutic Results | User_Needs
UN005 | clinical | The clinician needs to be able to use a range of monitoring accessories/devices: ETCO2 module, SpO2 module and sensors, Flow Sensor | incompatible accessory | Accessory Incompatible | User_Needs
UN006 | clinical | The clinician needs to be able to configure and connect Breathing Circuits: Connection of single limb circuit, Connection of dual limb circuit, Installation of exhalation valve membrane/holder, Installation of flow sensor within circuit | loss of connection | fitting problem | User_Needs
UN007 | clinical | The user needs to be able to configure and connect accessories to the ventilator and the breathing circuit: Flow sensor, ETCO2 sampling lines, Nebulizer, Heated humidifiers | fail to support accessory | Accessory Incompatible | User_Needs
UN008 | clinical | The user needs to be able to connect the ventilator to supplies of power and medical gas: Connection to wall AC power, Connection to pipeline or cylinder air and oxygen | fail to supply power | loose or intermittent connection | User_Needs
UN009 | clinical | The user needs to be able to power an electronic nebulizer from the ventilator | fail to power up | device difficult to setup or prepare | User_Needs
UN010 | clinical | The user needs to be able to connect the ventilator RS232 port to a patient data monitoring system | unable to connect | loose or intermittent connection | User_Needs
UN011 | clinical | The user needs to be able to download ventilator log data to a USB memory device | fail to record data | data issue | User_Needs
UN012 | clinical | The user needs to be able to see and interpret information provided by the ventilator using the user interface: LCD screen, Indicator lamps, Alarm banner, Digital monitor values, Graphical data | missing information | inadequate user interface | User_Needs
UN013 | clinical | The user needs to be able to operate the device user interface controls: LCD touchscreen controls, Hard key buttons (including Power on/off button), SW User Interface (menus, dialogs, screens) | unstable interface | inadequate user interface | User_Needs
UN014 | clinical | The user needs to be able to identify and acknowledge alarm indications: Alarm messages on the GUI, Alarm lamp(s), Alarm audible, Alarm audio pause and alarm reset functions | loss of alarm type indication | device alarm system issue | User_Needs
UN015 | clinical | The user needs to be able to transport the ventilator and accessories within a hospital: Moving and manipulation of the cart (handles, brakes), Carrying the ventilator (not in use) when not mounted to a cart | difficult to move | use of device issue | User_Needs
UN016 | clinical | The user needs to be able to position the ventilator and accessories in the care area: Unit movement under normal working loads (accessing user interface/pressing buttons), Positioning of water bag pole, Position of circuit support arm, Interface to a humidifier heater base | difficult to position | difficult to position | User_Needs
UN017 | clinical | The user needs to be able to clean, disinfect, sterilize (as appropriate) the ventilator and any reusable accessories that are subject to potential contamination: Reprocessing between patients for the exhalation valve membrane holder and membrane), Cleaning and disinfection of the ventilator enclosure, flow sensor cable, SpO2 module, EtCO2 module, Reprocessing between patients for the reusable flow sensors. | failed to be disinfected | failure to disinfect | User_Needs
UN018 | clinical | The user wants to be able to re-use the exhalation valve membrane holder for a minimum of 30 cleaning and disinfection cycles. | inappropriate reuse | device disinfection or sterilization issue | User_Needs
UN019 | clinical | The user needs to have validated reprocessing instructions for the exhalation valve membrane holder: Cleaning with pH neutral detergent (ValSure® Cleaner, Steris) at a temperature of 80 to 95 °C for 5 minutes; followed by Steam sterilization at 134 °C for 2 minutes (half cycle). | inadequate instruction | improper or incorrect procedure or method | User_Needs
UN020 | clinical | The user wants to be able to reprocess the ventilator enclosure for a minimum of 400 cleaning cycles. | fail to reach designated cycles | failure to disinfect | User_Needs
UN021 | clinical | The user needs to have validated cleaning instructions for the ventilator enclosure: CaviWipes™ (Metrex); 10% solution of 5.25% sodium hypochlorite bleach (e.g. Clorox), used with lint-free textile cleaning cloths; Hydrogen peroxide 5% (e.g. 35% food grade H2O2, diluted 1:6 with distilled water), used with lint-free textile cleaning cloths. | inadequate instruction | Device Cleaning Issue | User_Needs
UN022 | clinical | Users want to be able to use the ventilator on a shelf, pedestal or pendant | inappropriate user | user used incorrect product for intended use | User_Needs
UN023 | clinical | Users want to be able to use the ventilator on a mobile cart, including ventilating patients during transport between clinical areas within a hospital | overbalancing | unstable | User_Needs
UN024 | clinical | Hospitals want to be able to use single patient use breathing circuits with the ventilator | single use violated | device disinfection or sterilization issue | User_Needs
UN025 | clinical | The user wants the device to be safe in an environment where fluids and other devices are handled. The ventilator should be designed so that material and fluid filled devices can not be placed on top of the device | fluid ingress | Electrical Shorting | User_Needs
UN026 | clinical | Users want to be able to ventilate the patient during outage of the AC power supply, and during transport within the hospital. Run time with the internal battery should be at least two hours | low battery | loss of power | User_Needs
UN027 | clinical | The user wants to be able to measure tidal volume, flow and airway pressure for intubated patients, during conventional and high frequency ventilation, and not impacted by breathing circuit compliance. Note that this implies use of a wye flow sensor and proximal pressure monitoring capability. | inaccurate measurement | incorrect measurement | User_Needs
UN028 | clinical | The ventilator should connect to the hospital gas delivery system or to external gas cylinders when used for in hospital patient transport. The device shall connect to standard approved circuits and specified accessories. | failure of medical gas supply | medical gas supply problem | User_Needs
UN029 | clinical | The user wants to be able to monitor patient oxygen saturation via an external SpO2 monitor, for display on the ventilator user interface and use for closed loop control of ventilator function. | delayed monitoring | inadequate user interface | User_Needs
UN030 | clinical | The user wants to be able to monitor patient EtCO2 via an external main or side stream sampling capnograph, for display on the ventilator user interface | inadequate display message | inadequate user interface | User_Needs
UN031 | clinical | The user of the equipment needs to know how to carry out preventive maintenance and safety inspections, with a maintenance interval of not less than one year | lack of maintenance process | device maintenance issue | User_Needs
UN032 | clinical | Patient population: the ventilators will be used in patients from 300 Gram to 30 kg during treatment on an intensive care | inappropriate user | user used incorrect product for intended use | User_Needs
UN033 | clinical | Patient Size/Body Mass: Neonatal to Pediatric, Ideal Body Mass – 300g to 30kg | infusion or flow issue; connection issue; improper flow or infusion; obstruction within device | user used incorrect product for intended use | User_Needs
UN034 | clinical | Ventilation performance: Required Delivered Volumes (in VG ventilation) - 2-60mL (neonatal), 10-300mL (pediatric), Required Limit Volume (in Pressure ventilation) - 2-150mL (neonatal), 10-500mL (pediatric), Required Baseline Pressures (PEEP) - 0-30cmH2O, Required Pressures (PIP) - 0-80cmH2O | inaccurate volume | volume accuracy issue | User_Needs
UN035 | clinical | Specific Health Conditions: Spontaneously Breathing with respiratory insufficiency, Non-spontaneously Breathing | air delivery compromised | inaccurate delivery | User_Needs
UN036 | clinical | Optional Additional Physical Characteristics: Unable to communicate (neo/infant), Often incubated (neo/infant), Sensitive to sound/noise (neo/infant) | noise emitted | Over-sensing | User_Needs
UN037 | clinical | The ventilator shall provide Closed Loop Oxygen control, automated using external Masimo pulse oximeter | incompatible accessory | Device Inoperable | User_Needs
UN038 | clinical | The HFO model shall be capable to deliver high frequency oscillation based on an oscillating membrane (like Sensormedics 3100A) | fail to deliver | therapeutic or diagnostic output failure | User_Needs
UN039 | clinical | The device shall have an FOT (forced oscillation Technique) to monitor lung compliance with a GUI tool to find the optimal therapy level to protect the lung | inappropriate therapy | Unexpected Therapeutic Results | User_Needs
UN040 | clinical | Neonatal units are relatively small in relation to adult ICU, the device should be small sized (example Babylog 8000) | inappropriate size | user used incorrect product for intended use | User_Needs
UN041 | clinical | The fabian ventilator may be used for transport within a hospital or other institutional healthcare facility when attached to the accessory mobile cart | overbalancing | unstable | User_Needs
UN042 | clinical | The device shall attach to hospital air and oxygen pipeline systems, and/or to medical gas supply cylinder regulated pressure outlets. | fail to compatible with oxygen supply | medical gas supply problem | User_Needs
UN043 | clinical | The device shall connect to standard approved circuits and specified accessories | incompatible connection | fitting problem | User_Needs
UN044 | function | Starting the VENTILATOR from power off | fail to power up | Power Source Issue | User_Needs
UN045 | function | Turning off the ventilator | fail to shut-down device | Failure to Shut Off | User_Needs
UN046 | function | Performing a basic pre-use functional check of the VENTILATOR including the ALARM SYSTEM | fail to schedule check | failure of device to self-test | User_Needs
UN047 | function | Starting ventilation from standby | unable auto ventilation | Failure to Cycle | User_Needs
UN048 | function | Activating standby | fail to exit mode | Failure to Cycle | User_Needs
UN049 | function | Activating maneuvers that help assess lung function and/or the effectiveness of VENTILATOR parameter settings | fail to activate | Difficult or Delayed Activation | User_Needs
UN050 | function | Activating a closed-suctioning function | fail to activate | Difficult or Delayed Activation | User_Needs
UN051 | function | Observing respiratory gas concentrations | incorrect concentration | Incorrect Measurement | User_Needs
UN052 | function | Humidifying/conditioning gases delivered through the VBS | fail to deliver | failure to deliver | User_Needs
UN053 | function | Adding medication to the gas flowing into the PATIENT | fail to nebulization medication | medication error | User_Needs
UN054 | function | Suctioning the PATIENT’S airway | low suction | Decrease in Suction | User_Needs
UN055 | function | X-raying the PATIENT | unclear image | Image Resolution Poor | User_Needs
UN056 | function | Providing alternative means of ventilation with a manual resuscitator | incorrect ventilation mode | no flow | User_Needs
UN057 | function | Positioning the PATIENT | difficult to position | difficult to position | User_Needs
UN058 | function | Transport ventilator via stand / mount on stand / position on table top shelf | inappropriate transportation configuration | device handling issue | User_Needs
UN059 | function | Operate ventilator on battery supply while in transit | low battery | Low Battery | User_Needs
UN060 | function | Clean ventilator | fail to clean adequately | Device Cleaning Issue | User_Needs
UN061 | Usability | Acquire and Configure equipment: Mount and Transport, Configure and Connect Parts | incorrect configuration | device handling issue | User_Needs
UN062 | Usability | Setup and Initiate Ventilation: Starting Ventilator from Power Off, Pre-Use Functional Check including Alarm system, Switching between Ventilation Modes and Breath Types, Setting Control Parameters, Setting Alarm Limits, Inactivating Alarm Signals, Connecting Patient Connection Ports, Disconnecting Patient Connection Ports, switching between different ventilation modes | fail to initiate ventilation | Failure to Cycle | User_Needs
UN063 | Usability | Maneuvers and Actions: Configure and Connect Detachable Part to Ventilator, Activating Maneuvers (FOT PEEP titration) to Assess Lung Function, Humidifying/Conditioning Gases, Adding Medication to Gas using nebulizer, Suctioning Patient Airway (activate oxygen flush; audio pause alarms), X-ray (manual breath for inspiratory hold; standby function for expiratory hold), Activating Standby, Starting Ventilation from Standby | ventilation interrupted | therapeutic or diagnostic output failure | User_Needs
UN064 | Usability | Monitor and Response: Observing Respiratory Gas Concentrations, Observing Monitored Ventilation Parameters, Inactivating Alarm Signals, Operate Ventilator on Battery Supply in Transit, Adjusting Ventilation Control Parameters | delayed response | Communication or Transmission Issue | User_Needs
UN065 | Usability | Shutdown and Stow: Disconnecting Patient Connection Port to Patient Interface, Turning Off Ventilator, Configuring Ventilator of Detachable Parts, Transport to Stow, Clean Ventilator, Clean re-usable parts | incorrect shutdown procedure | device handling issue | User_Needs
UN066 | Regulatory | The device shall comply with REACH requirements for all intended markets. | REACH violated | environmental control or utility issue | User_Needs
UN067 | Regulatory | The device shall comply with RoHS requirements for all intended markets | rohs violated | environmental control or utility issue | User_Needs
UN068 | Regulatory | The device shall comply with the biocompatibility requirements of ISO 10993-1 (materials with human contact) and ISO 18562-1 (gas pathway materials). | material bio-incompatibility | Biocompatibility Issue | User_Needs
UN069 | Regulatory | The device shall operate as intended following storage within the range 0 to 40 °C, 20 to 80 %RH non-condensing. | inappropriate storage environment or condition | use of device issue | User_Needs
UN070 | Regulatory | The device shall operate as intended following shipping transportation within the range -20 to 60 °C, 10 to 95 %RH non-condensing. | inappropriate transportation environment or condition | environmental control or utility issue | User_Needs
UN071 | Regulatory | The device shall have an expected service life not less than 8 years or 40,000 operating hours, provided that maintenance is executed to the defined schedule. | reduced service life | improper device output | User_Needs
UN072 | Regulatory | Labelling on the ventilator and accessories shall meet the requirements of EN IEC 60601-1, EN ISO 80601-2-12, and EN 1041 | incorrect labeling | device markings issue | User_Needs
UN073 | Regulatory | Packaging labels shall meet the requirements of EN IEC 60601-1, EN ISO 80601-2-12, and EN 1041 | inadequate labeling | device markings issue | User_Needs
UN074 | Regulatory | The device shall have a printed IFU | missing ifu in package | inadequate instructions for healthcare professional | User_Needs
UN075 | Regulatory | The IFU should have precautions and warnings to disclose all residual risks identified in the Risk Management File | inadequate IFU | Instruction for Use Issue | User_Needs
UN076 | Regulatory | The IFU shall meet the requirements of EN IEC 60601-1, EN ISO 80601-2-12, and EN 1041 | invalid IFU | Instruction for Use Issue | User_Needs
UN077 | Market | The devices shall use Langugages including English, Chinese, Spanish, French/German, Polish, Thai, Italian, French, Swedish, Russian, Czech, Japanese, Norwegian, Dutch, Turkish, Portugese, Lithuanian, Greek, Lets, Hungarian, Danish | inadequate language | Instruction for Use Issue | User_Needs
UN078 | Regulatory | The devices will have UDI labelling to meet the requirements for EU and global market placement | invalid UDI | device markings issue | User_Needs
UN079 | Regulatory | The devices will be labelled with CE mark | lack of ce mark | device markings issue | User_Needs
UN080 | Regulatory | The devices will be labelled with an NRTL mark | lack of NRTL mark | device markings issue | User_Needs
UN081 | Regulatory | The device shall comply with EN ISO 80601-2-12:2020 | fail to comply with standards or regulation | compliance issue | User_Needs
UN082 | Regulatory | The device shall comply with ISO FDIS 80601-2-87:2020 (HFO mode only) | fail to comply with standards or regulation | compliance issue | User_Needs
UN083 | Regulatory | The device shall comply with EN IEC 60601-1:2020 and required collateral standards | fail to comply with standards or regulation | compliance issue | User_Needs
UN084 | Regulatory | The device shall comply with EN ISO 80601-2-55:2018 (FiO2 and EtCO2 monitoring) | fail to comply with standards or regulation | compliance issue | User_Needs
UN085 | Regulatory | The device shall comply with EN ISO 80601-2-61:2019 (pulse oximetry) | fail to comply with standards or regulation | compliance issue | User_Needs
FAB-H-REQ-18 | Accessories | A cart shall be available to allow the ventilator to be used for intra-facility patient transport. | overbalancing | unstable | PRS
FAB-H-REQ-19 | Accessories | Any cart indicated as appropriate for use with the fabian ventilator shall incorporate handles | handle damaged | break | PRS
FAB-H-REQ-20 | Accessories | Any cart indicated as appropriate for use with the fabian ventilator shall incorporate secure means of attachment for the ventilator. | fail to secure to cart | Migration of Device or Device Component | PRS
FAB-H-REQ-21 | Accessories | The fabian ventilator shall support connection of a third party electrically-powered nebulizer (Aerogen AeroNeb Solo). | incompatible connection | device-device incompatibility | PRS
FAB-H-REQ-22 | Accessories | The nebulizer shall be connected using the USB port | incompatible connection | Detachment of Device or Device Component | PRS
FAB-H-REQ-23 | Accessories | The fabian ventilator shall provide a uniquely-keyed Lemo-type connector port for connection of a third party capnometry device | non unique key | protective measure issue | PRS
FAB-H-REQ-24 | Accessories | The fabian ventilator shall support the following sidestream capnometry device: Oridion MicroPod (Microstream) | fail to support accessory | Device-Device Incompatibility | PRS
FAB-H-REQ-xxx | Accessories | The fabian ventilator shall support the following mainstream capnometry device: Respironics Capnostat | fail to support accessory | Device-Device Incompatibility | PRS
FAB-H-REQ-25 | Accessories | The fabian ventilator shall provide a uniquely-keyed Lemo-type connector port for connection of a third party pulse oximetry device | non unique key | protective measure issue | PRS
FAB-H-REQ-26 | Component | The fabian ventilator shall support the Masimo USB-HPLPM pulse oximetry module. | fail to support accessory | Device-Device Incompatibility | PRS
FAB-H-REQ-194 | breathing circuit | The fabian ventilator shall be qualified for use with specified Vyaire breathing circuit configurations: | incorrect circuit type | Component Incompatible | PRS
FAB-H-REQ-195 | breathing circuit | AirLife dual limb heated breathing circuit, infant/pediatric; Dual limb circuit with heater wires for use with MR850 heater base, with patient wye with ISO 5356-1 connector; This circuit is not indicated for < 4 lpm. | incorrect circuit type | Component Incompatible | PRS
FAB-H-REQ-196 | breathing circuit | AirLife single limb heated breathing circuit for Infant Flow LP | incorrect circuit type | Component Incompatible | PRS
FAB-H-REQ-197 | breathing circuit | AirLife single limb heated breathing circuit for oxygen therapy | incorrect circuit type | Component Incompatible | PRS
FAB-H-REQ-198 | Component | The fabian ventilator shall be qualified for use with generic breathing circuits that meet ISO 5367 and satisfy the following rated specification. | fail to comply with standards or regulation | compliance issue | PRS
FAB-H-REQ-199 | Component | Inspiratory resistance shall be: NEO <= 0.74 mbar / L/min at 2.5 L/min, PED <= 0.40 mbar / L/min at 15 L/min | resistance out of spec | Physical Resistance | PRS
FAB-H-REQ-200 | Component | Expiratory resistance shall be: NEO <= 0.74 mbar / L/min at 2.5 L/min, PED <= 0.40 mbar / L/min at 15 L/min | resistance out of spec | Physical Resistance | PRS
FAB-H-REQ-201 | Component | Compliance shall be NEO <= 2.0 mL / mbar at 60 mbar, PED <= 4.0 mL / mbar at 60 mbar | compliance out of specification | material rigid or stiff | PRS
FAB-H-REQ-202 | Component | leak shall be NEO <= 30 mL / min at 60 mbar, PED <= 40 mL / min at 60 mbar | leakage out of range | leak | PRS
FAB-H-REQ-xxx | Component | Nasal CPAP interfaces need to add Infant Flow LP, Infant Flow, Medin Medijet | inappropriate interface design | inadequate user interface | PRS
FAB-H-REQ-1 | Component | The fabian ventilator shall be qualified for use with the Fisher and Paykel MR 850 heater base, and with the AH290 (F&P compatible) humidification chambers. | incorrect usage | component or accessory incompatibility | PRS
FAB-H-REQ-27 | Component | The fabian ventilator shall be qualified and labelled for use with the following proprietary flow sensors and cables: | inadequate labeling | device markings issue | PRS
FAB-H-REQ-28 | Component | The fabian ventilator shall be qualified and labelled for use with hot wire flow sensor, single patient use, Infant flow sensor (151111 Flow Sensor, SPU, Infant) | invalid sensor | Component Incompatible | PRS
FAB-H-REQ-29 | Component | The fabian ventilator shall be qualified and labelled for use with hot wire flow sensor, single patient use, Paediatric flow sensor (151112 Flow Sensor, SPU, Pediatric) | inadequate labeling | device markings issue | PRS
FAB-H-REQ-30 | Component | The fabian ventilator shall be qualified and labelled for use with hot wire flow sensor cable, re-usable (151100.01 Flow Sensor Cable, reusable) | invalid cable | Component Incompatible | PRS
FAB-H-REQ-533 | Component | The fabian ventilator shall be labelled as compatible with medical air and oxygen cylinders that comply with applicable CGA / ISO standards. | inadequate labeling | device markings issue | PRS
FAB-H-REQ-2 | Component | The fabian ventilator shall be qualified for use with the AirLife Verso closed suction catheters and adapters. | unqualified use | component incompatible | PRS
FAB-H-REQ-3 | Component | The fabian ventilator shall be qualified and labelled for use with the AirLife 001851 non-conductive bacterial/viral breathing system filter at the ventilator outlet | invalid filter | component incompatible | PRS
FAB-H-REQ-4 | Component | The fabian ventilator shall be qualified for use with the iNOvent nitric oxide delivery system. | fail to conjunct with nitro oxide | medical gas supply problem | PRS
FAB-H-REQ-5 | Component | The fabian ventilator shall be provided with gas hoses for air and oxygen complying with ISO 5359. | component missing | Component Missing | PRS
FAB-H-REQ-31 | Modes | The Fabian HFO ventilator shall provide means for the operator to select between a range of modes for invasive and non-invasive ventilation, including nasal pressure controlled ventilation and high flow oxygen therapy. | fail to enter mode | use of device issue | PRS
FAB-H-REQ-xxx | Modes | ventilator modes include Invasive modes (IPPV, SIPPV, SIMV, SIMV+PSV, PSV, CPAP), Non-invasive modes (nCPAP, DuoPAP, High Flow Oxygen Therapy), High Frequency Ventilation (HFO) | inadequate therapy modes | therapeutic or diagnostic output failure | PRS
FAB-H-REQ-xxx | Modes | Conventional ventilation modes shall be provided using a dual limb breathing circuit, with gas flow provided at the INSP outlet. | inappropriate mode configuration | Component Incompatible | PRS
FAB-H-REQ-xxx | Modes | In all conventional modes, mandatory breaths shall have TCPL breath type, with a fixed inspiratory phase flow (I-Flow), and a fixed pressure limit (dependent on mode and trigger type) | inappropriate mode configuration | Use of Incorrect Control Settings | PRS
FAB-H-REQ-203 | Modes | In IPPV ventilation mode, the ventilator shall provide mandatory breaths with properties determined by the ventilator controls, at a fixed rate independent of any patient spontaneous respiratory effort. | inadequate breath type | Unexpected Therapeutic Results | PRS
FAB-H-REQ-204 | Modes | The Rate control shall have accuracy ± 1 bpm for IPPV mode rate control | low rate accuracy | Inaccurate Flowrate | PRS
FAB-H-REQ-205 | Modes | In SIPPV ventilation mode, the ventilator shall provide mandatory breaths with properties determined by the ventilator controls at a rate determined by patient trigger events, or at the set rate, whichever is greater. | -fail to initiate breath -incorrect breath type -inadequate breath type | -inaccurate delivery -output below specifications -excess flow or over-infusion | PRS
FAB-H-REQ-206 | Modes | The Rate control shall have accuracy ± 1 bpm for SIPPV mode breath rate | low rate accuracy | Inaccurate Flowrate | PRS
FAB-H-REQ-207 | Modes | In SIMV ventilation mode, the ventilator shall provide mandatory breaths with properties determined by the ventilator controls at the set rate, synchronized to patient trigger events if they occur within the appropriate triggering window. | incorrect breath type setting | inaccurate delivery | PRS
FAB-H-REQ-208 | Modes | The Rate control shall have accuracy ± 1 bpm for SIMV mode breath rate | low accuracy rate | imprecision | PRS
FAB-H-REQ-209 | Modes | In SIMV+PSV ventilation mode, the ventilator shall provide mandatory breaths with properties determined by the ventilator controls at the set rate, synchronized to patient trigger events where possible. | no breath | Use of Incorrect Control Settings | PRS
FAB-H-REQ-210 | Modes | In SIMV+PSV ventilation mode, any patient trigger events that occur at greater than the set rate shall be supported by a PSV breath. | fail to increase rate | Output Issue | PRS
FAB-H-REQ-211 | Modes | The Rate control (applicable to mandatory breaths) shall have accuracy ± 1 bpm for SIMV+PSV mode breath rate | low rate accuracy | Inaccurate Flowrate | PRS
FAB-H-REQ-32 | Modes | In PSV ventilation mode, any patient trigger events that occur shall be supported by a PSV breath for PSV mode breath delivery | no delivery | delivery system failure | PRS
FAB-H-REQ-33 | Modes | In CPAP ventilation mode, the pressure at the patient airway shall be maintained at the set pressure level. No mandatory or PSV breaths are provided. | fail to hold pressure | pressure issue | PRS
FAB-H-REQ-xxx | Modes | Non-invasive ventilation shall be provided using a single limb breathing circuit and a nasal CPAP variable-flow interface (e.g. Infant Flow), with gas flow to the breathing circuit provided at the FG outlet. This circuit configuration provides for respiratory support by nasal prongs or nasal mask only. | incompatible nasal prongs | Accessory Incompatible | PRS
FAB-H-REQ-xxx | Modes | A control shall be provided to allow the operator to specify the leak compensation mode, from following options: leak Compensation OFF, leak Compensation LOW (<= 10% leak compensation flow), leak Compensation MEDIUM (<= 20% leak compensation flow), leak Compensation HIGH (<= 30% leak compensation flow) | inappropriate leak compensation setting | Use of Incorrect Control Settings | PRS
FAB-H-REQ-216 | Modes | A visual indicator shall be provided during nCPAP or DuoPAP modes to show the selected interface. | invisible indication | no display or display failure | PRS
FAB-H-REQ-534 | Modes | The exhalation valve control shall be set to zero (valve fully open) in single limb ventilation modes (nCPAP, DuoPAP, High Flow) | fail to setup valve correctly | Device Difficult to Setup or Prepare | PRS
FAB-H-REQ-425 | Modes | Nasal CPAP mode shall be available with a single-limb breathing circuit configuration and following nCPAP generator types: Infant Flow LP, Infant Flow, Medin Medijet | inappropriate mode configuration | Component Incompatible | PRS
FAB-H-REQ-426 | Modes | In Nasal CPAP mode, the pressure at the patient airway shall be maintained at the set pressure level (no mandatory or PSV breaths are provided). | fail to hold pressure | pressure issue | PRS
FAB-H-REQ-427 | Modes | DuoPAP mode shall be available with a single-limb breathing circuit configuration and following nCPAP generator types: Infant Flow LP, Infant Flow, Medin Medijet | inappropriate mode configuration | Component Incompatible | PRS
FAB-H-REQ-428 | Modes | In DuoPAP mode, mandatory pressure control breaths (bilevel pressure) shall be provided at a fixed rate independent of any patient spontaneous respiratory effort. | inappropriate pressure level | pressure issue | PRS
FAB-H-REQ-429 | Modes | The Rate control shall have accuracy ± 1 bpm for DuoPAP mode rate | low accuracy rate | Inaccurate Flowrate | PRS
FAB-H-REQ-xxx | Modes | High Flow Oxygen Therapy shall be supported using a single limb breathing circuit. In this mode, the ventilator provides a flow of gas to nasal cannulae, there is negligible pressure provided to the patient airway. | incorrect circuit type | Component Incompatible | PRS
FAB-H-REQ-535 | Modes | High flow mode shall be available with a single-limb breathing circuit configuration. | inappropriate mode configuration | Component Incompatible | PRS
FAB-H-REQ-217 | Modes | In High Flow Oxygen Therapy mode, a continuous flow of oxygenated air shall be provided to the nasal cannulae, at a flowrate set by the operator. | inappropriate flow volume | improper flow or infusion | PRS
FAB-H-REQ-218 | Modes | The Flow control shall have accuracy ± (1 lpm + 10% of setting) for High Flow control accuracy | inaccurate flow rate | Inaccurate Flowrate | PRS
FAB-H-REQ-xxx | Modes | High Frequency Oscillation shall be supported using a dual limb breathing circuit. In this mode, the ventilator provides a constant flow of gas from the INSP outlet to the breathing circuit, while also enabling high frequency pressure delivery using the internal HFO module. | fail to deliver high frequency pressure | improper device output | PRS
FAB-H-REQ-xxx | Modes | In HFO ventilation mode, the pressure at the patient airway shall be oscillating around the set mean pressure level with operator set frequency and pressure amplitude. No mandatory or PSV breaths are provided. | fail to deliver oscillated pressure | pressure issue | PRS
FAB-H-REQ-xxx | Modes | In the HFO mode, the device shall provide the possibility to route the patient gas flow (bias flow) from the FG port than the oscillation is triggered | inappropriate gas flow | infusion or flow issue | PRS
FAB-H-REQ-xxx | Modes | When lung recruitment is enabled, the ventilator shall increase the mean pressure level to a specified pressure level, for a specified amount of time and at a specified frequency. | fail to raise pressure | pressure issue | PRS
FAB-H-REQ-220 | Modes | When Volume Guarantee is enabled, the ventilator shall adjust the inspiratory pressure for mandatory breaths to maintain the expired tidal volume Vte at the set target. | fail to adjust pressure | pressure issue | PRS
FAB-H-REQ-221 | Modes | The V-Guarantee control shall have accuracy ± (1.0 mL + 10% setting) for Volume guarantee volume accuracy | inaccurate volume | incorrect measurement | PRS
FAB-H-REQ-222 | Modes | The inspiratory pressure for mandatory breaths in Volume Guarantee shall be constrained to the range (PEEP + 2 mbar) to set Pmax (with maximum of 80 mbar). | pressure out of range | pressure issue | PRS
FAB-H-REQ-223 | Modes | Control of inspiratory pressure shall be maintained separately for patient-triggered and time-triggered mandatory breath | incorrect trigger | use of incorrect control settings | PRS
FAB-H-REQ-224 | Modes | The inspiratory pressure for mandatory breaths in Volume Guarantee shall be limited to change by no more than ± 3 mbar of previous value from breath to breath (of the same type, patient-triggered or time-triggered). | pressure out of range | pressure issue | PRS
FAB-H-REQ-225 | Modes | Volume Guarantee shall be disabled in the event of a flow sensor disconnection or flow sensor fault alarm. | fail to disable control | Device Remains Activated | PRS
FAB-H-REQ-226 | Modes | During Volume Guarantee ventilation, mandatory breaths shall be terminated if the inspired volume exceeds 130% of the set target volume. | fail to disable breath | excess flow or over-infusion | PRS
FAB-H-REQ-227 | Modes | On deactivation of Volume Guarantee, the set inspiratory pressure shall be returned to the value of the Pinsp control. | fail to deactivate | Device Remains Activated | PRS
FAB-H-REQ-228 | Modes | When Volume Limit is enabled, the ventilator shall terminate inspiratory phase of a mandatory breath when the delivered volume exceeds V-Limit. | fail to disable ventilation | Device Remains Activated | PRS
FAB-H-REQ-229 | Modes | The V-Limit control shall have accuracy ± (1.0 mL + 10% setting) for Volume Limit accuracy | inaccurate volume limit | volume accuracy issue | PRS
FAB-H-REQ-xxx | Modes | A Rise Time control shall be provided to allow capability like 'pressure controlled' ventilation, in which sufficient flow is provided to ensure a specified pressure rise time. | rise time control failed | device operates differently than expected | PRS
FAB-H-REQ-230 | Modes | When Rise Time control is enabled, the flow provided to the outlet shall be determined by the ventilator to ensure that the specified pressure rise time can be achieved. | fail to activate control | Device Inoperable | PRS
FAB-H-REQ-231 | Modes | When Rise-Time is enabled, the ventilator shall adjust the exhalation valve drive using a ramped increase in drive current to enable the pressure to reach the preset target within the time set by the Rise-Time control. | fail to adjust pressure | pressure issue | PRS
FAB-H-REQ-232 | Modes | In absence of spontaneous inspiratory effort, the Rise Time control shall have accuracy ± (0.05 s) for Rise Time accuracy | low rise time accuracy | Incorrect Measurement | PRS
FAB-H-REQ-xxx | Modes | Apnea ventilation shall be available in CPAP and PSV modes. In other modes that include a set rate, the set rate remains active throughout the period of any patient apnea. | inappropriate mode configuration | Use of Incorrect Control Settings | PRS
FAB-H-REQ-233 | Modes | In PSV mode, with the Apnea Alarm set (to a value other than Off), when the limit for apnea detection is exceeded, backup breaths shall be triggered using the set BackupRate. | fail to convert to back-up | Failure to Convert to Back-Up | PRS
FAB-H-REQ-234 | Modes | In PSV mode, with the Apnea Alarm set Off, backup breaths shall be triggered if the patient breath rate falls below the set BackupRate. | fail to convert to back-up | Failure to Convert to Back-Up | PRS
FAB-H-REQ-235 | Modes | In CPAP mode, when the limit for apnea detection is exceeded, from 1 to 5 backup breaths (as determined by the Backup control) shall be triggered using pre-determined settings. | fail to trigger to pre-determined settings | Use of Incorrect Control Settings | PRS
FAB-H-REQ-236 | Modes | The apnea timer shall be reset at the start of inspiration of each detected patient breath. | fail to reset timer | Failure to Reset | PRS
FAB-H-REQ-xxx | Pressure | In all modes intended for a dual limb breathing circuit configuration other than HFO and CPAP, the pressure target between breaths shall be determined by the PEEP control. | incorrect pressure calculation | pressure issue | PRS
FAB-H-REQ-xxx | Pressure | In CPAP, nCPAP and DuoPAP modes the pressure target between breaths shall be determined by the CPAP control. | incorrect pressure calculation | pressure issue | PRS
FAB-H-REQ-xxx | Pressure | In HFO mode the pressure target upon which high frequency oscillations are superimposed shall be determined by the Pmean control. | incorrect pressure calculation | pressure issue | PRS
FAB-H-REQ-237 | Pressure | Set CPAP / PEEP shall have accuracy ± (1 mbar + 3% of setting) for PEEP / CPAP pressure accuracy | low pressure accuracy | pressure issue | PRS
FAB-H-REQ-xxx | Pressure | In all modes intended for a dual limb breathing circuit configuration other than HFO and CPAP, the flow of gas to the breathing circuit between breaths is determined by the E-Flow control. | incorrect flow calculation | improper flow or infusion | PRS
FAB-H-REQ-238 | flow | The set flow during exhalation shall have accuracy ± (0.5 lpm + 10% setting) for Exhalation flow accuracy | low flow accuracy | imprecision | PRS
FAB-H-REQ-239 | oxygen | The volumetric percentage of oxygen in the gas delivered to the patient shall be determined by the O2 setting. | incorrect O2 setting | medical gas supply problem | PRS
FAB-H-REQ-240 | oxygen | The oxygen fraction shall have accuracy ± 3 % (provided that the oxygen sensor is correctly calibrated) for FiO2 accuracy | low oxygen fraction accuracy | medical gas supply problem | PRS
FAB-H-REQ-241 | oxygen | The oxygen fraction shall have accuracy ± 5 % (when the oxygen sensor is not calibrated or disabled) for FiO2 accuracy with non-calibrated / faulty O2 sensor | low oxygen fraction accuracy | medical gas supply problem | PRS
FAB-H-REQ-242 | oxygen | An oxygen flush control shall be provided to enable the volume fraction of oxygen delivered to the patient to be increased by a preset amount, for a maximum period of 120 seconds. | fail to enable oxygen flush | failure to deliver | PRS
FAB-H-REQ-xxx | PRICO | PRICO functions shall be available only if the PRICO license is activated on the device. | fail to activate prico license | failure to deliver | PRS
FAB-H-REQ-xxx | PRICO | PRICO function shall be available in all ventilation modes: IPPV, SIPPV, SIMV, SIMV+PSV, PSV, CPAP, nCPAP, DuoPAP, O2 Therapy, HFO | loss of function | improper device output | PRS
FAB-H-REQ-xxx | PRICO | PRICO shall provide the capability to the user to set the target SpO2 range (as low and high limit values). | fail to set limit correctly | infusion or flow issue | PRS
FAB-H-REQ-xxx | PRICO | The target SpO2 range shall be between 80-100% | fail to set limit correctly | use of incorrect control settings | PRS
FAB-H-REQ-xxx | PRICO | PRICO shall control the FiO2 to maintain the target SpO2 to be in the center of the set range. | fail to maintain value in the center | improper device output | PRS
FAB-H-REQ-xxx | PRICO | PRICO shall provide the capability to the user to set the limits of FiO2 within the PRICO can control. | incorrect limit setting | improper device output | PRS
FAB-H-REQ-xxx | PRICO | The FiO2 limits shall be settable between 21-100%. | fail to set limit correctly | use of incorrect control settings | PRS
FAB-H-REQ-xxx | PRICO | PRICO shall be deactivated and the target FiO2 shall be switched to the set FiO2 (the value of the FiO2 control prior to enabling PRICO) automatically in case any of the following alarm occur: Air supply pressure, Input pressure blender, O2 sensor calibration error, Oxygen sensor defect, Oxygen sensor used up, Oxygen supply pressure, SpO2 module disconnected, SpO2 sensor failure. | fail to deactivate | Device Remains Activated | PRS
FAB-H-REQ-xxx | PRICO | PRICO shall be deactivated and the target FiO2 shall be switched to the set FiO2 automatically if any of the following alarm is active. In case the alarm condition is resolved PRICO shall be re-activated automatically: Check ET tube, Clean flow sensor, Check SpO2 sensor, Flow sensor defect, Flow sensor not connected, Low SIQ, O2 value out of range, Patient disconnected, Tube Occlusion. | fail to activate alarm | device alarm system issue | PRS
FAB-H-REQ-34 | breath trigger | Breaths shall be triggered by time in IPPV mode, and in SIPPV, SIMV, SIMV+PSV and DuoPAP modes when no patient-trigger event occurs within a time window defined by (60 / Rate). | time trigger failed | decreased sensitivity | PRS
FAB-H-REQ-xxx | breath trigger | The Trigger Mode control shall determine the mechanism used to detect patient spontaneous inspiratory effort, as flow trigger, volume trigger or pressure trigger. | inappropriate trigger type | Use of Incorrect Control Settings | PRS
FAB-H-REQ-243 | breath trigger | Breaths shall be triggered by patient effort in SIPPV, SIMV, SIMV+PSV and PSV modes, and in DuoPAP mode when NIV trigger is enabled. | fail to initiate trigger | decreased sensitivity | PRS
FAB-H-REQ-xxx | breath trigger | When Trigger Mode is set to Flow Trigger, the Trigger control shall determine the threshold for inspiratory flow required to initiate a patient-triggered breath. | flow trigger failed | decreased sensitivity | PRS
FAB-H-REQ-430 | breath trigger | Trigger values of 1 to 10 shall represent a flow threshold of 0.12 to 1.20 lpm (neonatal) or 0.6 to 6.0 lpm (pediatric). | invalid trigger value | Use of Incorrect Control Settings | PRS
FAB-H-REQ-244 | breath trigger | When Trigger Mode is set to Pressure Trigger, the Trigger control determines the threshold for inspiratory pressure (as a decrement below baseline pressure) required to initiate a patient-triggered breath. | pressure trigger failed | pressure issue | PRS
FAB-H-REQ-xxx | breath trigger | When Trigger Mode is set to Volume Trigger, the Trigger control determines the threshold for inspiratory volume required to initiate a patient-triggered breath. | volume trigger failed | decreased sensitivity | PRS
FAB-H-REQ-431 | breath trigger | Trigger value of 1 to 10 shall represent thresholds of 5% to 15% of the measured exhaled tidal volume (Vte) for Volume Trigger control range | invalid trigger value | Use of Incorrect Control Settings | PRS
FAB-H-REQ-35 | breath trigger | A manual breath shall be triggered by action of the manual breath control in all modes. | fail to initiate trigger | decreased sensitivity | PRS
FAB-H-REQ-245 | breath type | Mandatory breaths shall have a set inspiratory duration, configured by the operator using the Inspiratory Time control. | inappropriate inspiratory time | Imprecision | PRS
FAB-H-REQ-246 | breath type | The inspiratory duration shall have accuracy ± (0.05 s). | low inspiratory duration accuracy | Imprecision | PRS
FAB-H-REQ-247 | breath type | Mandatory breaths shall be terminated within 200 ms on activation of the High PIP alarm. | long termination time | Failure to Auto Stop | PRS
FAB-H-REQ-xxx | breath type | The default mandatory breath type for all modes intended for a dual-limb breathing circuit shall be TCPL. This applies when the Rise-Time control is disabled. | fail to deactivate control | Device Remains Activated | PRS
FAB-H-REQ-432 | breath type | The ventilator shall provide a fixed flow of oxygenated air to the breathing circuit throughout the duration of the inspiratory phase of a TCPL breath at the set Insp Flow. | fluctuated flow | medical gas supply problem | PRS
FAB-H-REQ-433 | breath type | The set flow shall have accuracy ± (0.5 lpm + 10% setting). | low flow accuracy | improper flow or infusion | PRS
FAB-H-REQ-434 | breath type | With Volume Guarantee disabled, the ventilator shall modulate the exhalation valve force in order to set the opening pressure for the exhalation valve to a set level (P-Insp + PEEP). | fail to open valve | pressure issue | PRS
FAB-H-REQ-435 | breath type | The set pressure shall have accuracy ± (1 cmH2O + 3% of setting). | low pressure accuracy | pressure issue | PRS
FAB-H-REQ-436 | breath type | With Volume Guarantee enabled, the ventilator shall modulate the exhalation valve force in order to set the opening pressure for the exhalation valve to an algorithmically-determined pressure target as specified in section | fail to open valve | pressure issue | PRS
FAB-H-REQ-xxx | breath type | When the Rise Time control is enabled, the mandatory breath type shall be a modified form of TCPL breath, with auto-set inspiratory flow, and ramped control of the exhalation valve pressure target waveform. | fail to control valve | difficult to open or close | PRS
FAB-H-REQ-437 | breath type | When Rise Time control is enabled, the inspiratory flow shall be auto-set by the ventilator and independent of the I-Flow control | fail to auto-set flow | improper flow or infusion | PRS
FAB-H-REQ-438 | breath type | When Rise Time control is enabled, the ventilator shall modulate the exhalation valve force to a target pressure waveform with a ramped increase to the set pressure level over the set Rise Time, followed by a plateau for the remaining I-Time. | fail to modulate valve force | Mechanical Issue | PRS
FAB-H-REQ-xxx | breath type | Ventilator breaths, provided in response to patient trigger events between mandatory breaths in SIMV+PSV mode, and to all patient trigger events in PSV mode, shall be PSV breaths. | patient trigger failed | Use of Device Issue | PRS
FAB-H-REQ-248 | breath type | The ventilator shall provide a fixed flow of oxygenated air to the breathing circuit throughout the duration of the inspiratory phase of a PSV breath, at the set I-Flow value. | fluctuated flow | improper flow or infusion | PRS
FAB-H-REQ-249 | breath type | PSV breaths shall be terminated when the inspiratory flow falls below a set proportion of the peak inspiratory flow, determined by the Termination-PSV control | fail to terminate breath | Device Remains Activated | PRS
FAB-H-REQ-250 | breath type | The Termination-PSV control shall have accuracy ± (15 %). | low control accuracy | imprecision | PRS
FAB-H-REQ-251 | breath type | PSV breaths shall be terminated within 200 ms on activation of the High PIP alarm. | long termination time | Failure to Auto Stop | PRS
FAB-H-REQ-252 | breath type | The ventilator shall modulate the exhalation valve force in order to prevent the airway pressure from exceeding a set level (P-PSV + PEEP), above the baseline pressure. | fail to modulate valve force | Mechanical Issue | PRS
FAB-H-REQ-253 | breath type | The set flow shall have accuracy ± (0.5 lpm + 10% setting). | low flow accuracy | improper flow or infusion | PRS
FAB-H-REQ-254 | breath type | The set pressure shall have accuracy ± (1 cmH2O + 3% of setting). | low pressure accuracy | pressure issue | PRS
FAB-H-REQ-xxx | breath type | The DuoPAP breath type shall be applicable only to a single limb breathing circuit with nCPAP generator. | incorrect breath type | use of device issue | PRS
FAB-H-REQ-255 | breath type | DuoPAP breaths shall have a set inspiratory duration, configured by the operator using the I-Time control. | inappropriate inspiratory time | Use of Incorrect Control Settings | PRS
FAB-H-REQ-256 | breath type | DuoPAP breaths shall be terminated within 200 ms on activation of the High PIP alarm. | long termination time | Failure to Auto Stop | PRS
FAB-H-REQ-257 | breath type | The ventilator shall provide a fixed flow of oxygenated air to the breathing circuit throughout the duration of the inspiratory phase of a DuoPAP breath, pre-determined to achieve the target inspiratory pressure (P-High). | fluctuated flow | improper flow or infusion | PRS
FAB-H-REQ-258 | breath type | The set pressure shall have accuracy ± (1 cmH2O + 3% of setting). | low pressure accuracy | pressure issue | PRS
FAB-H-REQ-259 | breath type | Manual breaths shall be mandatory breaths in all modes, with inspiratory pressure determined by the Pmanual control; and with inspiratory phase continuing as long as the manual breath control is held, or until the maximum set time is reached. | fail to hold mandatory breath | device difficult to maintain | PRS
FAB-H-REQ-260 | breath type | In CPAP mode, manual breaths shall be TCPL breaths with the pressure limit determined by the P-Manual control. | impropriate pressure limit | pressure issue | PRS
FAB-H-REQ-261 | breath type | In nCPAP and DuoPAP modes, the ventilator shall provide a fixed flow of oxygenated air to the breathing circuit throughout the duration of the inspiratory phase of a manual breath, pre-determined to achieve the target pressure (P-Manual in nCPAP mode, P-High in DuoPAP mode) | fluctuated flow | improper flow or infusion | PRS
FAB-H-REQ-262 | breath type | The inspiratory duration of a manual breath shall be determined by the duration of the key-press: the breath continues (to a configured maximum within the range 2 to 30 seconds, see xxx) while the manual breath key remains depressed. | inappropriate inspiratory time | Use of Incorrect Control Settings | PRS
FAB-H-REQ-263 | breath type | Manual breaths shall be terminated within 200 ms on activation of the High PIP alarm. | long termination time | Failure to Auto Stop | PRS
FAB-H-REQ-264 | breath type | The set pressure shall have accuracy ± (1 cmH2O + 3% of setting). | low pressure accuracy | pressure issue | PRS
FAB-H-REQ-xxx | Non-ventilating states | The fabian ventilator shall implement a standby mode, in which flow to the breathing circuit is reduced to a value primarily intended to prevent a connected heated humidifier from over-heating, breath delivery ceases, and all patient alarms are disabled. | fail to disable alarm | device alarm system issue | PRS
FAB-H-REQ-265 | Non-ventilating states | In Standby mode, the flow to the breathing circuit shall be set to a constant 4 lpm. | fail to maintain constant minimum flow | no flow | PRS
FAB-H-REQ-266 | Non-ventilating states | In Standby mode, the exhalation valve control shall be constant set PEEP of 1.0 cmH2O (subject to tolerance for PEEP, see FAB-H-REQ-237). | fail to maintain constant pressure | pressure issue | PRS
FAB-H-REQ-267 | Non-ventilating states | In Standby mode, all patient alarms shall be disabled. | fail to disable alarm | device alarm system issue | PRS
FAB-H-REQ-268 | Non-ventilating states | During standby mode there shall be continuous visual indication of the standby mode status. | fail to generate indication | no display or display failure | PRS
FAB-H-REQ-269 | Non-ventilating states | With a dual-limb breathing circuit configured, the ventilator shall resume from Standby mode after a maximum of 120 s, or when the operator cancels by pressing the standby key. | fail to resume | Reset Issue | PRS
FAB-H-REQ-270 | Non-ventilating states | With a single-limb breathing circuit configured, the ventilator shall resume from Standby mode when the operator cancels by pressing the standby key. | fail to resume | Reset Issue | PRS
FAB-H-REQ-xxx | Non-ventilating states | Whenever the Tube Occlusion alarm is activated the fabian ventilator shall cease outlet gas flow and opens the exhalation valve to relieve pressure. | fail to open valve | difficult to open or close; pressure issue | PRS
FAB-H-REQ-271 | Non-ventilating states | The tube occlusion alarm state shall be activated when the proximal pressure has exceeded the specified threshold (above intended peak inspiratory pressure) for a period of 50 milliseconds. | fail to activate occlusion alarm | occlusion within device | PRS
FAB-H-REQ-272 | Non-ventilating states | With a dual limb breathing circuit configuration, the occlusion alarm pressure threshold shall be 12 cmH2O | inappropriate threshold setting | pressure issue | PRS
FAB-H-REQ-273 | Non-ventilating states | In ventilation modes nCPAP, DuoPAP, the occlusion alarm pressure threshold shall be 10 cmH2O. | fail to activate occlusion alarm | device alarm system issue | PRS
FAB-H-REQ-274 | Non-ventilating states | In ventilation mode High Flow Oxygen Therapy, the occlusion alarm shall be disabled. | fail to disable alarm | device alarm system issue | PRS
FAB-H-REQ-275 | Non-ventilating states | The tube occlusion alarm state shall be activated in presence of an obstruction of the primary gas outlet. | fail to activate occlusion alarm | occlusion within device | PRS
FAB-H-REQ-276 | Non-ventilating states | When the tube occlusion alarm state is activated the exhalation valve control shall be set to zero (exhalation valve fully opened). | fail to open valve | difficult to open or close | PRS
FAB-H-REQ-277 | Non-ventilating states | When the tube occlusion alarm state is activated the outlet gas flow shall be disabled (outlet flowrate set to zero). | fail to disable gas flow | Device Remains Activated | PRS
FAB-H-REQ-278 | Non-ventilating states | The tube occlusion alarm state shall be de-activated after the pressure has fallen below set PEEP, at the time of the next breath trigger, or after 1s in CPAP and NCPAP modes | fail to silence alarm | device alarm system issue | PRS
FAB-H-REQ-439 | circuit test | The circuit test shall allow the use of a flow test at either 2.5 L/min, or 15 L/min, as appropriate to the selected patient size. | inappropriate test configuration | incorrect or inadequate test results | PRS
FAB-H-REQ-440 | circuit test | The circuit test shall provide means to set a test flow through the circuit, with the exhalation valve open. | inappropriate flow setting | improper flow or infusion | PRS
FAB-H-REQ-441 | circuit test | The circuit test shall provide measurements of insp pressure and prox pressure to allow the operator to determine the pressure difference between the prox pressure sensor and ambient; and the pressure difference between the insp pressure sensor and the prox pressure sensor. | low pressure accuracy | incorrect measurement | PRS
FAB-H-REQ-442 | circuit test | The circuit test shall provide a threshold test for circuit compliance, based on the time taken for circuit pressure to reach a threshold with the exhalation valve closed and a fixed outlet flow. | inappropriate test configuration | incorrect or inadequate test results | PRS
FAB-H-REQ-443 | circuit test | The circuit test shall provide a threshold test for circuit leak, based on the rate of pressure drop after the end of the compliance test. | incorrect leaking result | Incorrect or Inadequate Result | PRS
FAB-H-REQ-444 | circuit test | The circuit test shall provide means to set a test flow through the circuit, with the exhalation valve open. | inappropriate test configuration | incorrect or inadequate test results | PRS
FAB-H-REQ-445 | circuit test | The circuit test shall provide indication of prox pressure. | lack of pressure indication | device alarm system issue | PRS
FAB-H-REQ-446 | circuit test | No circuit test shall be applicable to the HFOT circuit configuration. | inappropriate test configuration | inappropriate or unexpected reset | PRS
FAB-H-REQ-539 | GUI | While the screen is locked, an icon or other graphical representation will be provided to indicate the screen locked state. | lack of locked state indication | no display or display failure | PRS
FAB-H-REQ-446 | GUI | The screen shall automatically unlock on activation of any alarm. | fail to unlock screen | no display or display failure | PRS
FAB-H-REQ-447 | GUI | When the screen is locked, means shall be provided to unlock the screen using the rotary control. | fail to unlock screen | no display or display failure | PRS
FAB-H-REQ-280 | GUI | A Screen Freeze control shall be provided to allow the screen to be frozen (to allow review of saved waveforms and loops) | fail to freeze screen | incorrect display | PRS
FAB-H-REQ-282 | GUI | A Manual Breath control shall be provided to allow the operator to trigger a manual breath. | fail to trigger breath | Device Inoperable | PRS
FAB-H-REQ-448 | GUI | A Standby control be provided to allow the ventilator to be activate or deactivate Standby mode. | fail to activate mode | use of device issue | PRS
FAB-H-REQ-449 | GUI | The standby control shall require an extended keypress minimum of 4 s and a GUI dialog requiring a confirmation touch-key press to enter standby mode. | fail to activate mode | use of device issue | PRS
FAB-H-REQ-450 | GUI | A power control shall be provided to allow the ventilator to be switched on or off. | power on or off switch failed | loss of power | PRS
FAB-H-REQ-451 | GUI | The power on/off control shall require to be held for a minimum of 4 s and a GUI dialog requiring a confirmation touch-key press to cause the ventilator to switch off. | short confirmation time | use of device issue | PRS
FAB-H-REQ-452 | GUI | On switching the ventilator off, an audible alarm shall be activated for a minimum of 120 s, capable of being silenced by action of the Alarm Silence control. | short alarm period | device alarm system issue | PRS
FAB-H-REQ-284 | GUI | A control shall be provided to allow the operator to access settings and calibration functions. | fail to access | human-device interface issue | PRS
FAB-H-REQ-285 | GUI | A control shall be provided to allow the operator to switch between displays including waveform graphs and loop graphs | unable to switch | Use of Device Issue | PRS
FAB-H-REQ-286 | GUI | A control shall be provided to allow the operator to access the main screen. | fail to access | human-device interface issue | PRS
FAB-H-REQ-287 | GUI | A control shall be provided to allow the operator to access alarm threshold settings. | difficult to setup | Device Difficult to Setup or Prepare | PRS
FAB-H-REQ-453 | GUI | A control shall be provided to allow the operator to activate the alarm system Audio Pause function. | fail to activate alarm | device alarm system issue | PRS
FAB-H-REQ-454 | GUI | The Alarm Silence control shall be labelled using icon IEC 60417-5576. | incorrect labeling | compliance issue | PRS
FAB-H-REQ-288 | GUI | A control shall be provided to allow the operator to enter the patient weight, used for display of normalized tidal volume (volume per kg), with range 300 g to 30kg. | fail to enter weight data | Use of Device Issue | PRS
FAB-H-REQ-289 | GUI | A control shall be provided to allow the operator to specify the patient size as either neonate, or pediatric | fail to specify patient size | use of device issue | PRS
FAB-H-REQ-290 | GUI | The selected patient size shall be indicated continuously during clinical use. | fail to indicate patient size | no display or display failure | PRS
FAB-H-REQ-291 | GUI | On change of the selected patient size, an option shall be provided for all ventilation controls to revert to institutional default (if stored) or factory default for that control and patient size. | fail to restore factory defaults | use of incorrect control settings | PRS
FAB-H-REQ-292 | GUI | Means shall be provided to select a ventilation mode, and then adjust the ventilation parameters for that mode, prior to activating the mode change. | fail to select mode | use of device issue | PRS
FAB-H-REQ-293 | GUI | It shall not be possible to select a ventilation mode that is not compatible with the specified breathing circuit type. | ventilation mode incomptable with circuit type | component incompatible | PRS
FAB-H-REQ-294 | GUI | Visual highlighting shall be provided to indicate that a control is selected and may be changed. | invisible indication | no display or display failure | PRS
FAB-H-REQ-295 | GUI | All changes to ventilation controls shall require a confirmation step. | lack of confirmation | use of device issue | PRS
FAB-H-REQ-296 | GUI | If two or more controls interact, a visual cue shall be provided to indicate the parameter requiring modification. | inadeqaute indication | no display or display failure | PRS
FAB-H-REQ-455 | GUI | The O2-Flush control shall be at least 2% above O2 setting when O2 setting is less than or equal to 98%, and no greater than 100% | low oxygen concentration | medical gas supply problem | PRS
FAB-H-REQ-456 | GUI | The Volume Limit control shall be available only if Volume Guarantee is disabled | fail to trigger volume limit control | decreased sensitivity | PRS
FAB-H-REQ-457 | GUI | The Volume Guarantee control shall be available only if Volume Limit is disabled | fail to trigger volume guarantee control | decreased sensitivity | PRS
FAB-H-REQ-458 | GUI | The P-Insp control shall be limited such that Pinsp + PEEP ≥ 4 cmH2O (intial inspiratory pressure) | inspiratory pressure out of range | pressure issue | PRS
FAB-H-REQ-459 | GUI | The P-PSV control shall be less than or equal to P-Insp (inspiratory pressure for the PSV breath) | inspiratory pressure out of range | pressure issue | PRS
FAB-H-REQ-460 | GUI | The P-High control shall be no less than CPAP + 2 cmH2O when mode is DuoPAP | pressure over upper limit | pressure issue | PRS
FAB-H-REQ-461 | GUI | The P-Manual control shall comply P-Manual ≥ CPAP + 2 cmH2O (mode NCPAP) | pressure out of range | pressure issue | PRS
FAB-H-REQ-462 | GUI | The Rise Time shall comply Rise-Time shall be less than or equal to I-Time | rise time out of range | imprecision | PRS
FAB-H-REQ-463 | GUI | The Rate control (modes that include mandatory breaths) and Backup Rate control (PSV mode only) shall each control the same internally stored parameter | parameter misprogrammed | use of incorrect control settings | PRS
FAB-H-REQ-xxx | GUI | Specified controls shall implement a 'locked' behaviour that requires confirmation to set an extreme value. | lack of confirmation | use of device issue | PRS
FAB-H-REQ-464 | GUI | Pinsp control shall have locked range PEEP + P-Insp > 25 cmH2O | inspiratory pressure out of range | pressure issue | PRS
FAB-H-REQ-465 | GUI | CPAP control shall have locked range CPAP > 10 cmH2O (continuous positive airway pressure) | pressure out of range | pressure issue | PRS
FAB-H-REQ-466 | GUI | PEEP control shall have locked range PEEP> 10 cmH2O (positive end-expiratory pressure) | expiratory pressure out of range | pressure issue | PRS
FAB-H-REQ-467 | GUI | PEEP control shall have locked range PEEP< 2 cmH2O | expiratory pressure out of range | pressure issue | PRS
FAB-H-REQ-468 | GUI | PSV control shall have locked range PEEP + P-PSV > 25 cmH2O (pressure support ventilation) | pressure out of range | pressure issue | PRS
FAB-H-REQ-469 | GUI | V-Guarantee control shall have locked range V-Guarantee > 30 mL | volume out of specification | tidal volume fluctuations | PRS
FAB-H-REQ-470 | GUI | V-Limit control shall have locked range V-Limit > 30 mL | volume out of specification | volume accuracy issue | PRS
FAB-H-REQ-471 | GUI | P-Max control shall have locked range P-Max > 25 cmH2O | pressure over upper limit | pressure issue | PRS
FAB-H-REQ-472 | GUI | The rate control shall have the following range (in bpm): IPPV, NEO: 2 – 200, SIPPV, NEO: 2 – 200, SIMV, NEO: 2 – 100, SIMV+PSV, NEO: 2 – 100, IPPV, PED: 2 – 100, SIPPV, NEO: 2 – 100, SIMV, NEO: 2 – 75, SIMV+PSV, NEO: 2 – 75, PSV: N/A, CPAP: N/A, HFO: N/A | rate out of range | improper device output | PRS
FAB-H-REQ-473 | GUI | The backup rate control shall have the following range (in bpm): IPPV: N/A, SIPPV: N/A, SIMV: N/A, SIMV+PSV: N/A, PSV, NEO: 2 – 200, PSV, PED: 2 - 100, CPAP: N/A, HFO: N/A | rate out of range | improper device output | PRS
FAB-H-REQ-474 | GUI | The I-Time (inspiratory time) control shall have the following range (in seconds): IPPV, NEO: 0.10 to 2.00, SIPPV, NEO: 0.10 to 2.00, SIMV, NEO: 0.10 to 2.00, SIMV+PSV, NEO: 0.10 to 2.00, PSV, NEO: 0.10 to 2.00, IPPV, PED: 0.30 to 2.00, SIPPV, PED: 0.30 to 2.00, SIMV, PED: 0.30 to 2.00, SIMV+PSV, PED: 0.30 to 2.00, PSV, PED: 0.30 to 2.00, CPAP: N/A, HFO: N/A | inspiratory time out of range | use of incorrect control settings | PRS
FAB-H-REQ-475 | GUI | E-Time (determined from Rate and I-Time) shall have the following range (in seconds): IPPV: 0.20 to 30.00, SIPPV: 0.20 to 30.00, SIMV: 0.50 to 30.00, SIMV+PSV: 0.50 to 30.00, PSV: 0.20 to 30.00, CPAP: N/A, HFO: N/A (expiratory time) | expiratory time out of range | use of incorrect control settings | PRS
FAB-H-REQ-476 | GUI | The I-Flow control (inspiratory flow) shall have the following range (in lpm): IPPV: 1 to 32, SIPPV: 1 to 32, SIMV: 1 to 32, SIMV+PSV: 1 to 32, PSV: 1 to 32, CPAP: N/A, HFO: N/A | inspiratory flow out of range | improper flow or infusion | PRS
FAB-H-REQ-477 | GUI | The E-Flow control shall have the following range (in lpm): IPPV: 1 to 32, SIPPV: 1 to 32, SIMV: 1 to 32, SIMV+PSV: 1 to 32, PSV: 1 to 32, CPAP: N/A, HFO: N/A (expiratory flow) | expiratory flow out of range | use of incorrect control settings | PRS
FAB-H-REQ-xxx | GUI | The Flow- control shall have the following range (in lpm): HFO: 5 to 20, All other modes: N/A | flow out of range | use of incorrect control settings; improper flow or infusion | PRS
FAB-H-REQ-478 | GUI | The Flow-Min control shall have the following range (in lpm): CPAP: 4 to 16, All other modes: N/A | slow flow | improper flow or infusion | PRS
FAB-H-REQ-479 | GUI | The Rise Time control shall have the following range (in seconds): IPPV, NEO: 0.10 to 2.00, SIPPV, NEO: 0.10 to 2.00, SIMV, NEO: 0.10 to 2.00, SIMV+PSV, NEO: 0.10 to 2.00, IPPV, PED: 0.30 to 2.00, SIPPV, PED: 0.30 to 2.00, SIMV, PED: 0.30 to 2.00, SIMV+PSV, NEO: 0.30 to 2.00, PSV: N/A, CPAP: N/A, HFO: N/A | rise time out of range | use of incorrect control settings | PRS
FAB-H-REQ-480 | GUI | The PEEP control shall have the following range (in mbar): IPPV: 1 to 30, SIPPV: 1 to 30, SIMV: 1 to 30, SIMV+PSV: 1 to 30, PSV: 1 to 30, CPAP: N/A, HFO: N/A | expiratory pressure out of range | pressure issue | PRS
FAB-H-REQ-481 | GUI | The CPAP control shall have the following range (in mbar): IPPV: N/A, SIPPV: N/A, SIMV: N/A, SIMV+PSV: N/A, PSV: N/A, CPAP: 1 to 30, HFO: N/A | pressure out of range | pressure issue | PRS
FAB-H-REQ-482 | GUI | The P-insp control shall have the following range (in mbar), active only when VG is disabled: IPPV: 4 to 80, SIPPV: 4 to 80, SIMV: 4 to 80, SIMV+PSV: 4 to 80, PSV: 4 to 80, CPAP: N/A, HFO: N/A | inspiratory pressure out of range | pressure issue | PRS
FAB-H-REQ-483 | GUI | The P-max control shall have the following range (in mbar): IPPV: 4 to 80, SIPPV: 4 to 80, SIMV: 4 to 80, SIMV+PSV: 4 to 80, PSV: 4 to 80, CPAP: N/A, HFO: N/A | pressure out of range | pressure issue | PRS
FAB-H-REQ-484 | GUI | The P-PSV control shall have the following range (in mbar): IPPV: N/A, SIPPV: N/A, SIMV: N/A, SIMV+PSV: 3 to 80, PSV: 3 to 80, CPAP: N/A, HFO: N/A | inspiratory pressure out of range | pressure issue | PRS
FAB-H-REQ-485 | GUI | The P-manual control shall have the following range (in mbar): IPPV: N/A, SIPPV: N/A, SIMV: N/A, SIMV+PSV: N/A, PSV: N/A, CPAP: 4 to 80, HFO: 6 to 80 | pressure out of range | pressure issue | PRS
FAB-H-REQ-xxx | GUI | The Pmean control shall have the following range (in mbar): HFO: 5 to 50 | pressure out of range | pressure issue | PRS
FAB-H-REQ-xxx | GUI | The Amp control shall have the following range (in mbar): HFO: 5 to 100 (Initial oscillating pressure amplitude) | pressure amplitude out of range | use of incorrect control settings | PRS
FAB-H-REQ-xxx | GUI | The Amp,max control shall have the following range (in mbar): HFO: 5 to 100 (Maximum oscillating pressure amplitude) | pressure amplitude out of range | use of incorrect control settings | PRS
FAB-H-REQ-xxx | GUI | The HF Freq control shall have the following range (in Hz): HFO: 5 to 20 | frequency out of range | electronic property issue | PRS
FAB-H-REQ-xxx | GUI | The I:E control shall have the following range (no units): HFO: 1:1 to 1:3 (inspiration to expiration ratio) | ratio out of range | use of incorrect control settings | PRS
FAB-H-REQ-xxx | GUI | The Pmean,rec control shall have the following range (in mbar): HFO: 7 to 50 | pressure out of range | pressure issue | PRS
FAB-H-REQ-xxx | GUI | The I-time,rec control shall have the following range (in seconds): HFO: 2 to 60 | inspiratory time out of range | use of incorrect control settings | PRS
FAB-H-REQ-xxx | GUI | The Freq,req control shall have the following range: HFO: Off, 1/hour to 59/hour, 1/min to 4/min | frequency out of range | electronic property issue | PRS
FAB-H-REQ-486 | GUI | The O2 control shall have the following range (in percent): IPPV: 21 to 100, SIPPV: 21 to 100, SIMV: 21 to 100, SIMV+PSV: 21 to 100, PSV: 21 to 100, CPAP: 21 to 100, HFO: 21 to 100 | oxygen out of specification | medical gas supply problem | PRS
FAB-H-REQ-487 | GUI | The O2 Flush control shall have the following range (in percent): IPPV: 23 to 100, SIPPV: 23 to 100, SIMV: 23 to 100, SIMV+PSV: 23 to 100, PSV: 23 to 100, CPAP: 23 to 100, HFO: 23 to 100 | oxygen out of specification | medical gas supply problem | PRS
FAB-H-REQ-488 | GUI | The V-limit control shall have the following range (in mL), active only when VG is disabled: IPPV, NEO: 2 to 150, SIPPV, NEO: 2 to 150, SIMV, NEO: 2 to 150, SIMV+PSV, NEO: 2 to 150, PSV, NEO: 2 to 150, IPPV, PED: 10 to 500, SIPPV, PED: 10 to 500, SIMV, PED: 10 to 500, SIMV+PSV, NEO: 10 to 500, PSV, PED: 10 to 500, CPAP: N/A, HFO: N/A | volume out of specification | volume accuracy issue | PRS
FAB-H-REQ-489 | GUI | The V-guarantee control shall have the following range (in mL): IPPV, NEO: 2 to 60, SIPPV, NEO: 2 to 60, SIMV, NEO: 2 to 60, SIMV+PSV, NEO: 2 to 60, PSV, NEO: 2 to 60, IPPV, PED: 10 to 300, SIPPV, PED: 10 to 300, SIMV, PED: 10 to 300, SIMV+PSV, NEO: 10 to 300, PSV, PED: 10 to 300, CPAP: N/A, HFO: 0.3 to 30 | volume out of specification | tidal volume fluctuations | PRS
FAB-H-REQ-490 | GUI | The Trigger Mode control shall have the following range (as enumerated list): IPPV: N/A, SIPPV: Pressure, Flow, Volume, SIMV: Pressure, Flow, Volume, SIMV+PSV: Pressure, Flow, Volume, PSV: Pressure, Flow, Volume, CPAP: N/A, HFO: N/A | invalid trigger mode | use of incorrect control settings | PRS
FAB-H-REQ-491 | GUI | The Trigger control shall have the following range (as A.U.): IPPV: N/A, SIPPV: 1 to 10, SIMV: 1 to 10, SIMV+PSV: 1 to 10, PSV: 1 to 10, CPAP: N/A, HFO: N/A | invalid trigger control | use of incorrect control settings | PRS
FAB-H-REQ-492 | GUI | The Termination-PSV control shall have the following range (as percentage): IPPV: N/A, SIPPV: N/A, SIMV: N/A, SIMV+PSV: 1 to 85, PSV: 1 to 85, CPAP: N/A, HFO: N/A | pressure out of range | pressure issue | PRS
FAB-H-REQ-493 | GUI | The Backup control shall have the following range (as number of breaths): IPPV: N/A, SIPPV: N/A, SIMV: N/A, SIMV+PSV: N/A, PSV: N/A, CPAP: 1 to 5, HFO: N/A | control out of range | use of incorrect control settings | PRS
FAB-H-REQ-494 | GUI | The rate control shall have the following range (in bpm): nCPAP: N/A, DuoPAP: 2 to 60, O2 Therapy: N/A | rate out of range | improper device output | PRS
FAB-H-REQ-495 | GUI | The I-Time control shall have the following range (in seconds): nCPAP: N/A, DuoPAP: 0.15 to 15.00, O2 Therapy: N/A | inspiratory time out of range | use of incorrect control settings | PRS
FAB-H-REQ-496 | GUI | E-Time (determined from Rate and I-Time) shall have the following range (in seconds): nCPAP: N/A, DuoPAP: 0.20 to 30.00, O2 Therapy: N/A | expiratory time out of range | use of incorrect control settings | PRS
FAB-H-REQ-497 | GUI | The Flow control shall have the following range (in lpm): nCPAP: N/A, DuoPAP: N/A, O2 Therapy, NEO: 1 to 15, O2 Therapy, PED: 1 to 30 | flow out of range | improper flow or infusion | PRS
FAB-H-REQ-498 | GUI | The CPAP control shall have the following range (in mbar): nCPAP: 2 to 13, DuoPAP: 2 to 13, O2 Therapy: N/A | pressure out of range | pressure issue | PRS
FAB-H-REQ-499 | GUI | The P-manual control shall have the following range (in mbar): nCPAP: 5 to 15, DuoPAP: N/A, O2 Therapy: N/A | pressure out of range | pressure issue | PRS
FAB-H-REQ-500 | GUI | The P-high control shall have the following range (in mbar): nCPAP: N/A, DuoPAP: 5 to 15, O2 Therapy: N/A | pressure over upper limit | pressure issue | PRS
FAB-H-REQ-501 | GUI | The O2 control shall have the following range (in percentage): nCPAP: 21 to 100, DuoPAP: 21 to 100, O2 Therapy: 21 to 100 | oxygen out of specification | medical gas supply problem | PRS
FAB-H-REQ-502 | GUI | The O2 Flush control shall have the following range (in percentage): nCPAP: 23 to 100, DuoPAP: 23 to 100, O2 Therapy: 23 to 100 | oxygen out of specification | medical gas supply problem | PRS
FAB-H-REQ-xxx | GUI | The FOT amplitude shall be fixed: 5 mbar (forced oscillation technique) | pressure out of range | pressure issue | PRS
FAB-H-REQ-xxx | GUI | The FOT forced oscillation frequency shall be fixed: 10 Hz | frequency out of range | electronic property issue | PRS
FAB-H-REQ-xxx | GUI | The FOT forced oscillation I:E ratio shall be fixed: 1:1 | ratio out of range | use of incorrect control settings | PRS
FAB-H-REQ-xxx | GUI | The FOT period shall be less than or equal to 18 seconds. | period out of range | use of incorrect control settings | PRS
FAB-H-REQ-xxx | GUI | The FOT stabilization period shall be less than or equal to 180 seconds. | period out of range | electronic property issue | PRS
FAB-H-REQ-xxx | GUI | The FOT algorithm step count shall be set to an odd number within the range 1 to 21 steps. | step count out of range | use of incorrect control settings | PRS
FAB-H-REQ-xxx | GUI | The FOT mean pressure (FOT during HFO ventilation) is stepped from Pmean,low to Pmean,high and back in FOT steps. | inappropriate pressure steps | pressure issue | PRS
FAB-H-REQ-xxx | GUI | Pmean low shall have range 5 to 49 mbar, and Pmean high shall have range 6 to 50 mbar | pressure out of range | pressure issue | PRS
FAB-H-REQ-xxx | GUI | The FOT PEEP (FOT during conventional ventilation modes) is stepped from PEEP,low to PEEP,high and back in FOT steps. | inappropriate pressure steps | use of incorrect control settings | PRS
FAB-H-REQ-xxx | GUI | PEEP low shall have range 3 to 29 mbar, and PEEP high shall have range 4 to 30 mbar | pressure out of range | pressure issue | PRS
FAB-H-REQ-297 | GUI | The ventilator shall provide indication of expiratory time, calculated from set I-Time and Rate. | lack of time indication | no display or display failure | PRS
FAB-H-REQ-298 | GUI | The E-Time displayed shall cover at least the range 0.2 to 30.0 s. | incorrect time display range | incorrect display | PRS
FAB-H-REQ-299 | GUI | E-Time values shall have accuracy ±(0.05 s). | low time accuracy | imprecision | PRS
FAB-H-REQ-300 | GUI | The ventilator shall provide indication of I:E Ratio, calculated from set I-Time and Rate. | incorrect ratio calculation | parameter calculation error due to software problem | PRS
FAB-H-REQ-301 | GUI | I:E ratio values shall be displayed as 1:xx.x (inspiratory time less than or equal to expiratory time) or xx.x:1 (inspiratory time greater than expiratory time). | incorrect display format | incorrect display | PRS
FAB-H-REQ-302 | GUI | I:E ratio values shall have accuracy ±(0.1 in the larger value). | low ratio accuracy | imprecision | PRS
FAB-H-REQ-xxx | GUI | PIP_max shall be calculated from the high pressure setting and the ventilation’s actual ΔP, PIP_max is the maximally reached peak pressure during the whole recruitment procedure, PIP_max is limited to 100 mbar in HFO mode respective 80 mbar for Conventional modes.. | data calculated incorrectly | incorrect software programming calculations | PRS
FAB-H-REQ-xxx | GUI | Step shall be calculated as the pressure difference between two consecutive FOT measurements. It is automatically calculated from the difference between the low and high-pressure settings during FOT. The minimum step is 1 mbar. | data calculated incorrectly | incorrect software programming calculations | PRS
FAB-H-REQ-36 | Measured values | The PIP monitor shall be available in following modes: IPPV, SIPPV, SIMV, SIMV+PSV, PSV, DuoPAP, HFO | lack of mode monitor | inadequate user interface | PRS
FAB-H-REQ-37 | Measured values | The Pmean monitor shall be available in following modes: IPPV, SIPPV, SIMV, SIMV+PSV, PSV, CPAP, nCPAP, DuoPAP, O2 Therapy, HFO | lack of mode monitor | inadequate user interface | PRS
FAB-H-REQ-38 | Measured values | The PEEP monitor shall be available in following modes: IPPV, SIPPV, SIMV, SIMV+PSV, PSV, DuoPAP | lack of mode monitor | inadequate user interface | PRS
FAB-H-REQ-39 | Measured values | The Vte monitor shall be available in following modes: IPPV, SIPPV, SIMV, SIMV+PSV, PSV, CPAP, HFO | lack of mode monitor | inadequate user interface | PRS
FAB-H-REQ-40 | Measured values | The MV-exp monitor shall be available in following modes: IPPV, SIPPV, SIMV (mand + spon), SIMV+PSV (mand + spon), PSV, CPAP | lack of mode monitor | inadequate user interface | PRS
FAB-H-REQ-41 | Measured values | The Vte/kg monitor shall be available in following modes: IPPV, SIPPV, SIMV (mand + spon), SIMV+PSV (mand + spon), PSV, CPAP | lack of mode monitor | inadequate user interface | PRS
FAB-H-REQ-42 | Measured values | The Vti monitor shall be available in following modes: IPPV, SIPPV | lack of mode monitor | inadequate user interface | PRS
FAB-H-REQ-43 | Measured values | The leak monitor shall be available in following modes: IPPV, SIPPV, SIMV, SIMV+PSV, PSV, CPAP, HFO | lack of mode monitor | inadequate user interface | PRS
FAB-H-REQ-44 | Measured values | The O2 monitor shall be available in following modes: IPPV, SIPPV, SIMV, SIMV+PSV, PSV, CPAP, nCPAP, DuoPAP, O2 Therapy, HFO | lack of mode monitor | inadequate user interface | PRS
FAB-H-REQ-45 | Measured values | The Rate monitor shall be available in following modes: IPPV, SIPPV, SIMV, SIMV+PSV, PSV, CPAP | lack of mode monitor | inadequate user interface | PRS
FAB-H-REQ-46 | Measured values | The Compliance monitor (C) shall be available in following modes: IPPV, SIPPV, SIMV, SIMV+PSV, PSV | lack of mode monitor | inadequate user interface | PRS
FAB-H-REQ-47 | Measured values | The C20/C monitor shall be available in following modes: IPPV, SIPPV, SIMV, SIMV+PSV, PSV | lack of mode monitor | inadequate user interface | PRS
FAB-H-REQ-48 | Measured values | The Resistance monitor (R) shall be available in following modes: IPPV, SIPPV, SIMV, SIMV+PSV, PSV | lack of mode monitor | inadequate user interface | PRS
FAB-H-REQ-49 | Measured values | The Time Constant monitor (RC) shall be available in following modes: IPPV, SIPPV, PSV | lack of mode monitor | inadequate user interface | PRS
FAB-H-REQ-50 | Measured values | The Spont %MV monitor shall be available in following modes: SIMV, SIMV+PSV | lack of mode monitor | inadequate user interface | PRS
FAB-H-REQ-51 | Measured values | The Flow monitor shall be available in following modes: nCPAP, DuoPAP, O2 Therapy | lack of mode monitor | inadequate user interface | PRS
FAB-H-REQ-52 | Measured values | The EtCO2 monitor shall be available in following modes (when capnometry module is connected and enabled): IPPV, SIPPV, SIMV, SIMV+PSV, PSV, CPAP, nCPAP, DuoPAP, O2 Therapy, HFO | lack of mode monitor | inadequate user interface | PRS
FAB-H-REQ-53 | Measured values | The FiO2 monitor shall be available in following modes: IPPV, SIPPV, SIMV, SIMV+PSV, PSV, CPAP, nCPAP, DuoPAP, O2 Therapy, HFO | lack of mode monitor | inadequate user interface | PRS
FAB-H-REQ-54 | Measured values | The SpO2 monitor shall be available in following modes (when pulse oximetry module is connected and enabled): IPPV, SIPPV, SIMV, SIMV+PSV, PSV, CPAP, nCPAP, DuoPAP, O2 Therapy, HFO | lack of mode monitor | inadequate user interface | PRS
FAB-H-REQ-55 | Measured values | The Pulse Rate monitor shall be available in following modes (when pulse oximetry module is connected and enabled): IPPV, SIPPV, SIMV, SIMV+PSV, PSV, CPAP, nCPAP, DuoPAP, O2 Therapy, HFO | lack of mode monitor | inadequate user interface | PRS
FAB-H-REQ-56 | Measured values | The Perfusion monitor shall be available in following modes (when pulse oximetry module is connected and enabled): IPPV, SIPPV, SIMV, SIMV+PSV, PSV, CPAP, nCPAP, DuoPAP, O2 Therapy, HFO | lack of mode monitor | inadequate user interface | PRS
FAB-H-REQ-xxx | Measured values | The Vte HFO monitor shall be available in HFO mode only. | inappropriate mode monitor | data issue | PRS
FAB-H-REQ-xxx | Measured values | The HFO Amp monitor shall be available in HFO mode only. | inappropriate mode monitor | data issue | PRS
FAB-H-REQ-xxx | Measured values | The HFO Freq monitor shall be available in HFO mode only. | inappropriate mode monitor | data issue | PRS
FAB-H-REQ-xxx | Measured values | The DCO2 monitor shall be available in HFO mode only. | inappropriate mode monitor | data issue | PRS
FAB-H-REQ-xxx | Measured values | Measured values derived from time-varying measurements of airway pressure shall share a common specification for accuracy. Airway Pressure measurements use a proximal line to provide for monitoring of airway pressure at the patient interface. | low measurement accuracy | imprecision | PRS
FAB-H-REQ-57 | Measured values | Measurements of airway pressure shall cover at least the range -10 to +125 mbar. | measurement out of range | pressure issue | PRS
FAB-H-REQ-58 | Measured values | Airway pressure values shall have accuracy ±(1 mbar + 4% of actual value).PIP | low pressure accuracy | pressure issue | PRS
FAB-H-REQ-303 | Measured values | In all modes excluding SIMV+PSV, the value PIP shall represent the peak airway pressure during the period of each mandatory or PSV breath, updated prior to the start of the next mandatory or PSV breath. | fail to update pressure | pressure issue | PRS
FAB-H-REQ-517 | Measured values | In mode SIMV+PSV, the value PIP shall represent the peak airway pressure during the period of each mandatory breath, updated prior to the start of the next mandatory breath. | fail to update pressure | pressure issue | PRS
FAB-H-REQ-304 | Measured values | PIP shall be displayed within the range 0 to +120 mbar | incorrect display range | incorrect display | PRS
FAB-H-REQ-305 | Measured values | The value Pmean shall represent the average airway pressure during the whole of the inspiratory and expiratory period of each breath, updated prior to the start of the next mandatory or PSV breath. | fail to update pressure | pressure issue | PRS
FAB-H-REQ-306 | Measured values | Pmean shall be displayed within the range 0 to +120 mbar. | incorrect display range | incorrect display | PRS
FAB-H-REQ-307 | Measured values | The value PEEP shall represent the airway pressure at the end of the expiratory period of each breath, excluding any period during which airway pressure is reduced by patient spontaneous inspiratory effort, updated prior to the start of the next mandatory or PSV breath. | fail to update pressure | pressure issue | PRS
FAB-H-REQ-308 | Measured values | PEEP shall be displayed within the range 0 to +60 mbar. | incorrect display range | incorrect display | PRS
FAB-H-REQ-xxx | Measured values | Measured values derived from time-varying measurements of airway flow shall share a common specification for range and accuracy. Volume measurements use a proximal hot-wire flow sensor to provide for monitoring of airway flow at the patient connection port. Volume measurements are not available with the single limb breathing circuit configurations, or with the Hot Wire Flow Sensor disabled. | low measurement accuracy | volume accuracy issue | PRS
FAB-H-REQ-59 | Measured values | Measurements of tidal volume shall cover at least the range 0 to +999 mL. | inappropriate measurement range | measurement system incompatibility | PRS
FAB-H-REQ-60 | Measured values | Tidal Volume values shall have accuracy ±(1 mL + 8 % of measurement). | low volume accuracy | volume accuracy issue | PRS
FAB-H-REQ-61 | Measured values | The value Vte shall represent the exhaled tidal volume during the period of each mandatory or PSV breath, updated prior to the start of the next mandatory or PSV breath. | fail to update tidal volume | tidal volume fluctuations | PRS
FAB-H-REQ-62 | Measured values | The value Vti shall represent the delivered tidal volume during the period of each mandatory or PSV breath, updated prior to the start of the next mandatory or PSV breath. | fail to update tidal volume | tidal volume fluctuations | PRS
FAB-H-REQ-309 | Measured values | The value leak shall represent the fraction of delivered volume that is lost to airway leak and not returned to the breathing circuit as expired volume. | incorrect leak calculation | parameter calculation error due to software problem | PRS
FAB-H-REQ-310 | Measured values | The leak parameter shall be indicated within range 10 to 50 %. | incorrect leak range | leak | PRS
FAB-H-REQ-311 | Measured values | The leak parameter shall have accuracy ± 10 %. | low leakage parameter accuracy | incorrect measurement | PRS
FAB-H-REQ-312 | Measured values | The value Ve shall represent the exhaled minute ventilation averaged over a period of one minute (exhaled volume). | volume out of specification | tidal volume fluctuations | PRS
FAB-H-REQ-313 | Measured values | The Ve parameter shall be indicated within range 0 to 9.9 L. | incorrect display range | incorrect display | PRS
FAB-H-REQ-314 | Measured values | The Ve parameter shall have accuracy ± (0.1 L + 8 % of measurement). | low volume accuracy | volume accuracy issue | PRS
FAB-H-REQ-315 | Measured values | The value Spont % Ve shall represent the exhaled minute ventilation for spontaneous breaths, as a fraction of total exhaled minute ventilation, averaged over a period of one minute. | incorrect volume calculation | parameter calculation error due to software problem | PRS
FAB-H-REQ-316 | Measured values | The Spont % Ve parameter shall be indicated within range 0 to 100 %. | incorrect display range | incorrect display | PRS
FAB-H-REQ-63 | Measured values | The value Trigger Volume shall represent the inspired volume prior to initiation of a patient-triggered breath. | volume trigger failed | decreased sensitivity | PRS
FAB-H-REQ-64 | Measured values | The ventilator shall provide measurement of volume fraction of oxygen at the outlet. | loss of measurement | measurement system incompatibility | PRS
FAB-H-REQ-65 | Measured values | The O2 measurement shall cover at least the range 18 to 100 %. | inappropriate measurement range | incorrect measurement | PRS
FAB-H-REQ-66 | Measured values | O2 values shall have accuracy ±(3 % absolute). | low oxygen fraction accuracy | medical gas supply problem | PRS
FAB-H-REQ-67 | Measured values | The ventilator shall provide indication of total breathing rate (including mandatory and PSV breaths) provided that the flow sensor is calibrated and connected to the patient circuit. | inadeqaute indication | no display or display failure | PRS
FAB-H-REQ-68 | Measured values | The Rate measurement shall cover at least the range 0 to 150 bpm. | inappropriate measurement range | incorrect measurement | PRS
FAB-H-REQ-69 | Measured values | Rate values shall have accuracy ±(1 bpm). | low rate accuracy | imprecision | PRS
FAB-H-REQ-317 | Measured values | The ventilator shall provide measurement of dynamic compliance. | inadeqaute measurement | incorrect measurement | PRS
FAB-H-REQ-318 | Measured values | Compliance shall be displayed within range 0 to 300 mL/cmH2O. | incorrect display range | incorrect display | PRS
FAB-H-REQ-319 | Measured values | Compliance shall have accuracy ±(8 % of measurement). | low compliance accuracy | incorrect measurement | PRS
FAB-H-REQ-320 | Measured values | The ventilator shall provide measurement of airway resistance. | inadeqaute resistance measurement | physical resistance | PRS
FAB-H-REQ-321 | Measured values | Resistance shall be displayed within range 0 to 2000 cmH2O / L/s | resistance out of spec | physical resistance | PRS
FAB-H-REQ-322 | Measured values | Resistance shall have accuracy ±(8 % of measurement). | low resistance accuracy | incorrect measurement | PRS
FAB-H-REQ-323 | Measured values | The ventilator shall provide measurement of respiratory time constant. | inadequate time measurement | incorrect measurement | PRS
FAB-H-REQ-324 | Measured values | RC shall be displayed within range 0 to 2.5 s. | incorrect display range | incorrect display | PRS
FAB-H-REQ-70 | Measured values | The ventilator shall provide measurement of Rapid Shallow Breathing Index, defined as the ratio of breath rate and tidal volume. | inaccurate measurement item | incorrect measurement | PRS
FAB-H-REQ-325 | Measured values | The ventilator shall provide measurement of C20/C, defined as the compliance calculated over the last 20 % of the inspired volume, divided by the compliance calculated over the whole breath. | inaccurate measurement item | incorrect measurement | PRS
FAB-H-REQ-326 | Measured values | C20/C shall be displayed within range 0.0 to 5.0 | display value out of range | incorrect display | PRS
FAB-H-REQ-xxx | Measured values | Measurement of pulmonary reactance shall be available when the FOT option is enabled and selected. This estimates pulmonary reactance using high-frequency pressure changes imposed onto the ventilation baseline. | inadeqaute reactance measurement | incorrect measurement | PRS
FAB-H-REQ-xxx | Measured values | Xrx shall be displayed within range of –500 to 100 mbar/(L/sec). | display value out of range | incorrect display | PRS
FAB-H-REQ-xxx | Measured values | FOT shall be available in all invasive ventilation modes using a dual limb breathing circuit (Forced Oscillation Technique) | inadequate ventilation mode | failure to cycle | PRS
FAB-H-REQ-xxx | Measured values | FOT shall be available for use together with Volume Guarantee, Volume Limit, SpO2 measuring and PRICO. | inadequate ventilation mode | failure to cycle | PRS
FAB-H-REQ-xxx | Measured values | In conventional modes, the forced oscillation shall be created by the exhalation membrane. During the measurement, ventilation mode is temporarily changed to CPAP. | fail to initiate forced oscillation | failure to cycle | PRS
FAB-H-REQ-xxx | Measured values | In HFO mode, the forced oscillation shall be created by the HFO module. During the measurement, HFO is temporarily paused. | fail to initiate forced oscillation | failure to cycle | PRS
FAB-H-REQ-71 | Measured values | The ventilator shall provide measurement of Xrs, defined as the energy storage capacity of the respiratory system, determined from both elastic and inertive properties, as an option when using a dual-limb breathing circuit configuration only. | inadeqaute measurement | incorrect measurement | PRS
FAB-H-REQ-xxx | Measured values | The pulmonary reactance measurement requires a period of CPAP (to allow lung pressure to stabilize), followed by reactance measurement, prior to resumption of the prior ventilation. The ventilator shall initiate the measurement sequence only in response to a direct user interaction. | inadeqaute reactance measurement | incorrect measurement | PRS
FAB-H-REQ-72 | Measured values | Xrs shall be displayed within range -500 to +100 cmH2O / L/s. | incorrect display range | incorrect display | PRS
FAB-H-REQ-73 | Measured values | The ventilator shall provide for measurement of Xrs at a sequence of PEEP settings, with PEEP incremented with a preset number of steps each with a preset incremental pressure. | inadeqaute measurement | incorrect measurement | PRS
FAB-H-REQ-74 | Measured values | The ventilator shall provide display of Xrs plotted on the Y axis versus PEEP plotted on the X axis. | inadequate message display | incorrect display | PRS
FAB-H-REQ-xxx | Measured values | The FOT Graph is cleared if any of the following actions are performed: a new FOT sequence is started, the device is restarted, the patient range is changed, device date / time is changed | fail to reset graph | image display error | PRS
FAB-H-REQ-xxx | Measured values | FOT is turned OFF and cannot be re-enabled while any of the following actions occurs or alarms are active: Any system alarm, Any flow sensor alarm, Any system failure, alarm actions | fail to disable forced oscillation | device remains activated | PRS
FAB-H-REQ-xxx | Measured values | The system shall export the FOT settings and the measurement results (Xrs and Resistance) at each pressure level via system log files | fail to export settings | failure to convert to back-up | PRS
FAB-H-REQ-xxx | Measured values | The ventilator shall provide measurements of capnometry parameters when a capnometer device is in use. | fail to measure | measurement system incompatibility | PRS
FAB-H-REQ-xxx | Measured values | The ventilator shall provide for the operator to configure CO2 values to be displayed in units of mmHg, kPa, or Vol%. | incorrect display format | incorrect display | PRS
FAB-H-REQ-75 | Measured values | The ventilator shall provide for the operator to configure CO2 values to be displayed in units of mmHg, kPa, or Vol%. | fail to setup configuration | device difficult to setup or prepare | PRS
FAB-H-REQ-327 | Measured values | The ventilator shall provide display of end-tidal CO2, EtCO2. | inadequate message display | incorrect display | PRS
FAB-H-REQ-328 | Measured values | EtCO2 shall be displayed within range 0 to 150 mmHg, or 0 to 20 kPa. (partial pressure of CO2, end-tidal CO2) | pressure measurement out of range | pressure issue | PRS
FAB-H-REQ-329 | Measured values | EtCO2 with the Oridion micropod option shall have accuracy ± 2 mmHg (0 to 38 mmHg), ± (5% of reading + 8% of (reading – 39 mmHg)) (39 to 150 mmHg) | low pressure accuracy | imprecision | PRS
FAB-H-REQ-xxx | Measured values | EtCO2 with the Respironics capnostat option shall have accuracy | low accuracy | imprecision | PRS
FAB-H-REQ-330 | Measured values | The ventilator shall provide display of inspired CO2, FiCO2. | fail to measure | measurement system incompatibility | PRS
FAB-H-REQ-331 | Measured values | FiCO2 shall be displayed within range 0 to 150 mmHg, or 0 to 20 kPa. | pressure measurement out of range | pressure issue | PRS
FAB-H-REQ-332 | Measured values | FiCO2 with the Oridion micropod option shall have accuracy ± 2 mmHg (0 to 38 mmHg), ± (5% of reading + 8% of (reading – 39 mmHg)) (39 to 150 mmHg) | low pressure accuracy | imprecision | PRS
FAB-H-REQ-xxx | Measured values | FiCO2 with the Respironics capnostat option shall have accuracy | low pressure accuracy | imprecision | PRS
FAB-H-REQ-333 | Measured values | The ventilator shall provide a waveform graphic display of partial pressure of CO2. | confused graph | image display error | PRS
FAB-H-REQ-xxx | Measured values | The ventilator shall provide measurements of pulse oximetry parameters when a pulse oximeter device is in use. | inadeqaute oxygen measurement | incorrect measurement | PRS
FAB-H-REQ-335 | Measured values | The ventilator shall provide display of peripheral oxygen saturation,--> SpO2 | inadequate oxygen display | incorrect display | PRS
FAB-H-REQ-336 | Measured values | SpO2 shall be displayed within range 1 to 100 %. | incorrect display range | incorrect display | PRS
FAB-H-REQ-337 | Measured values | SpO2 shall have accuracy specified only within the range 70 to 100%, as ± 2% pediatric patients, no motion conditions ± 3% pediatric patients, motion conditions, and all neonatal patients. | low oxygen accuracy | imprecision | PRS
FAB-H-REQ-338 | Measured values | The ventilator shall provide display of pulse rate. | inadequate message display | incorrect display | PRS
FAB-H-REQ-339 | Measured values | Pulse Rate shall be displayed within range 25 to 240 bpm. | incorrect display range | incorrect display | PRS
FAB-H-REQ-340 | Measured values | Pulse Rate shall have accuracy ± 3 bpm (no motion conditions) or ± 5 bpm (motion conditions). | low measurement accuracy | incorrect measurement | PRS
FAB-H-REQ-341 | Measured values | The ventilator shall provide display of perfusion index. | inadequate message display | no display or display failure | PRS
FAB-H-REQ-342 | Measured values | Perfusion Index shall be displayed within range 0.02 to 20.0 %. | incorrect display range | incorrect display | PRS
FAB-H-REQ-343 | Measured values | The ventilator shall provide a waveform graphic display of the pulse oximeter plethysmographic signal. | inadequate graphic display | image display error | PRS
FAB-H-REQ-344 | Measured values | The ventilator shall provide a graphic display of SIQ. | inadequate graphic display | incorrect display | PRS
FAB-H-REQ-345 | Measured values | The ventilator shall allow configuration of pulse oximeter averaging time as 2-4, 4-6, 8, 10, 12, 14 or 16 seconds. | incorrect time configuration | use of incorrect control settings | PRS
FAB-H-REQ-346 | Measured values | The ventilator shall allow configuration of pulse oximeter sensitivity as Normal, Max, or APOD (advanced probe off detection). | inappropriate sensitivity setting | use of incorrect control settings | PRS
FAB-H-REQ-347 | Measured values | The ventilator shall allow configuration of pulse oximeter FastSat algorithm as On or Off. | inappropriate algorithm setting | use of incorrect control settings | PRS
FAB-H-REQ-76 | Measured values | Waveform graphic screens shall be provided that plot pressure, flow and / or volume against time using user-configurable axis scaling | inadequate graphic display | image display error | PRS
FAB-H-REQ-77 | Measured values | In waveform graphic screens that show breathing activity plotted against time, patient-triggered breaths shall be identified by use of a distinct color. | illegible graph | image display error | PRS
FAB-H-REQ-6 | Measured values | Loop graphic screens shall be provided that plot volume against pressure, and / or Flow against volume, using user-configurable axis scaling. | incorrect graph | image display error | PRS
FAB-H-REQ-78 | Measured values | A Trend screen shall be provided that plots measured values, averaged over a sampling period of 30 seconds. | inadequate graphic display | image display error | PRS
FAB-H-REQ-79 | Measured values | The trend screen shall provide for the following parameters to be selected for display: Pinsp, Pmean, FiO2, Vte, Compliance, Mve, RSBI, % Mvmand, Resistance, leak, SPO2, SPO2 PI, SPO2 PR, ETCO2, Freq | inadequate graphic display | incorrect display | PRS
FAB-H-REQ-80 | Measured values | The trend screen shall allow the operator to select the timebase (time axis scaling) and to scroll through available trend data. | fail to setup timebase | use of device issue; use of incorrect control settings | PRS
FAB-H-REQ-81 | Measured values | The ventilator shall provide for up to 5 days continuous trend data storage. | inadequate data storage | data issue | PRS
FAB-H-REQ-82 | Alarm system | The ventilator shall provide audible indication of alarm activation using the priority encoding specified in IEC 60601-1-8. | unnoticeable alarm | low audible alarm | PRS
FAB-H-REQ-83 | Alarm system | The ventilator shall provide visual indication of alarm activation to comply with the 1m and 4m indication classes specified in IEC 60601-1-8. | invisible indication | alarm not visible | PRS
FAB-H-REQ-84 | Alarm system | A control shall be provided to allow the volume of the audible alarm signal to be set to a minimum of three distinct volume settings, within range 49 to 65 dB(A). | fail to configure audio alarm | device alarm system issue | PRS
FAB-H-REQ-7 | Alarm system | The audio pause function shall cause the alarm audible indication to be silenced for a period not exceeding 120 s. | fail to silence alarm | device alarm system issue | PRS
FAB-H-REQ-xxx | Alarm system | An alarm limits screen shall be provided to allow the operator to change the alarm limits for patient alarms applicable to the current mode of ventilation. | fail to set limit correctly | device alarm system issue | PRS
FAB-H-REQ-85 | Alarm system | All alarm conditions shall be implemented as non-latching alarms. | fail to activate alarm | device alarm system issue | PRS
FAB-H-REQ-86 | Alarm system | When an alarm deactivates, if no other alarm remains active, an informational message shall be displayed showing the last alarm active. | incorrect message display | device alarm system issue | PRS
FAB-H-REQ-87 | Alarm system | The informational message showing the last alarm active shall be cleared by action of the alarm audio pause control. | fail to clear message | device alarm system issue | PRS
FAB-H-REQ-88 | Alarm system | An alarm autoset function shall be provided to allow alarm limits to be preset based on measured values. | fail to preset limit correctly | device alarm system issue | PRS
FAB-H-REQ-89 | Alarm system | Alarm thresholds that are user configurable and related to a monitor value shall be indicated alongside that monitor value. | fail to configure alarm threshold | device alarm system issue | PRS
FAB-H-REQ-90 | Alarm system | All alarms when active shall suppress indication of any alarm with the same, or higher, sequence value as tabulated below, with specified exceptions: Flat battery! Alarm over-rides Charge Battery <15 mins | fail to suppress low priority alarms | device alarm system issue | PRS
FAB-H-REQ-349 | Alarm system | The Cooling fan defect alarm shall have following parameters: Range: N/A, Autoset: N/A, Priority: High, Sequence: 1 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-350 | Alarm system | The COM interface alarm shall have following parameters: Range: N/A, Autoset: N/A, Priority: High, Sequence: 1 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-351 | Alarm system | The DIO interface alarm shall have following parameters: Range: N/A, Autoset: N/A, Priority: High, Sequence: 1 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-352 | Alarm system | The I2C interface alarm shall have following parameters: Range: N/A, Autoset: N/A, Priority: High, Sequence: 1 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-353 | Alarm system | The Parallel interface alarm shall have following parameters: Range: N/A, Autoset: N/A, Priority: High, Sequence: 1 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-354 | Alarm system | The SPI interface alarm shall have following parameters: Range: N/A, Autoset: N/A, Priority: High, Sequence: 1 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-355 | Alarm system | The Flat Battery alarm shall have following parameters: Range: N/A, Autoset: N/A, Priority: High, Sequence: 2 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-356 | Alarm system | The Charge Battery < 15 mins alarm shall have following parameters: Range: N/A, Autoset: N/A, Priority: High, Sequence: 2 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-357 | Alarm system | The Checksum Controller PIC alarm shall have following parameters: Range: N/A, Autoset: N/A, Priority: High, Sequence: 2 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-358 | Alarm system | The Checksum Monitor PIC alarm shall have following parameters: Range: N/A, Autoset: N/A, Priority: High, Sequence: 2 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-359 | Alarm system | The Low Memory – please reboot! alarm shall have following parameters: Range: N/A, Autoset: N/A, Priority: High, Sequence: 2 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-360 | Alarm system | The Air supply pressure alarm shall have following parameters: Range: N/A, Autoset: Fixed threshold 200 kPa, Priority: High, Sequence: 2 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-361 | Alarm system | The Oxygen supply pressure alarm shall have following parameters: Range: N/A, Autoset: Fixed threshold 200 kPa, Priority: High, Sequence: 2 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-362 | Alarm system | The Safety relay defect alarm shall have following parameters: Range: N/A, Autoset: N/A, Priority: High, Sequence: 2 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-363 | Alarm system | The Tube occlusion alarm shall have following parameters: Range: See section 3.6.2, Autoset: N/A, Priority: High, Sequence: 2 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-364 | Alarm system | The Voltage Monitoring alarm shall have following parameters: Range: N/A, Autoset: N/A, Priority: High, Sequence: 2 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-365 | Alarm system | The Battery Defect alarm shall have following parameters: Range: N/A, Autoset: N/A, Priority: High, Sequence: 2 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-366 | Alarm system | The Exhalation Calibration alarm shall have following parameters: Range: N/A, Autoset: N/A, Priority: High, Sequence: 2 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-367 | Alarm system | The Check Blender alarm shall have following parameters: Range: N/A, Autoset: N/A, Priority: High, Sequence: 2 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-368 | Alarm system | The Patient Disconnect alarm shall have following parameters: Range: N/A, Autoset: N/A, Priority: High, Sequence: 3 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-369 | Alarm system | The Check ET tube alarm shall have following parameters: Range: N/A, Autoset: N/A, Priority: High, Sequence: 3 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-370 | Alarm system | The Clean flow sensor alarm shall have following parameters: Range: N/A, Autoset: N/A, Priority: High, Sequence: 4 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-371 | Alarm system | The Flow sensor defect alarm shall have following parameters: Range: N/A, Autoset: N/A, Priority: High, Sequence: 4 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-372 | Alarm system | The Flow sensor not connected alarm shall have following parameters: Range: N/A, Autoset: N/A, Priority: High, Sequence: 4 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-373 | Alarm system | The O2 sensor calibration failed alarm shall have following parameters: Range: N/A, Autoset: N/A, Priority: High, Sequence: 4 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-374 | Alarm system | The Oxygen sensor defect alarm shall have following parameters: Range: N/A, Autoset: N/A, Priority: High, Sequence: 4 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-375 | Alarm system | The Oxygen sensor used up alarm shall have following parameters: Range: N/A, Autoset: N/A, Priority: High, Sequence: 4 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-376 | Alarm system | The O2 value out of range alarm shall have following parameters: Range: N/A, Autoset: Measured O2 outside range, max(Set O2 – 5 %, 18%) to (Set O2 + 5 %) , Alarm delay: 60s, Priority: High, Sequence: 4 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-377 | Alarm system | The SpO2 module disconnected alarm shall have following parameters: Range: N/A, Autoset: N/A, Priority: High, Sequence: 5 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-378 | Alarm system | The SpO2 sensor failure alarm shall have following parameters: Range: N/A, Autoset: N/A, Priority: High, Sequence: 5 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-379 | Alarm system | The check SpO2 sensor alarm shall have following parameters: Range: N/A, Autoset: N/A, Priority: High, Sequence: 5 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-380 | Alarm system | The etCO2 FilterLine disconnected alarm shall have following parameters: Range: N/A, Autoset: N/A, Priority: High, Sequence: 5 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-381 | Alarm system | The etCO2 module disconnected alarm shall have following parameters: Range: N/A, Autoset: N/A, Priority: High, Sequence: 5 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-382 | Alarm system | The occlusion in etCO2 sampling line alarm shall have following parameters: Range: N/A, Autoset: N/A, Priority: High, Sequence: 5 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-383 | Alarm system | The etCO2 sensor faulty alarm shall have following parameters: Range: N/A, Autoset: N/A, Priority: High, Sequence: 5 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-384 | Alarm system | The Apnea alarm shall have following parameters: Range: 2 to 20 s, Off; Autoset: 10s; Priority: High; Sequence: 6 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-385 | Alarm system | The High PIP alarm shall have following parameters: Range, dual limb excluding CPAP: 10 to 90 cmH2O Range, CPAP: 10 to 40 mbar; Range, nCPAP + DuoPAP: 1 to 30 mbar; Autoset, dual limb excluding CPAP: Set PEEP + P-Insp + 5cmH2O; Autoset, CPAP: Set CPAP + 5cmH2O; Autoset, nCPAP + DuoPAP: Set CPAP + 5 mbar; Priority: High; Sequence: 7 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-386 | Alarm system | The High PEEP alarm shall have following parameters: Range: N/A (autoset), Autoset: Set PEEP + 6 cmH2O, Priority: High, Sequence: 7 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-387 | Alarm system | The Low SIQ alarm shall have following parameters: Range: Off, 5 to 100 %, Autoset: 50% of measured, Priority: High, Sequence: 7 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-388 | Alarm system | The Low PIP alarm shall have following parameters: Range: off, 0 to 89 cmH2O, Autoset: N/A, Priority: Medium, Sequence: 8 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-389 | Alarm system | The Charge Battery < 30 mins alarm shall have following parameters: Range: N/A, Autoset: N/A, Priority: Medium, Sequence: 8 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-390 | Alarm system | The Charge Battery < 60 mins alarm shall have following parameters: Range: N/A, Autoset: N/A, Priority: Medium, Sequence: 8 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-391 | Alarm system | The High Breath Rate alarm shall have following parameters: Range: 10 to 200 bpm, Off; Autoset: 150% of measured; Priority: Medium; Sequence: 8 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-392 | Alarm system | The High EtCO2 alarm shall have following parameters: Range: Off, 1 to 149 mmHg; Autoset: 150% of measured; Priority: Medium; Sequence: 8 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-393 | Alarm system | The Low EtCO2 alarm shall have following parameters: Range: 0 to 147 mmHg; Autoset: 50% of measured; Priority: Medium; Sequence: 8 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-394 | Alarm system | The High FiCO2 alarm shall have following parameters: Range: 3 to 149 mmHg; Off; Autoset: N/A; Priority: Medium; Sequence: 8 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-395 | Alarm system | The High Tube leak alarm shall have following parameters: Range: Off, 10 to 50 %; Autoset: 150% of measured; Priority: Medium; Sequence: 8 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-396 | Alarm system | The High Minute Volume alarm shall have following parameters: Range, NEO: 0.05 to 7.00 L; Range, PED: 0.50 to 10.0 L; Autoset: 180% of measured; Priority: Medium; Sequence: 8 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-397 | Alarm system | The Low Minute Volume alarm shall have following parameters: Range, NEO: Off, 0.01 to 6.9 L; Range, PED: Off, 0.01 to 9.9 L; Autoset: 50% of measured; Priority: Medium; Sequence: 8 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-398 | Alarm system | The Low PEEP alarm shall have following parameters: Range, dual limb excluding CPAP: 0 to 60 cmH2O; Range, CPAP: Off, 0 to 39 cmH2O; Range, nCPAP + DuoPAP: Off, 0 to 30 cmH2O; Autoset, dual limb excluding CPAP: Set PEEP – 3cmH2O; Autoset, CPAP: Set CPAP – 5cmH2O; Autoset, nCPAP + DuoPAP: Set CPAP– 3cmH2O; Priority: Medium; Sequence: 8 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-399 | Alarm system | The High Pulse Rate alarm shall have following parameters: Range: 35 to 235 bpm, Off; Autoset: 150% of measured; Priority: Medium; Sequence: 8 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-400 | Alarm system | The Low Pulse Rate alarm shall have following parameters: Range: Off, 30 to 230 bpm; Autoset: 50% of measured; Priority: Medium; Sequence: 8 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-401 | Alarm system | The High SpO2 alarm shall have following parameters: Range: 2 to 99 %, Off; Autoset: 150% of measured; Priority: Medium; Sequence: 8 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-402 | Alarm system | The Low SpO2 alarm shall have following parameters: Range: Off, 1 to 97 %; Autoset: 50% of measured; Priority: Medium; Sequence: 8 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-403 | Alarm system | The Low Perfusion Index alarm shall have following parameters: Range: Off, 0.03 to 19.0 %; Autoset: 50% of measured; Priority: Medium; Sequence: 8 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-404 | Alarm system | The Volume Not Reached alarm shall have following parameters: Range: Based on V-Guarantee, see controls; Autoset: N/A; Priority: Medium; Sequence: 8 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-405 | Alarm system | The Volume Not Reached alarm shall have following parameters: Range: Based on V-Guarantee, see controls; Autoset: N/A; Priority: Medium; Sequence: 8 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-406 | Alarm system | The Insp Press Not Reached alarm shall have following parameters: Range: Determined by system from PEEP control; Autoset: Low PEEP limit + 2 cmH2O; Priority: Low; Sequence: 9 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-407 | Alarm system | The Power Failure alarm shall have following parameters: Range: N/A; Autoset: N/A; Priority: Low; Sequence: 9 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-xxx | Alarm system | The Tidal Volume Limit alarm shall have following parameters: Range: Determined by V-Limit, see controls; Autoset: N/A; Priority: Low; Sequence: 9 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-xxx | Alarm system | The Acoustic audio alarm test failed alarm shall have following parameters: Range: N/A; Autoset: N/A; Priority: High; Sequence: 1 | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-xxx | Alarm system | The High DCO2 alarm shall have following parameters: Range: 2 to 10,000 mL2/sec, Off; Autoset: N/A; Priority: xxx; Sequence: xxx | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-xxx | Alarm system | The Low DCO2 alarm shall have following parameters: Range: Off, 1 to 9,900 mL2/sec, Off; Autoset: N/A; Priority: xxx; Sequence: xxx | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-xxx | Alarm system | The High Pmean alarm shall have following parameters: Range: 1 to 55 mbar; Autoset: N/A; Priority: xxx; Sequence: xxx | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-xxx | Alarm system | The Low Pmean alarm shall have following parameters: Range: 0 to 54 mbar; Autoset: N/A; Priority: xxx; Sequence: xxx | incorrect alarm configuration | device alarm system issue | PRS
FAB-H-REQ-408 | Alarm system | In the event of a pneumatic fault (oxygen supply pressure, air supply pressure, tube occlusion, blender fault, or exhalation valve fault) the site of the fault shall be indicated graphically on the user interface annotated onto a pneumatic diagram. | fail to indicate site of fault graphically | no display or display failure | PRS
FAB-H-REQ-409 | Alarm system | The ventilator shall provide switch to using the alternative gas in the event of one supply falling below the low inlet pressure alarm limit. | lack of backup gas | medical gas supply problem | PRS
FAB-H-REQ-503 | Alarm system | The ventilator shall provide an alarm indication for occlusion of the inspiratory, expiratory or proximal pressure sense lines. | fail to activate occlusion alarm | device alarm system issue | PRS
FAB-H-REQ-504 | Alarm system | In the event of activation of the occlusion alarm, the ventilator shall terminate outlet flow for a period of at least 1s , prior to attempting to resume ventilation. | fail to disable ventilation | device remains activated | PRS
FAB-H-REQ-410 | Alarm system | In the event of a blender fault, the Check Blender alarm shall be accompanied by a code indicating the nature of the fault. | fail to indicate site of fault graphically | no display or display failure | PRS
FAB-H-REQ-505 | Alarm system | The ventilator shall use a dual oxygen sensor. | fail to convert to back-up | failure to convert to back-up | PRS
FAB-H-REQ-506 | Alarm system | On activation of the oxygen sensor fault alarm, the operator shall be provided means to either re-calibrate, or to disable , the oxygen sensor. | fail to recalibrate | failure to recalibrate | PRS
FAB-H-REQ-411 | Alarm system | On activation of the flow sensor fault alarm, the operator shall be provided means to either re-calibrate, or to disable , the proximal flow sensor. | fail to disable sensor | device remains activated | PRS
FAB-H-REQ-412 | Alarm system | The ventilator shall shut down no sooner than one minute following activation of the 'flat battery' alarm. | slow shut down | failure to shut off | PRS
FAB-H-REQ-512 | Alarm system | The ventilator shall switch to using the internal battery without interruption to ventilation on activation of the Power Failure alarm. | battery switched inappropriately | battery issue | PRS
FAB-H-REQ-91 | Alarm system | The ventilator shall provide High Priority alarm indications for the following capnometer fault conditions: etCO2 module disconnected etCO2 filter line disconnected Occlusion in etCO2 sampling line etCO2 sensor faulty | fail to raise high priority alarm | device alarm system issue | PRS
FAB-H-REQ-536 | Alarm system | All capnometry measurements (numerics and graphical) shall be disabled during any EtCO2 module fault or disconnect alarm state. | fail to disable measurement | device remains activated | PRS
FAB-H-REQ-92 | Alarm system | The ventilator shall provide alarm indications for the following oximeter fault conditions (sub-bullets are additional informational messages, not discrete alarm states): SpO2 Module Disconnected (High Priority); SpO2 Sensor Failure (High Priority); spO2 Check Sensor (Medium Priority) | fail to raise high priority alarm | device alarm system issue | PRS
FAB-H-REQ-93 | Alarm system | The ventilator shall provide a technical alarm indication (using a red alarm lamp) in the event of any system fault that prevents the display of IEC 60601-1-8 compliant visual alarm indication (for example, failure of the NetDCU processor, alarm PIC, or their power supply rails). | fail to generate visual alarm | device alarm system issue | PRS
FAB-H-REQ-94 | Alarm system | All indicated alarms are recorded into an alarm log file, together with the alarm priority and a timestamp. | inadequate log file | patient data issue | PRS
FAB-H-REQ-95 | Alarm system | The system shall provide capability to review the content of the alarm log. | missing data | data issue | PRS
FAB-H-REQ-96 | Alarm system | The ventilator shall provide storage for up to 1,000 alarm messages. | insufficient storage space | inadequate storage | PRS
FAB-H-REQ-97 | Alarm system | The alarm log shall be preserved during power failure. | unable to record event | data issue | PRS
FAB-H-REQ-98 | Patient Data | The ventilator user interface shall incorporate a patient data screen with provision for entry of the following patient data: Patient's hospital identifier; Patient's date of birth; Patient's bodyweight; Patient's full name; Remark, alphanumeric field for free text notes | inadequate patient data | patient data issue | PRS
FAB-H-REQ-99 | Patient Data | The patient data screen shall incorporate means to reset all patient data to blank. | fail to reset patient data | patient data issue | PRS
FAB-H-REQ-100 | Patient Data | The patient data screen shall incorporate means to accept data entered by the operator. | fail to enter patient data | patient data issue | PRS
FAB-H-REQ-101 | Patient Data | On accepting changes to the patient data, the ventilator shall prompt the user to confirm whether the stored trend data is to be deleted. | undesired patient data changes | patient data issue | PRS
FAB-H-REQ-102 | Patient Data | On confirmation by the operator, the stored trend data shall be deleted. | fail to delete data | data issue | PRS
FAB-H-REQ-103 | Patient Data | The patient data screen shall incorporate means to cancel changes to the patient data. | unauthorized deletion | unauthorized access to computer system | PRS
FAB-H-REQ-104 | Patient Data | The patient data screen shall incorporate means for the operator to change the data displayed. | unauthorized deletion | unauthorized access to computer system | PRS
FAB-H-REQ-105 | Patient Data | The patient data reset function shall reset the display of trend data so that only data collected after the reset is visible to the operator. | fail to reset patient data | patient data issue | PRS
FAB-H-REQ-106 | Patient Data | The patient data reset function shall be actioned on powering on the ventilator. | fail to reset patient data | patient data issue | PRS
FAB-H-REQ-122 | User Interface | The ventilator shall incorporate a user interface with a color graphical display with touch screen control, a rotary control, and dedicated keys. | inadequate user interface | inadequate user interface | PRS
FAB-H-REQ-123 | User Interface | The ventilator shall have dedicated lamps to indicate the following functions: System Fault, Battery Power (operating from internal battery), Battery Charging Status (not charging; charging; fully charged), Mains Power (operating from external AC supply) | confused indication | inadequate user interface | PRS
FAB-H-REQ-124 | User Interface | The ventilator shall have a set of ten dedicated membrane keys for dedicated functions. | indistinct button | human-device interface issue | PRS
FAB-H-REQ-125 | User Interface | The ventilator user interface shall incorporate dedicated indication (icons or text) for the following: Patient size (neonatal or pediatric); Patient data entered; Nurse call active; Battery charge status; Ventilation mode; Alarm audio pause remaining time (if active); Screen lock (if active) | indistinct indication | incorrect display | PRS
FAB-H-REQ-507 | Connections, inputs and outputs | The ventilator outlet shall be a 22 mm male / 15 mm female coaxial conical connector per ISO 5356-1:2015. | diameter out of specification | defective component | PRS
FAB-H-REQ-508 | Connections, inputs and outputs | The ventilator shall have means to detect an obstruction at the primary outlet. | fail to detect obstruction | occlusion within device | PRS
FAB-H-REQ-509 | Connections, inputs and outputs | A secondary outlet (the 'FG' or 'Fresh Gas' outlet) shall be provided for connection to single limb circuits. | inadequate connection | loose or intermittent connection | PRS
FAB-H-REQ-510 | Connections, inputs and outputs | The secondary outlet shall be a hose barb connector, with nominal OD 5.3 mm, hose barb OD 6.3 mm. | diameter out of specification | loose or intermittent connection | PRS
FAB-H-REQ-511 | Connections, inputs and outputs | The exhalation valve shall incorporate a removable polycarbonate diaphragm holder that includes all parts wetted by the exhaled patient gases. | valve stuck | difficult to open or close | PRS
FAB-H-REQ-512 | Connections, inputs and outputs | The exhalation valve diaphragm holder shall incorporate a 22 mm male / 15 mm female coaxial conical connector per ISO 5356-1:2015 for connection from the breathing circuit expiratory limb. | connector mismatched | fitting problem | PRS
FAB-H-REQ-513 | Connections, inputs and outputs | The ventilator shall incorporate a pressure sensor to monitor the pressure at the patient airway interface using a proximal pressure line. | sensor measurement interfered | incorrect measurement | PRS
FAB-H-REQ-514 | Connections, inputs and outputs | The connector for the proximal pressure line shall be a cylindrical hose connector, with nominal OD 4.0 mm. | diameter out of specification | defective component | PRS
FAB-H-REQ-515 | Connections, inputs and outputs | The ventilator shall provide purge flow to the proximal pressure line. | inappropriate purge flow | device cleaning issue | PRS
FAB-H-REQ-516 | Connections, inputs and outputs | The ventilator shall incorporate means to detect a faulty proximal pressure sensor. | fail to detect sensor error | failure to sense | PRS
FAB-H-REQ-413 | Connections, inputs and outputs | The port for connection of the proximal flow sensor shall be a uniquely keyed, proprietary, dedicated latching connector. | undedicated connector | misconnection | PRS
FAB-H-REQ-414 | Connections, inputs and outputs | The port for connection of a capnometer module shall be a uniquely keyed, proprietary, dedicated latching connector with shrouded contacts. | undedicated connector | misconnection | PRS
FAB-H-REQ-415 | Connections, inputs and outputs | The port for connection of a pulse oximeter module shall be a uniquely keyed, proprietary, dedicated latching connector. | undedicated connector | misconnection | PRS
FAB-H-REQ-416 | Connections, inputs and outputs | A USB-A socket shall be provided for connection of the nebulizer, or a FLASH memory device. | incompatible connection | connection issue | PRS
FAB-H-REQ-126 | Ventilator mass | The fabian ventilator shall have a mass not exceeding 16 kg (excluding cart and accessories). | mass out of specification | use of device issue | PRS
FAB-H-REQ-127 | Ventilator mass | The fabian ventilator shall have dimensions not exceeding 30 cm (w) x 37 cm (h) x 42 cm (d). | out of dimension | use of device issue | PRS
FAB-H-REQ-136 | Ventilator mass | The ventilator shall incorporate an RS232 port for connection to a PDMS. | port mismatched | component incompatible | PRS
FAB-H-REQ-137 | Ventilator mass | The RS232 port shall incorporate a 5V DC output capable of providing a minimum of 500 mA. | voltage out of range | electrical issue | PRS
FAB-H-REQ-138 | Ventilator mass | The ventilator shall incorporate an Ethernet port, with a tamper-proof cover. | port mismatched | fitting problem | PRS
FAB-H-REQ-139 | Ventilator mass | The ventilator shall provide means to communicate with a PDMS using the RS232 port using the Philips Intellibridge/VueLink protocol. | loss of communication | Communication or Transmission Issue | PRS
FAB-H-REQ-140 | Ventilator mass | The ventilator shall provide means to communicate with a PDMS using the RS232 port using a proprietary protocol. | loss of communication | Communication or Transmission Issue | PRS
FAB-H-REQ-142 | Ventilator mass | The ventilator shall incorporate a dedicated port for connection to a Nurse Call system, with normally open and normally closed contacts, enabled during activation of any alarm. | port mismatched | fitting problem | PRS
FAB-H-REQ-143 | Ventilator mass | The Nurse Call connection shall be compatible with switching up to 1A DC or AC, at signaling voltages up to 30 VDC. | unqualified switch | electrical issue | PRS
FAB-H-REQ-14 | Ventilator mass | The ventilator shall support the use of a FLASH memory device connected to the USB port, for export of trend data, alarm logs, and system information. | incompatible accessory | accessory incompatible | PRS
FAB-H-REQ-15 | Ventilator mass | The ventilator shall provide an HDMI video output, labelled as not for clinical use. | loss of video output | output issue | PRS
FAB-H-REQ-149 | Ventilator mass | The ventilator shall provide means for the user to specify patient size as either neonatal or pediatric. | fail to specify patient size | patient data issue | PRS
FAB-H-REQ-150 | Ventilator mass | The ventilator shall provide means for the user to specify units of measure for capnometry values as mmHg or kPa (default). | inappropriate measure unit | incorrect measurement | PRS
FAB-H-REQ-151 | Ventilator mass | The ventilator shall provide means for the user to specify the local time and date. | lack of time indication | no display or display failure | PRS
FAB-H-REQ-152 | Ventilator mass | The ventilator shall provide means for the user to specify the altitude at the site of use, to allow for barometric compensation for controlled and measured values, with range 0 to 10,000 feet. | lack of barometric compensation | imprecision | PRS
FAB-H-REQ-153 | Ventilator mass | The ventilator shall provide means for the user to specify the maximum duration for a manual breath, in the range 2 to 30 s, with accuracy ± 1s. | inappropriate maximum duration | use of incorrect control settings | PRS
FAB-H-REQ-154 | Ventilator mass | The ventilator shall provide means for the user to specify the maximum duration for the oxygen flush function, as Off, 30, 60 or 120 s. | inappropriate maximum duration | use of incorrect control settings | PRS
FAB-H-REQ-16 | Ventilator mass | The ventilator shall provide means for the user to store and retrieve the current control settings as 'Hospital Settings'. | fail to retrieve data | data issue | PRS
FAB-H-REQ-8 | Enviornmental and safety | The ventilator shall conform to specifications when used within the following environmental conditions: Temperature +10 to +40 °C; Relative Humidity 10 to 90%, non-condensing; Ambient pressure 70 to 106 kPa (10.2 to 15.0 PSIA) | temperature out of range | temperature issue | PRS
FAB-H-REQ-107 | Enviornmental and safety | The ventilator shall conform to specifications following shipping and transport within the following environmental conditions: Temperature -20 to +60 °C; Relative Humidity 10 to 90%, non-condensing; Ambient pressure 50 to 106 kPa (7.2 to 15.0 PSIA) | inappropriate shipping environment or condition | environmental control or utility issue | PRS
FAB-H-REQ-108 | Enviornmental and safety | Each ventilator shall be provided with a shipping carton. | shipping damage | break | PRS
FAB-H-REQ-109 | Enviornmental and safety | The shipping carton shall be labelled to show limit on stacking height, environmental conditions for shipping and storage. | inadequate labeling | device markings issue | PRS
FAB-H-REQ-110 | Enviornmental and safety | The ventilator in its intended shipping carton or packaging shall conform to the required tested by ASTM D4169 (Distribution Cycle DC13, Assurance Level 1) or to ISTA 2A. | shipping test failed | failure of device to self-test | PRS
FAB-H-REQ-111 | Enviornmental and safety | The ventilator shall conform to specifications when used with gas supplies within the following range of inlet pressure: Inlet pressure 200 to 600 kPa (29 to 87 PSIG) | pressure out of range | pressure issue | PRS
FAB-H-REQ-112 | Enviornmental and safety | The ventilator shall be provided with a particulate filter at each gas inlet. | filter mismatched | fitting problem | PRS
FAB-H-REQ-113 | Enviornmental and safety | Pressure regulators at each gas inlet shall be rated to at least 10 bar inlet pressure. | inappropriate rated pressure | pressure issue | PRS
FAB-H-REQ-114 | Enviornmental and safety | The gas inlets for oxygen and air shall use gas-specific DISS connectors. | connector mismatched | component incompatible | PRS
FAB-H-REQ-115 | Enviornmental and safety | The ventilator shall be designed to draw no more than 40 lpm (liter per minutes) from either inlet connection. | flow rate overspeed | excess flow or over-infusion | PRS
FAB-H-REQ-116 | Enviornmental and safety | A water trap shall be available for use at the air inlet. | water splashed | fluid leak | PRS
FAB-H-REQ-117 | Enviornmental and safety | The ventilator shall conform to specifications when used with a wall AC supply within the following rated range: Voltage 100 to 240 VAC, 0.9 to 0.5 A; Frequency 50 / 60 Hz; Power≤ 70 VA (operating); Power≤ 35 VA (standby, charging internal battery) | inappropriate rated AC power | electrical power problem | PRS
FAB-H-REQ-118 | Enviornmental and safety | The wall AC inlet connector shall be a socket for a power cord with an IEC 60320 connector and a retention clip. | lack of retention clip | defective component | PRS
FAB-H-REQ-119 | Enviornmental and safety | The ventilator shall incorporate visual indication of the status of AC power connection. | fail to indicate AC power | no display or display failure | PRS
FAB-H-REQ-120 | Enviornmental and safety | The ventilator shall incorporate a rechargeable internal battery capable of run time of at least 2 hours. | reduced battery capacity | Low Battery | PRS
FAB-H-REQ-121 | Enviornmental and safety | The battery shall be capable of recharging from fully discharged in no more than 5 hours. | fail to charge battery to full | battery issue | PRS
FAB-H-REQ-128 | Enviornmental and safety | The fabian ventilator shall have ingress protection classification IP21 in accordance with IEC 60529:1989+A1:1999+A2:2013. | water proof failed | electrical shorting | PRS
FAB-H-REQ-129 | Enviornmental and safety | The fabian ventilator shall be electrical safety Class I / Internally Powered (protectively earthed when connected to wall power, internally powered when not connected to external power). | lack of earth protective | electrical shorting | PRS
FAB-H-REQ-130 | Enviornmental and safety | The patient applied parts shall have isolation Type BF | loss of electrical isolation | electrical shorting | PRS
FAB-H-REQ-9 | Enviornmental and safety | Excluding audible alarm indication, the fabian ventilator shall produce sound pressure level not exceeding 52 dB(A) in normal operation. | sound pressure level out of range | audible noise | PRS
FAB-H-REQ-10 | Enviornmental and safety | The ventilator shall incorporate a mechanical pressure relief to prevent the pressure at the ventilator outlet from exceeding 125 cmH2O. | fail to relieve pressure properly | pressure issue | PRS
FAB-H-REQ-11 | Enviornmental and safety | Electronic pressure relief shall be provided during periods of active Tube Occlusion alarm, see section 3.6.2 | fail to relieve pressure properly | pressure issue | PRS
FAB-H-REQ-12 | Enviornmental and safety | In the event of loss of pressure at one gas inlet, such that the required flow for that gas cannot be provided, the ventilator shall switch to ventilation with the remaining inlet gas. | fail to switch backup gas | medical gas supply problem | PRS
FAB-H-REQ-13 | Enviornmental and safety | In the event of loss of software control, the ventilator shall disable the blender flow control valves. | fail to control valve | improper flow or infusion | PRS
FAB-H-REQ-xxx | Enviornmental and safety | Total system resistance as measured at the patient connection port shall comply with following specifications: | resistance out of spec | physical resistance | PRS
FAB-H-REQ-131 | Enviornmental and safety | The inspiratory resistance when using a pediatric dual limb breathing circuit, at 15 lpm shall not exceed 6 hPa (i.e. 24 hPa/L/s) when measured at the patient connection port. | inspiratory resistance out of spec | physical resistance | PRS
FAB-H-REQ-132 | Enviornmental and safety | The expiratory resistance when using a pediatric dual limb breathing circuit, at 15 lpm shall not exceed 6 hPa (i.e. 24 hPa/L/s) when measured at the patient connection port. | expiratory resistance out of spec | physical resistance | PRS
FAB-H-REQ-133 | Enviornmental and safety | The inspiratory resistance when using an infant dual limb breathing circuit, at 2.5 lpm shall not exceed 6 hPa (i.e. 144 hPa/L/s) when measured at the patient connection port. | Inspiratory resistance out of spec | physical resistance | PRS
FAB-H-REQ-134 | Enviornmental and safety | The expiratory resistance when using an infant dual limb breathing circuit, at 2.5 lpm shall not exceed 6 hPa (i.e. 144 hPa/L/s) when measured at the patient connection port. | expiratory resistance out of spec | physical resistance | PRS
FAB-H-REQ-523 | Enviornmental and safety | The ventilator shall verify all sensors and their calibration as part of the start-up self test, and prevent ventilation if fault detected. | out of calibration | calibration issue | PRS
FAB-H-REQ-135 | Enviornmental and safety | The ventilator shall implement a self test at startup that allows the operator to confirm correct operation of the following: The audible alarm indicator; The visual alarm lamp; the nurse call indicator (if connected) | lack of confirmation | use of device issue | PRS
FAB-H-REQ-167 | Enviornmental and safety | The ventilator shall be designed for an operating life of 40,000 hours, or eight years (whichever is lesser) assuming all planned maintenance is carried out per the published schedule. | reduced service life | shelf life exceeded | PRS
FAB-H-REQ-168 | key components | The ventilator shall be designed for continuous operation for a period of at least 30 days. | reduced service life | shelf life exceeded | PRS
FAB-H-REQ-169 | key components | The blender construction shall prevent reverse gas flow from the outlet manifold into the blender. | reversed gas flow | improper flow or infusion | PRS
FAB-H-REQ-170 | key components | All internal cable connections shall incorporate means for mechanical retention. | weak mechanical strength | mechanical issue | PRS
FAB-H-REQ-171 | key components | All PCBAs and other internal components shall be mounted with vibration-proof fasteners (for example, screws with spring washers or thread-lock). | fastener loose due to vibration | detachment of device component | PRS
FAB-H-REQ-172 | key components | There shall be a cooling fan mounted between the power electronics and the blender and designed to push air out of the enclosure. | fan detachment | detachment of device component | PRS
FAB-H-REQ-173 | key components | The cooling fan outlet shall be protected with a grille. | fan broken or damaged | break | PRS
FAB-H-REQ-174 | key components | All mains parts shall have dual insulation to the enclosure and ventilator electronic control circuits. | insulation broken or damaged | electrical shorting | PRS
FAB-H-REQ-175 | key components | The internal AC/DC power supply shall be a safety-marked component. | missing safety mark | device markings issue | PRS
FAB-H-REQ-527 | key components | The powerboard shall be classified as a high integrity component subject to purchasing controls. | circuit board broken or damaged | circuit failure | PRS
FAB-H-REQ-526 | key components | The rechargeable battery shall be classified as a high integrity component subject to purchasing controls. | charge cycles out of specification | charging issue | PRS
FAB-H-REQ-525 | key components | The nurse call relay shall be classified as a high integrity component subject to purchasing controls. | reduced life cycle | shelf life exceeded | PRS
FAB-H-REQ-528 | key components | The pressure regulators shall be classified as high integrity components subject to purchasing controls. | fail to inform regulator on problem | Communication or Transmission Issue | PRS
FAB-H-REQ-529 | key components | The proportional valves shall have proof pressure at least 10 bar, and be classified as high integrity components subject to purchasing controls. | low proof pressure | pressure issue | PRS
FAB-H-REQ-531 | key components | The pressure sensors shall be classified as high integrity components subject to purchasing controls. | low measurement accuracy | imprecision | PRS
FAB-H-REQ-532 | key components | The membrane switch assembly shall be classified as a high integrity component subject to purchasing controls. | self-activation | self-activation or keying | PRS
FAB-H-REQ-540 | key components | All membrane switches shall have edge-triggered activation. | lack of edge-triggered activation | difficult or delayed activation | PRS
FAB-H-REQ-176 | Enviornmental and safety | The ventilator shall provide over-voltage protection to protect from power regulator failure. | lack of over-voltage protection | protective measure issue | PRS
FAB-H-REQ-519 | Enviornmental and safety | The ventilator shall provide isolation to the external SIP/SOPs (RS232, Ethernet, and Nurse Call) to meet requirements of IEC 60601-1. | insulation broken or damaged | electrical shorting | PRS
FAB-H-REQ-520 | Enviornmental and safety | The blender flow sensors shall be hot-wire type sensors. | inaccurate sensor measurement data | incorrect measurement | PRS
FAB-H-REQ-192 | Regulatory | All parts of the fabian ventilator that contact the patient, patient cells or fluids, the user, or the breathing gas shall be evaluated to ISO 10993-1 and ISO 18562-1 for biocompatibility. | material bio-incompatibility | biocompatibility issue | PRS
FAB-H-REQ-193 | Regulatory | The ventilator shall comply with the requirements of FDA guidance "Final Guidance on "Use of International Standard ISO 10993-1, Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process". | material biocompatibility compromised | biocompatibility issue | PRS
FAB-H-REQ-109 | packaging and labelling | The shipping carton shall be labelled to show limit on stacking height, environmental conditions for shipping and storage. | inadequate labeling | device markings issue | PRS
FAB-H-REQ-518 | packaging and labelling | The fabian ventilator shall be labelled as intended for hospital and institutional healthcare facility use, including during intra-hospital transport. | inappropriate usage | user used incorrect product for intended use | PRS
FAB-H-REQ-522 | packaging and labelling | The ventilator shall be labelled as intended for use in settings in which vital signs monitoring is mandatory standard of care. | inadequate labeling | user used incorrect product for intended use | PRS
FAB-H-REQ-177 | packaging and labelling | The ventilator shall be provided together with printed Instructions for Use. | missing IFU in package | Instruction for Use Issue | PRS
FAB-H-REQ-178 | packaging and labelling | Any safety sign applied to the ventilator shall comply to ISO 7010. | fail to comply with standards or regulation | device markings issue | PRS
FAB-H-REQ-179 | packaging and labelling | Marking on the device shall include a safety sign that the video output is not for use on patients. | missing safety sign | device markings issue | PRS
FAB-H-REQ-180 | packaging and labelling | Marking on the device shall include a safety sign to connect to wall power when not in transit use. | missing safety sign | device markings issue | PRS
FAB-H-REQ-181 | packaging and labelling | The ventilator shall be marked with the "wheelie bin" symbol denoting "Disposal as WEEE". | missing disposal mark WEEE | device markings issue | PRS
FAB-H-REQ-182 | packaging and labelling | The instructions for use shall include specification of planned maintenance service intervals. | lack of maintenance interval | lack of maintenance documentation or guidelines | PRS
FAB-H-REQ-xxx | packaging and labelling | The ventilator shall be marked with the ISO 7010-M002 symbol, “Refer to instruction manual / booklet” | inadequate markings | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The ventilator labels shall include any particular storage and/or handling instructions. Specific instructions: labeling for ingress protection, labeling of mass of product | inadequate instruction | Instruction for Use Issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The ventilator labels shall include any relevant warnings and cautions. Specific warnings and cautions: do not obstruct chassis fan exhaust, read instructions (symbol) | inadequate warning | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The ventilator shall have a label specifying prescription use only (e.g. “RX Only”). | inappropriate usage | use of device issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The ventilator shall have a UDI label | invalid UDI | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The ventilator shall be labeled with the rated range of supply voltage and frequency. | inappropriate power supply | Power Source Issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The ventilator shall be labeled with the range of rated input, specified in amperes. | inappropriate current supply | Power Source Issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The ventilator shall be labeled with the ingress protection classification according to IEC 60529. | incorrect ingress classification | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | Ports for connection of the following applied parts shall be labeled with the symbols for Type BF Applied Part, as defined in IEC 60417-5333: HWFS, EtCO 2 sensor, Pulse oximetry sensor, Nebulizer | incorrect labeling | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | Each of the high-pressure gas inlet ports shall be labeled with the gas name or chemical symbol “AIR” for medical nominal 4 bar, and “O2” for oxygen). | fail to label gas name | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The ventilator shall be labeled with the following rated supply pressure ranges: O2 : 200 - 600 kPa (29 - 87 psig) ; Air: 200 - 600 kPa (29 - 87 psig) | fail to label pressure range | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The ventilator shall be labeled with the rated flow requirement for both high-pressure gas inlets. | fail to label gas name | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The ventilator inspiratory gas outlet shall be labeled with the text, “Insp” | fail to label text | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The ventilator proximal pressure port shall be labeled with the text, “Prox”. | fail to label text | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The exhaled gas inlet port shall be marked with the text, “Exp”. | fail to label text | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The ventilator cooling fan exhaust shall be labeled with the ISO 7000-0434 symbol, 'Caution' and the IEC 60417-5035 symbol, 'Output'. | fail to label text | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The ventilator shall be labeled with its mass (excluding cart and accessories). | fail to label mass | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The ventilator shall be labeled with the name or trademark and contact information for the manufacturer, together with ISO 7000-3082 symbol for 'Manufacturer'. | fail to label manufacturer information | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The ventilator shall be labeled with a type reference (part number), serial number and the date of manufacture. | missing product serial number | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The ventilator label(s) shall allow space for the addition of NRTL scheme certification (i.e. ETL mark) labels. | lack of ETL mark | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The ventilator shall be labeled with reference to information for proper disposal to comply with European Directive 2002/96/EC, Waste Electrical and Electronic Equipment. | inproper disposal | environmental control or utility issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The ventilator shall be labeled with the type and full rating for the fuse-holder(s). | fail to label rating | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The Ethernet & RS-232 ports shall be labeled with the IEC 60417-5850 symbol for 'Serial Interface'. | fail to label symbol | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The nurse call connector port shall be labeled with the IEC 60417-5307 symbol for “alarm, general”. | fail to label alarm symbol | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The nebulizer output port shall be labeled with the text, 'NEB'. | fail to label text | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The HDMI port shall be labeled with the text, 'Video Out'. | fail to label text | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The HDMI port shall be labeled with the text, 'Service Only'. | fail to label text | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The HDMI port shall be marked with the ISO 7010-M002 symbol, “Refer to instruction manual / booklet”. | fail to label standard symbol | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The equipotential connection shall be labeled with the IEC 60417-5021 symbol for 'Equipotentiality'. | fail to label equipotentiality label | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The HWFS connector port shall be labeled with the text, 'Flow sensor'. | fail to label text | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The connector port for pulse oximeter connection shall be labeled with the text, 'SpO2'. | fail to label SpO2 text | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The connector port for capnometer connection shall be labeled with the text, 'EtCO2 '. | fail to label EtCO2 text | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The Warning LED shall be labeled with the ISO 7010-W001 symbol for 'General warning sign'. | fail to label symbol | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The Battery LEDs shall be labeled with the ISO 7000-0247 symbol for “Battery charging”;. | fail to label symbol | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The Mains LED shall be labeled with the ISO 60417-5534 symbol for 'Power plug'. | fail to label symbol | device markings issue | PRS
FAB-H-REQ-xxx | Market | The ventilator shall be marked with the proprietary symbol for Masimo Signal Extraction Technology (SET®). | fail to label proprietary symbol | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The following operator-detachable components of the ventilator shall be marked with the name or trademark of the manufacturer and a model or type reference: Exhalation valve diaphragm holder , Power cord | fail to label manufacturer information | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The Manual Breath membrane button shall be marked with the following icon: | fail to mark icon | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The O2 Flush membrane button shall be marked with the following icon: | fail to mark icon | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The Menu/Calibration membrane button shall be marked with the following icon: | fail to mark icon | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The Nebulizer membrane button shall be marked with the following icon: | fail to mark icon | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The Graphics membrane button shall be marked with the following icon: | fail to mark icon | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The Home membrane button shall be marked with the following icon: | fail to mark icon | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The Alarm Limits membrane button shall be marked with the following icon: | fail to mark icon | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The On/Off membrane button shall be marked with the following symbols: IEC 60417-5264 symbol for 'ON', IEC 60417-5265 symbol for 'OFF' | fail to mark symbol | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The Mute Alarm membrane button shall be labeled with the following icon: | fail to mark icon | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The Start/Stop membrane button shall be labeled with the following icon: | fail to mark icon | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The UDP shall be marked with the ISO 7010-M002 symbol for 'Refer to instruction manual/booklet'. | fail to mark standard symbol | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The UDP shall be labeled with the following icon: | fail to mark icon | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The ventilator internal battery shall be marked with the ISO 7010-M002 symbol for 'Refer to instruction manual/booklet'. | fail to mark standard symbol | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The ground post internally to the ventilator at the AC inlet shall be labeled as a protective grounding terminal using the IEC 60417-5019 symbol. | fail to label standard symbol | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | Accessories shall be labeled with the name or trademark and contact information for the manufacturer. | fail to label manufacturer information | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | Accessories shall be labeled with the ISO 7000-2492 symbol for 'Batch code' or the ISO 7000-2498 symbol for 'Serial number' with a batch or serial number, respectively, and the ISO 7000-2493 symbol for 'Catalogue number' with a type reference (part number or model designator). | fail to label serial number | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | Accessories shall be labeled with the ISO 7000-2497 symbol for 'date of manufacture', or the ISO 7000-2607 symbol for 'Use by date'. | fail to label date of manufacture | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The following accessory components shall each have a UDI label: VBS filters, filter kit,HWFS | fail to label UDI | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The HWFS shall be marked with the direction of flow. | fail to mark flow direction | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | All markings and labels shall be clearly legible. | illegible label | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | All markings and labels shall be sufficiently durable to remain clearly legible during the expected service life of the product. | label smeared | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The meaning of all symbols and safety signs shall be explained in the instructions for use. | sign loss of explanation | device markings issue | PRS
FAB-H-REQ-xxx | packaging and labelling | Packaging for all accessories shall be marked with a description of the contents. | missing content description | Packaging Issue | PRS
FAB-H-REQ-xxx | packaging and labelling | All shipping packages for the ventilator and accessories shall have UDI labeling appropriate to the package and content. | fail to label UDI | Packaging Issue | PRS
FAB-H-REQ-xxx | packaging and labelling | Packaging for accessories intended for single-use shall be marked with the at least one of the following words: 'SINGLE USE' , 'DO NOT REUSE', ISO 7000-1051 symbol for 'Do not reuse'. | fail to label DO NOT REUSE | Packaging Issue | PRS
FAB-H-REQ-xxx | packaging and labelling | Packaging for accessories shall include the following symbols: ISO 7000-2492 symbol for 'Batch code' or ISO 7000-2498 symbol for 'Serial number' with a batch or serial number, respectively; ISO 7000-2493 symbol for 'Catalogue number with a type reference (part number or model designator). | fail to label serial number | Packaging Issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The shipping carton for the ventilator shall be labeled to show limit on stacking height. | excessive stack | unstable | PRS
FAB-H-REQ-xxx | packaging and labelling | Unit packaging (for shipping) of the ventilator shall include the permissible environmental conditions along with the following symbols: ISO 7000-0621 symbol for 'Fragile', ISO 7000-0632 symbol for 'Temperature limit', ISO 7000-2620 symbol for 'Humidity limit', ISO 7000-2621 symbol for 'Atmospheric pressure limit'. | fail to label Atmospheric pressure limit | Packaging Issue | PRS
FAB-H-REQ-xxx | packaging and labelling | Packaging for a replacement battery shall include the UN 3481 symbol and the manufacturer's contact information (e.g. telephone number). | fail to label standard symbol | Packaging Issue | PRS
FAB-H-REQ-xxx | packaging and labelling | Each ventilator saleable unit packaging shall include an IFU in print form, as appropriate for the Intended Market. | missing IFU in package | Packaging Issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall quantify and state any degradation in performance or accuracy specification at altitude. | loss of function due to altitude change | use of device issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall disclose the rated range of supply voltage and frequency for the input from the supply mains. | fail to disclose input voltage and frequency | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall disclose the rated input, specified in amperes, volt-amperes, or (only if the power factor exceeds 0.9) watts. | fail to disclose input current | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall disclose the rated range of pressure at each high-pressure gas inlet. | fail to disclose pressure range | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall disclose the 10 seconds average input flow required by the ventilator for each gas at a pressure of 280 kPa, measured at the gas input port. | fail to disclose input flow | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall disclose the maximum transient input flow averaged for 3 seconds required by the ventilator for each gas at a pressure of 280 kPa, measured at the gas input port. | fail to disclose input flow | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall include disclosure of the measurement uncertainty for each disclosed tolerance in the technical descriptions. | fail to disclose measurement accuracy | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | If marking any accessory (with indication of any limitations or adverse effects on the ventilator performance) is not practicable, this information shall be placed in the IFU. | fail to disclose accessory limitation | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall include a statement that latex is not incorporated into any materials that contact the patient gas pathway or the accessible surfaces of the device or proprietary accessories. | fail to disclose toxic material | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | All disclosures of performance specification shall use SI units, except that the units listed below may be used as alternative: Respiratory pressure: mbar; cmH2O; Volume: L or mL; Time: hour, minutes, day; Pressure of inlet gases: bar, kPa, psig; O2 concentration: %; CO2 concentration: mmHg or kPa. | fail to disclose specification | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | All performance specifications relating to gas flow or volume within the ventilator breathing system shall be expressed under BTPS conditions. | fail to disclose specification | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | All performance specifications relating to gas flow or volume not within the ventilator breathing system shall be expressed under STPD20 conditions. | fail to disclose specification | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | If there are any labels for which there is insufficient space to place the affirmative statement together with the associated safety sign on the ventilator equipment, the affirmative statement shall be placed in the IFU. | fail to address affirmative statement | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | Safety signs, including any supplementary symbol or text, shall be explained in the IFU. | sign loss of explanation | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The meaning of the symbols used for marking shall be explained in the IFU. | symbol loss of explanation | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The ventilator equipment shall be accompanied by documents containing at least the IFU and a technical description. | missing IFU in package | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The accompanying documents shall identify the equipment by including, as applicable, the following: Name or trade name of the manufacturer; Contact information to which the responsible organization can refer; Model or type reference. | fail to provide manufacturer information | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The accompanying documents may be provided electronically (e.g. electronic file format on CD-ROM), in which case the Usability Engineering File shall include specification as to which information needs to be provided as hard copy or as markings on the equipment. | fail to provide electronical document | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The accompanying documents shall specify any special skills, training and knowledge required of the intended operator or the responsible organization and any restrictions on locations or environments in which the ventilator equipment can be used. | inadequate document | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall document: The use of the ventilator equipment as intended by the manufacturer, The frequently used functions, Any known contraindication(s) to the use of the equipment, Those parts of the equipment that shall not be serviced or maintained while in-patient use. | inadequate IFU | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall include the name or trademark and address of the manufacturer. | lack of legal manufacturer name and address | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall include the model or type reference for the device. | inadequate ifu | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | For any single-use component of the ventilator, its parts or accessories, the IFU shall disclose the information on known characteristics and technical factors that could pose a risk if the ventilator, its parts or accessories would be reused. | fail to disclose single-use only | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall disclose the date of issue or the revisions of the IFU. | fail to disclose IFU revision | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall specify the use of the ventilator and accessory regarding the patient population. | fail to specify patient population | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall specify the use of the ventilator and accessory regarding the part of the body or type of tissue to which it is applied. | fail to specify applied tissue | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall include all warnings and safety notices. | lack of warning | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall include a warning statement to the effect: “WARNING: To avoid the risk of electric shock, this equipment must only be connected to a supply mains with protective earth.” | fail to disclose protective earth required | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall provide the operator or responsible organization with warnings regarding any significant risks of reciprocal interference posed by the presence of the ventilator during specific investigations or treatments. | inadequate ifu | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall include information regarding potential electromagnetic or other interference between the ventilator and other devices together with advice on ways to avoid or minimize such interference. | fail to advice EMI | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall include a warning statement to the effect that the ventilator shall not be covered or positioned in such a way that the operation or performance of the ventilator is adversely affected, including applicable examples. | fail to advice position | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall include a warning statement to the effect that, in case of ventilator failure, the lack of immediate access to appropriate alternative means of ventilation can result in patient death. | fail to advice backup ventilation | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall include a warning statement to the effect that adding attachments or other components or sub-assemblies to the ventilator breathing system can change the pressure gradient across the ventilator breathing system and that such changes to the ventilator breathing system can adversely affect the ventilator performance. | fail to disclose pressure changing | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall include a warning statement to the effect that nebulization or humidification can increase the resistance of breathing system filters and that the operator needs to monitor the breathing system filter frequently for increased resistance and blockage. | fail to disclose resistance increasion | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall include a warning statement to the effect that the ventilator shall not be used in a hyperbaric chamber. | fail to warn usage in hyperbaric chamber | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall include recommendation against the use of a gas-powered nebulizer and a warning statement to the effect that the ventilator accuracy can be affected by the gas added by use of a nebulizer. | fail to warn use of gas-powered nebulizer | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall include a warning statement in the HFO section to the effect that when the device is using the internal battery, HFO amplitude is limited to 60mbar | fail to warn use of internal battery | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall include a warning statement in the HFO section to the effect that during HFO ventilation the distinct recognition of disconnection and measuremente of the minute volume is limited | fail to disclose measurement accuracy | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall include a warning statement in the HFO section to the effect that the use of a heated inspiratory and expiratory tubing system is strongly recommended in HFO mode | fail to recommend heated tubing | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall include a warning statement in the FOT section to the effect that the “Manual Breath” can be activated during FOT, but this will deactivate FOT. To re-enable FOT after a Manual Breath /Sustained Inflation, you must restart FOT from the beginning. | lack of warning | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall include a warning statement in the HFO section to the effect that the Start / Stop button can be used to interrupt mechanical ventilation for two minutes, but this will deactivate FOT. | lack of warning | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall include a warning statement in the FOT section to the effect that any FOT measurements taken during Standby are NOT valid. | inadequate IFU | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall include a warning statement in the FOT section to the effect that inaccurate measurements may be caused by: Flow sensor not calibrated or defective; Incorrect pressure measurement line application or use; Excessive patient spontaneous breathing (strong active respiratory drive); | fail to disclose measurement accuracy | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall include the location of all natural-rubber-latex-based components. | fail to disclose toxic material | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall contain: a reference to where the installation instructions are to be found; or contact information for persons designated by the manufacturer as qualified to perform the installation. | fail to contain installation instruction | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall contain an instruction not to position the ventilator so that it is difficult to disconnect the power cord from the wall outlet. | improper device position | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall contain the necessary information for the operator to bring the ventilator into operation including such items as any initial control settings, connection to or positioning of the patient, etc. | fail to provide default setting | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall detail any treatment or handling needed before the ventilator, its parts, or accessories can be used. | insufficient treatment definition | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall disclose a method by which all of the alarm signals can be functionally tested to determine if they are operating correctly. | fail to disclose alarm test | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall include the following: a method of verifying all operator-adjustable alarm system functions; the time duration from start-up to providing essential performance. | fail to disclose warm up period | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall include a step by step description for the usage of FOT | lack of brief usability description | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall list all system messages, error messages and fault messages that are generated, unless these messages are self-explanatory. | fail to disclose error message | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall include an explanation of messages including important causes, and possible action(s) by the operator, if any, that are necessary to resolve the situation indicated by the message. | fail to disclose trouble shooting | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall contain the necessary information for the operator to safely terminate the operation of the ventilator. | fail to disclose terminate device | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall contain, for ventilator parts or accessories that can become contaminated through contact with the patient or with body fluids or expired gases during normal use and single fault condition, the following: details about cleaning and disinfection or sterilization methods that may be used, and list the applicable parameters such as temperature, pressure, humidity, time limits and number of cycles that such ventilator parts or accessories can tolerate. | fail to disclose contamination | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall identify which portions of the gas pathways through the ventilator can become contaminated with body fluids or expired gases during both normal condition and single fault condition. | fail to disclose contamination | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall instruct the operator or responsible organization in sufficient detail concerning preventive inspection, maintenance and calibration to be performed by them, including the frequency of such maintenance. | fail to disclose preventative maintenance | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall provide information for the safe performance of such routine maintenance necessary to ensure the continued safe use of the ventilator. | fail to disclose preventative maintenance | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall identify the parts on which preventative inspection and maintenance shall be performed by service personnel, including the periods to be applied, but not necessarily including details about the actual performance of such maintenance. | fail to disclose service personnel | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall include the procedures for calibration before or during use. | fail to disclose calibration | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall include the methods and frequency of routine inspection and testing. | fail to disclose maintenance frequency | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall include advice on the proper disposal of calibration gases. | inappropriate disposal of calibration gases | instruction for use issue | PRS
FAB-H-REQ-xxx | packaging and labelling | The IFU shall contain the information specified below: The permissible environmental conditions of use including conditions for transport and storage; All characteristics of the ventilator, including range(s), accuracy, and precision of the displayed values or an indication where they can be found; Any special installation requirements such as the maximum permissible apparent impedance of supply mains; A description of the means of isolating the ventilator from the supply mains, if such means is not incorporated in the ventilator; A warning statement that addresses the hazards that can result from unauthorized modification of the ventilator. | inadequate IFU | Inadequate instructions for healthcare professional | PRS
FAB-H-REQ-417 | Sterilization | The exhalation valve diaphragm shall be a single patient use. | single use violated | disinfection or sterilization issue at user location | PRS
FAB-H-REQ-418 | Sterilization | The exhalation valve diaphragm holder shall be capable of cleaning and disinfection per the labelled instructions for no less than 30 complete cleaning cycles, or for a maximum of one year (if sooner). | disinfection inefficacy | disinfection or sterilization issue at user location | PRS
FAB-H-REQ-419 | Sterilization | The instructions for use shall include instructions for emptying the exhalation valve and diaphragm assembly of water and secretions. | inable to reduce viable microorganism | disinfection or sterilization issue at user location | PRS
FAB-H-REQ-420 | Sterilization | The exhalation valve membrane holder shall be compatible with the following cleaning and disinfection process: Cleaning with pH neutral detergent (ValSure® Cleaner, Steris) at a temperature of 80 to 95 °C for 5 minutes; followed by Steam sterilization at 134 °C for 2 minutes (half cycle). | improper disinfection | disinfection or sterilization issue at user location | PRS
FAB-H-REQ-421 | Sterilization | The ventilator enclosure shall be designed for a minimum of 400 cleaning cycles. | inappropriate cleaning cycle | device cleaning issue | PRS
FAB-H-REQ-422 | Sterilization | The ventilator enclosure shall be compatible with cleaning with each of the following materials, when used in accordance with manufacturers instructions: CaviWipes™ (Metrex); 10% solution of 5.25% sodium hypochlorite bleach (e.g. Clorox), used with lint-free textile cleaning cloths; Hydrogen peroxide 5% (e.g. 35% food grade H2O2, diluted 1:6 with distilled water), used with lint-free textile cleaning cloths. | incompatible solvent | device cleaning issue | PRS
FAB-H-REQ-423 | Sterilization | The flow sensor cable shall be designed for a minimum of 400 cleaning cycles. | inappropriate cleaning cycle | device cleaning issue | PRS
FAB-H-REQ-424 | Sterilization | The flow sensor cable shall be compatible with cleaning with each of the following materials, when used in accordance with manufacturers instructions: CaviWipes™ (Metrex); 10% solution of 5.25% sodium hypochlorite bleach (e.g. Clorox), used with lint-free textile cleaning cloths; Hydrogen peroxide 5% (e.g. 35% food grade H2O2, diluted 1:6 with distilled water), used with lint-free textile cleaning cloths. | inappropriate cleanser | device cleaning issue | PRS
FAB-H-REQ-xxx | Regulatory | The ventilator shall be CE marked to demonstrate conformity to the GSPR of the EU Medical Devices Directive | fail to comply with CE Mark | compliance issue | PRS
FAB-H-REQ-183 | Regulatory | The fabian ventilator shall comply with ISO 80601-2-12:2020 ("Medical electrical equipment - Part 2-12: Particular requirements for basic safety and essential performance of critical care ventilators"). | fail to comply with ventilator safety requirement | compliance issue | PRS
FAB-H-REQ-184 | Regulatory | The fabian ventilator shall comply with EN IEC60601-1:2005 + A1:2012 + A2:2020 Medical electrical equipment – Part 1: General requirements for basic safety and essential performance | fail to comply with electrical safety requirement | compliance issue | PRS
FAB-H-REQ-185 | Regulatory | The fabian ventilator shall comply with EN IEC 60601-1-2:2014 + A1:2020 Medical electrical equipment –Part 1-2: General requirements for basic safety and essential performance – Collateral Standard: Electromagnetic disturbances – Requirements and tests. | fail to comply with EM safety requirement | compliance issue | PRS
FAB-H-REQ-186 | Regulatory | The fabian ventilator shall comply with EN IEC 60601-1-6:2010+A1:2013 + A2:2020 Medical electrical equipment –Part 1-6: General requirements for basic safety and essential performance – Collateral Standard: Usability. | fail to comply with usability safety requirement | compliance issue | PRS
FAB-H-REQ-187 | Regulatory | The fabian ventilator shall comply with EN IEC 60601-1-8:2006+A1:2012+A2:2020 Medical electrical equipment –Part 1-8: General requirements for basic safety and essential performance – Collateral Standard: General requirements, tests and guidance for alarm systems in medical electrical equipment and medical electrical systems. | fail to comply with alarm system requirement | compliance issue | PRS
FAB-H-REQ-188 | Regulatory | The fabian ventilator oxygen monitoring and etCO2 monitoring functions shall comply with EN ISO 80601-2-55:2018 Medical electrical equipment -- Part 2-55: Particular requirements for the basic safety and essential performance of respiratory gas monitors. | fail to comply requirement for respiratory gas monitor | compliance issue | PRS
FAB-H-REQ-189 | Regulatory | The fabian ventilator pulse oximetry functions shall comply with EN ISO 80601-2-61:2017 Medical electrical equipment -- Part 2-61: Particular requirements for basic safety and essential performance of pulse oximeter equipment. | -iso regulation violated -incorrect triggered pulse -inadequate safety information | compliance issue | PRS
FAB-H-REQ-190 | Regulatory | The fabian ventilator shall comply with IEC 62304:2006+A1:2016 Medical device software – Software life cycle processes. | fail to comply with software requirement | compliance issue | PRS
FAB-H-REQ-191 | Regulatory | The fabian ventilator shall comply with IEC 62366-1:2016 Medical devices – Application of usability engineering to medical devices. | fail to comply with usability requirement | compliance issue | PRS
FAB-H-REQ-xxx | Regulatory | The fabian ventilator when operating in HFO mode shall comply with ISO FDIS 80601-2-87:2021 Medical electrical equipment - Part 2-87: Particular requirements for basic safety and essential performance of high frequency ventilators. | fail to comply requirement for high frequency ventilator | compliance issue | PRS
FAB-H-REQ-17 | configuration | The ventilator shall provide means to enter license codes to enable and / or disable the following options: FOT, Volume Guarantee, Volume Limit, Trend Display, NCPAP / Biphasic, O2-Therapy, HFO, PRICO | invalid license code | human-device interface issue | PRS
FAB-H-REQ-144 | Installation and servicing | The ventilator shall provide means for the user to calibrate the proximal flow sensor (drift correction). | miscalibration | failure to calibrate | PRS
FAB-H-REQ-145 | Installation and servicing | The ventilator shall provide a prompt to the operator to calibrate the proximal flow sensor at startup (if it is connected) and on each occasion that a proximal flow sensor is connected to the ventilator. | difficult to calibrate | failure to calibrate | PRS
FAB-H-REQ-146 | Installation and servicing | The ventilator shall provide a notification on each occasion that a proximal flow sensor is disconnected from the ventilator. | fail to trigger calibration | failure to calibrate | PRS
FAB-H-REQ-147 | Installation and servicing | The ventilator shall provide automated means for calibration of the oxygen sensor, at start-up and at intervals not exceeding 24 hours, without interruption to the patient inspired oxygen concentration. | fail to trigger calibration | failure to calibrate | PRS
FAB-H-REQ-530 | Installation and servicing | The oxygen sensor calibration process in use shall not cause the oxygen calibration to become corrupt due to a slow response from the oxygen cell. | calibration failed | failure to calibrate | PRS
FAB-H-REQ-148 | Installation and servicing | When a capnometry device is in use, the ventilator shall provide means for the user to perform zero calibration for the capnometry sensor. | miscalibration | failure to calibrate | PRS
FAB-H-REQ-155 | Installation and servicing | Planned maintenance shall be specified in a Service Manual, and / or service procedure documents. | fail to disclose preventative maintenance | lack of maintenance documentation or guidelines | PRS
FAB-H-REQ-156 | Installation and servicing | Planned maintenance documentation shall include a procedure to calibrate the blender flow sensors. | fail to disclose calibration | instruction for use issue | PRS
FAB-H-REQ-157 | Installation and servicing | Planned maintenance documentation shall include a procedure to characterize the exhalation valve. | fail to disclose characterization of exhalation valve | lack of maintenance documentation or guidelines | PRS
FAB-H-REQ-158 | Installation and servicing | Planned maintenance documentation shall include a procedure to calibrate the inspiratory and proximal pressure sensors. | fail to disclose sensor calibration | lack of maintenance documentation or guidelines | PRS
FAB-H-REQ-524 | Installation and servicing | Planned maintenance documentation shall include replacement interval and instructions for the RTC battery. | fail to disclose battery replacement | lack of maintenance documentation or guidelines | PRS
FAB-H-REQ-159 | Installation and servicing | All replacement parts shall be provided with instructions for installation, and which confirm what software version, software re-installation, calibration and processes are required following installation. | fail to disclose installation instruction | installation-related problem | PRS
FAB-H-REQ-160 | Installation and servicing | The ventilator shall provide a display of the battery charge / discharge cycles, and provide means for the user to measure remaining battery capacity. | fail to identify remaining battery | inadequate service | PRS
FAB-H-REQ-161 | Installation and servicing | The ventilator shall provide a counter for operating hours. | fail to count operating hours | failure to service | PRS
FAB-H-REQ-162 | Installation and servicing | The ventilator shall incorporate means to set, and display, the next due service date. | fail to set next service date | failure to service | PRS
FAB-H-REQ-163 | Installation and servicing | Planned maintenance activities shall not be required at more frequent than 12-month intervals. | inappropriate interval settings | maintenance does not comply to manufacturers recommendations | PRS
FAB-H-REQ-164 | Installation and servicing | The ventilator shall provide means to temporarily disable alarms, only accessible through a password-controlled service screen. | unauthorized access | unauthorized access to computer system | PRS
FAB-H-REQ-165 | Installation and servicing | The alarm system shall be re-enabled on power-cycle. | fail to reset alarm automatically | device alarm system issue | PRS
FAB-H-REQ-166 | Installation and servicing | A software installation procedure shall be documented that includes verification of the complete installation hardware and software configuration and that generates an installation report to be provided to the manufacturer. | fail to disclose verification of software installation | problem with software installation | PRS
FAB-MHFO-001 | Interface Connectors Requirements | The HFO pneumatic subsystem shall have HFO outlet port connected to Inspiratory tube port at manifold through HFO flexible hose, specification of HFO flexible hose mentioned below. HFO outlet port shall connect patient-connection port close to the patient for connection to an airway device. | connection detached | disconnection | Hardware_pneumatic
FAB-MHFO-xxx | Interface Connectors Requirements | HFO flexible hose: Adult breathing hose with sleeves for connectors with ISO 22 mm on both ends. Length of breathing hose shall be determined based on placement of connecting ports. Biocompatibility requirement: ISO 18562-1: 2017, ISO 15001:2010. | improper length | defective component | Hardware_pneumatic
FAB-MHFO-002 | Interface Connectors Requirements | Specification of HFO outlet port mentioned below. HFO outlet port: Barb connector compatible with 22 mm hose HFO flexible hose. HFO outlet port shall connect patient-connection port close to the patient for connection to an airway device. Material: Stainless steel | size out of specification | defective component | Hardware_pneumatic
FAB-MHFO-003 | Operational requirement | The HFO pneumatic subsystem shall include a mechanism to shut off/on pneumatic output at HFO outlet port that consists of HFO valve placed between HFO generator and HFO output port. | lack of shut off mechanism | protective measure issue | Hardware_pneumatic
FAB-MHFO-004 | Operational requirement | The HFO pneumatic subsystem shall provide interface between HFO PCBA and HFO Valve. HFO Valve will be controlled as CLOSED (not actuated) or OPEN (actuated) by HFO PCBA and shall have the following control specifications. Interface type: PWM drive through GPIO Notes: For details please refer HFO PCBA subsystem requirement documents. | false output | False Device Output | Hardware_pneumatic
FAB-MHFO-005 | Operational requirement | HFO valve shall be electrically operated by 2/2 way actuation valve, normally closed (shutoff). HFO valve shall have below specifications. (Requires detailed data sheet to confirm below specifications) • Insulation class of the winding: Class F • Reference Temperature: 0 to 50 deg C • Operating mode: S1 (100 %) • Rated Voltage: 12 V(DC), • Rated current: 0.35 amp • Holding current: 0.08 amp • Peak current: 0.39 amp • Power: 4,12 Watt Notes: For more details please refer HFO PCBA subsystem requirement documents. | fail to control valve | circuit failure | Hardware_pneumatic
FAB-MHFO-006 | Operational requirement | The HFO pneumatic subsystem shall have interface between HFO PCBA and High frequency oscillation generator by electrical input parameters through Class D amplifier for generating the High Frequency (HF) oscillations. High Frequency (HF) oscillations controls different output parameters as mentioned below. • Frequency 5 to 20 Hz with resolution 1 Hz with accuracy of ± (0.5 Hz). • 1:1 to 1:3 I:E ratio with resolution 1 (i.e. range is 1:1, 1:2, 1:3) | frequency out of range | electronic property issue | Hardware_pneumatic
FAB-MHFO-007 | Operational requirement | The HFO module includes an electromagnetic solenoid and horns attached with diaphragm which converts the electrical signal come from the PCB, converting it to a pneumatic output mentioned below. • 0 to 120 mbar peak to peak pressure with resolution <= 1 mbar and accuracy ± (0.5 mbar+ 4%). • Displacement volume per oscillation 0 to TBD millilitres. Notes: For details of controlling parameters please refer HFO PCBA subsystem requirement documents. | solenoid failed to energize | output issue | Hardware_pneumatic
FAB-MHFO-008 | Operational requirement | HFO pneumatic subsystem with HFO PCBA shall provide High frequency oscillation in presence of mean airway pressure in range 0 to 50 mbar. | fail to deliver high frequency pressure | pressure issue | Hardware_pneumatic
FAB-MHFO-009 | Operational requirement | The mechanical horn coils shall vibrate the Membrane which in turn shall generate the pneumatic oscillations. Specifications of Mechanical Horn coils mentioned below: • Rated Voltage: TBD • Rated current: TBD • Holding current: TBD • Peak current: TBD • unidirectional from holding current Notes: For details of controlling parameters please refer HFO PCBA subsystem requirement documents. | fail to generate pneumatic oscillation | mechanical issue | Hardware_pneumatic
FAB-MHFO-010 | Operational requirement | The HFO valve shall be operated by HFO PIC controller. | fail to control valve | circuit failure | Hardware_pneumatic
FAB-MHFO-011 | Operational requirement | The Mechanical Horn coils shall be operated by HFO PIC controller. | fail to activate control | device inoperable | Hardware_pneumatic
FAB-MHFO-012 | Alarm functions | HFO pneumatic subsystem shall trigger alarm to main board PCBA for HFO related fault happening with HFO oscillation. • HFO module shutoff valve electrical fault HFO module oscillator electrical fault | fail to activate alarm | device alarm system issue | Hardware_pneumatic
FAB-MHFO-013 | Construction requirement (Subsystem/Process) | The HFO generator diaphragm shall sustain minimum and maximum pressure up to -60 to +140 mbar at HFO Oscillation generator. | fail to hold pressure | pressure issue | Hardware_pneumatic
FAB-MHFO-014 | Construction requirement (Subsystem/Process) | HFO generator Body shall have exhaust hole to assist pneumatic oscillation and free movement of diaphragm. | hole blocked | Blockage within Device or Device Component | Hardware_pneumatic
FAB-MHFO-015 | Construction requirement (Subsystem/Process) | All pneumatic connections shall incorporate means for mechanical retention. | loose connection | detachment of device component | Hardware_pneumatic
FAB-MHFO-016 | Construction requirement (Subsystem/Process) | The HFO PCBA & HFO pneumatic subsystem shall have proper mounting to avoid self-dismantling or damages during the function / transportation. | self-dismantling | detachment of device component | Hardware_pneumatic
FAB-MHFO-017 | Construction requirement (Subsystem/Process) | The HFO PCBA & HFO pneumatic subsystem when operating in HFO mode shall comply with ISO 80601-2-87:2021 Medical electrical equipment - Part 2-87: Particular requirements for basic safety and essential performance of high frequency ventilators. Current protection: TBD | fail to comply requirement for high frequency ventilator | compliance issue | Hardware_pneumatic
FAB-MHFO-018 | Material requirement | The material in contact with medical gas shall be compatible with oxygen and AIR. Note 1: Corrosion resistance includes resistance to moisture and surrounding material. Note 2: Compatibility with oxygen involves both combustibility and ease of ignition. Material which bun in AIR will burn violently in pure oxygen. | moisture induced corrosion | Corrosion | Hardware_pneumatic
FAB-MHFO-019 | Material requirement | All pneumatic connections shall consist of following Gas contact components and other components that can be planned with aligned material as mentioned below. Gas Contact Materials: Aluminum (6082 alloy or 6061 alloy), Nitrile rubber, Silicon Rubber, Stainless steel Non-Gas Contact Materials: Aluminum (6082 alloy or 6061 alloy), Low Carbon Steel, Spring Steel, Rare-earth magnet, Off shelf item | fail to compatible with Oxygen | material erosion | Hardware_pneumatic
FAB-MHFO-XXX | Material requirement | Regulatory compliance applicable are REACH requirements and RoHS requirements for all intended markets. | REACH violated | environmental control or utility issue | Hardware_pneumatic
FAB-MHFO-020 | Physical requirement | The HFO pneumatic subsystem shall have the maximum physical dimensions placed inside as per below specifications: 1. Length – 190mm 2. Width – 170mm 3. Thickness – 170mm | out of dimension | defective component | Hardware_pneumatic
FAB-MHFO-021 | Physical requirement | The HFO pneumatic subsystem shall be mounted on plate with 2 Nos of M5 screws and same plate shall be mounted with mounting holes of TBD mm aligned to mounting threaded inserts at enclosures. | low mount strength | mechanical issue | Hardware_pneumatic
FAB-MHFO-022 | Physical requirement | The HFO PCBA shall have 3 NPTH Mounting holes of 3.2±0.1mm for mounting the card to the studs in enclosure. The mounting holes shall have routing clearance of min. 1mm around their circumference. | out of dimension | defective component | Hardware_pneumatic
FAB-MHFO-023 | Physical requirement | The HFO PCBA & HFO pneumatic subsystem shall be weighted not more than TBDkg. | weight out of range | patient-device incompatibility | Hardware_pneumatic
FAB-MHFO-024 | Operating environmental | The HFO PCBA & HFO pneumatic subsystem shall withstand and perform as intended within the following environmental conditions. • Temperature +0 to +50 °C • Relative Humidity 10 to 90%, non-condensing • Ambient pressure 70 to 106 kPa (10.2 to 15.0 PSIA) Note: the ambient pressure range is intended to equate to approximately -1350 to +10,000 feet altitude. Note: Temperature range of HFO pneumatic subsystem should withstand and perform at higher range of temperature +/-10°C w.r.t device since its generating High Frequency pneumatic oscillation in terms of flow and pressure. | inappropriate environment condition | environmental control or utility issue | Hardware_pneumatic
FAB-MHFO-025 | Operating environmental | The HFO PCBA & HFO pneumatic subsystem shall withstand and conform to specifications following shipping & distribution (transport) for up to 72h within the following environmental conditions. • Temperature -20 to +60 °C • Relative Humidity 10 to 95%, non-condensing • Ambient pressure 50 to 106 kPa (7.2 to 15.0 PSIA) Note: the ambient pressure range is intended to equate to approximately -1350 to +10,000 feet altitude. | inappropriate shipping environment or condition | Shipping Damage or Problem | Hardware_pneumatic
FAB-MHFO-026 | Other requirement | HFO pneumatic subsystem shall qualify leak test with pressure of 250 mbar. Leak rate should not exceed 10 cc/min when HFO valve open and HFO generator isolated between flexible hose outlet and subsystem. | fail to pass leak test | failure of device to self-test | Hardware_pneumatic
FAB-MGAS-001 | Pneumatic Interface Requirements | Gas Input subsystem shall provide standard connector to connect separate pneumatic interfaces for 02 as gas input at rear panel for following purpose. 1. Connection for compressed oxygen “O2” supply 2.0 – 6.0 bar (200 - 600 kPa) compatible with flow rate not greater than 40 L/min. Notes: Gas input subsystem shall be capable to supplying sufficient range of flow and pressure to deliver various modes of ventilation/therapy requirement. | non-standardized connector | nonstandard device or device component | Hardware_gas
FAB-MGAS-002 | Pneumatic Interface Requirements | Gas Input subsystem shall provide standard connector to connect separate pneumatic interfaces for AIR as gas input at rear panel for following purpose. 1. Connection for compressed AIR supply 2.0 – 6.0 bar (200 - 600 kPa) compatible with flow rate not greater than 40 L/min. Notes: Gas input subsystem shall be capable to supplying sufficient range of flow and pressure to deliver various modes of ventilation/therapy requirement. | fail to supply range of flow | insufficient flow or underinfusion | Hardware_gas
FAB-MGAS-003 | Pneumatic Interface Requirements | O2 inlet connector at rear panel shall have NIST connector (NIST body adaptor) as per ISO 18082:2014. The specifications of the standard connector as below: Inlet Port: Gas specific NIST Body adaptor with threaded connector. Dimensions of NIST body adaptor is per image and table mentioned below: | non-standardized connector | nonstandard device or device component | Hardware_gas
FAB-MGAS-004 | Pneumatic Interface Requirements | AIR inlet connector at rear panel shall have NIST connector (NIST body adaptor) as per ISO 18082:2014. The specifications of the standard connector as below: Inlet Port: Gas specific NIST Body adaptor with threaded connector. Dimensions of NIST body adaptor is per image and table mentioned below: | non-standardized connector | nonstandard device or device component | Hardware_gas
FAB-MGAS-005 | Pneumatic Interface Requirements | O2 inlet connector at rear panel shall be compatible with Gas(O2) specific NIST low pressure connectors (NIST nipple) for medical gasses (attached at gas hose along with NIST nut) only. Connector size: Gas specific NIST Nipple shall have dimensions are per image and table mentioned below: | non-standardized connector | nonstandard device or device component | Hardware_gas
FAB-MGAS-006 | Pneumatic Interface Requirements | AIR inlet connector at rear panel shall be compatible with Gas (AIR) specific NIST low pressure connectors (NIST nipple) for medical gasses (attached at gas hose) only. Connector size: Gas specific NIST Nipple shall have dimensions are per image and table mentioned below: | non-standardized connector | nonstandard device or device component | Hardware_gas
FAB-MGAS-007 | Pneumatic Interface Requirements | Gas input subsystem shall be compatible with detachable water trap (filter) at AIR input gas supply line, It provides NIST body adaptor (male) as input connector tilted 90 downwards and NIST nipple and nut (female) as outlet connector coaxial with AIR Gas input subsystem body adaptor without gauge mounting. | fail to compatible with filter | component or accessory incompatibility | Hardware_gas
FAB-MGAS-008 | Pneumatic Interface Requirements | Water trap have following specification. Pressure bearing capacity shall be 10 bar. Note: The inlet filters provide protection for the valves and blender flow sensor from contamination. | fail to withstand pressure | mechanical issue | Hardware_gas
FAB-MGAS-009 | Pneumatic Interface Requirements | Gas input subsystem of O2 supply line) shall have separate output port as standard push pneumatic connector. Maximum Inlet Pressure rating: 10 Bar Type of Connector: Push fit tubing connector Standard polyurethane Tubing: 6 mm OD Connector size: Straight - Straight connector Threaded/6 mm tubing Tubing Color: Gray | fail to separate output port | protective measure issue | Hardware_gas
FAB-MGAS-010 | Pneumatic Interface Requirements | Gas input subsystem of AIR supply line shall have separate output port as standard push pneumatic connector. Maximum Inlet Pressure rating: 10 Bar Type of Connector: Push fit tubing connector Standard polyurethane tubing: 6 mm OD Connector size: Straight - Straight connector Threaded/6 mm tubing Tubing Color: Black | non-standardized connector | nonstandard device or device component | Hardware_gas
FAB-MGAS-011 | Operational requirement | Gas input subsystem’s each component and joint (sealed with O ring / sealant/ Loctite) shall sustain 10 bar maximum inlet pressure. | fail to withstand pressure | mechanical issue | Hardware_gas
FAB-MGAS-012 | Operational requirement | Gas input subsystem shall comply following functional leak requirements. Leak test: Leakage test for subsystem shall not exceed 50ml/min at following test parameters. a. Compressed gas pressure 6.0 bar with blocked output’s b. At (23+/-2 deg C) Test Gas: Carry out test specific gas for which gas inlet is designed with dry gas with maximum moisture content of 50 microgram/gram, corresponding to dew point of -48 deg C. | fail to pass leak test | failure of device to self-test | Hardware_gas
FAB-MGAS-013 | Operational requirement | Gas input subsystem shall have inbuilt sintered filters having a pore size of less than or equal to 100 microns to particulate inlet filters to provide protection for the valves and blender flow sensor from contamination. Specification of sintered filer is mentioned below. Shape: Conical 100 microns perforation Maximum flow rate: 40 L/min TBD Note: Sintered filter to be provided with a particulate filter at each gas (O2 and AIR) subsystem. | pore size out of range | Filtration Issue | Hardware_gas
FAB-MGAS-014 | Operational requirement | Gas input subsystem (with inbuilt filter) shall comply following functional flow restriction requirements. The pneumatic output shall be 40 l/min least with at 2 bar pressure applied at gas input subsystem, pressure drop of input to output to be less than 0.2 bar. Notes: Carry out test at (23+/-2 deg C). | flow out of range | improper flow or infusion | Hardware_gas
FAB-MGAS-015 | Construction Requirements | The ventilator shall provide the gas O2/AIR input construct/connector in such way so that they are non-interchangeable to prevent wrong pneumatic connections. Note: Purpose of this requirement is to verify Probe of O2 and Air can be connected to their specific NIST connector only in order to prevent wrong connection. This design feature is implemented by NIST nipple and body adapter in accordance with ISO: 18082. | wrong connector | misconnection | Hardware_gas
FAB-MGAS-016 | Construction Requirements | AIR/O2 Gas input subsystem shall have locking mechanism (anti rotation) at rear panel internally mounted to avoid self-dismantling or damages due to vibration during the function / transportation. | lack of locking mechanism | Detachment of Device or Device Component | Hardware_gas
FAB-MGAS-017 | Construction Requirements | All pneumatic connections at Gas input subsystem components shall be constructed by following material. • NIST Connector: – Stainless steel • Lock Nut: Stainless steel • Lock pin: Stainless steel • O-Ring - Nitrile rubber (NBR) • Any other seal -Silicon Rubber • Sealant/Loctite: TBD | incorrect chemical composition | material erosion | Hardware_gas
FAB-MGAS-018 | Labeling Requirement | Each of the NIST connector shall be laser marked or labelled with the gas name or chemical symbol (“AIR” and “O2” for oxygen) for proper identification while assembly and storage, since they are identical in design. | fail to label text | device markings issue | Hardware_gas
FAB-MGAS-019 | Biocompatibility Requirements | The material in contact with medical gas shall be compatible with oxygen and AIR. . Note 1: Corrosion resistance includes resistance to moisture and surrounding material. Note 2: Compatibility with oxygen involves both combustibility and ease of ignition. Material which bun in AIR will burn violently in pure oxygen. | fail to compatible with Oxygen | material erosion | Hardware_gas
FAB-MGAS-020 | Physical Requirements | Gas input subsystem shall be place inside the ventilator enclosure and mounted on rear panel with proper mechanical retention. | self-dismantling | detachment of device component | Hardware_gas
FAB-MGAS-021 | Physical Requirements | The Gas input subsystem shall have the maximum physical dimensions as per below specifications: Approximate 50% of length portion shall be projected outside of enclosure(rear panel). 1. Length – 100mm 2. Width – 40 mm 3. Thickness – 40mm | out of dimension | defective component | Hardware_gas
FAB-MGAS-022 | Environmental requirements | The gas input subsystem shall withstand and perform as intended within the following environmental conditions. • Temperature +10 to +40 °C • Relative Humidity 10 to 90%, non-condensing • Ambient pressure 70 to 106 kPa (10.2 to 15.0 PSIA) Note: The ambient pressure range is intended to equate to approximately -1350 to +10,000 feet altitude. | inappropriate environment condition | environmental control or utility issue | Hardware_gas
FAB-MGAS-023 | Environmental requirements | The gas input subsystem shall withstand and conform to specifications following shipping & distribution (transport) for up to 72h within the following environmental conditions. • Temperature -20 to +60 °C • Relative Humidity 10 to 95%, non-condensing • Ambient pressure 50 to 106 kPa (7.2 to 15.0 PSIA) | inappropriate shipping environment or condition | Shipping Damage or Problem | Hardware_gas
FAB-SAD-001 | system function | The Fabian HFO is intended for premature infants, new-born as well as children weighing up to 30 kg, “in-patient use” in professional healthcare environment including intra-hospital patient transport. | inappropriate user | patient-device incompatibility | SAD
FAB-SAD-002 | system function | The Fabian HFO Ventilator is an innovative, compact, microprocessor-based device for accurate control and monitoring of respiration. The pneumatic circuits are based on a novel control concept using two fast micro controllers for monitoring and control. The user interface is controlled by a single board computer based on a RISC-Processor that also supervises the pneumatic controller circuits. For monitoring flow the Fabian HFO uses a small, low dead space plastic sensor that is adapted to the ET tube close to the patient. | processor overheat | overheating of device or device component | SAD
FAB-SAD-003 | system function | Oxygen is metered by the integrated Air/O2 blender. The oxygen concentration is measured internally with a galvanic oxygen sensor. | low measurement accuracy | incorrect measurement | SAD
FAB-SAD-004 | system overview | This section of the System Architecture Definition document includes design output as well as design specifications. Tags in this section do not denote requirements that are intended to be tested at the system level. | ambiguous specification | physical property issue | SAD
FAB-SAD-005 | Theory of Operation | The blender module consists of 1 proportional valve for air and 1 proportional valve for oxygen. Each valve can deliver a flow from 0 to the maximum. The highest flow through one valve is 32 LPM. The delivered flow of each valve is measured by a flow sensor and a local control loop maintains the correct value of flow. | low proof pressure | pressure issue | SAD
FAB-SAD-006 | Theory of Operation | The air input pressure and oxygen input pressure are read by pressure transducers on the blender module board. | transducer overheating | overheating of device or device component | SAD
FAB-SAD-007 | Theory of Operation | The respiratory gas delivered to the patient is a mixture of medical air and O2. Based on the clinical need, the O2 concentration in the blend is programmable from 21% to 100%. The relationship between O2 and air concentrations are described in the equations below: | inappropriate mixture of medical air and O2 | medical gas supply problem | SAD
FAB-SAD-008 | Theory of Operation | The mechanism with which the Ventilator components achieve the target blend ratio is described in section 7.2 Pneumatics Architecture. | incorrect ratio calculation | parameter calculation error due to software problem | SAD
FAB-SAD-009 | Theory of Operation | The total outlet flow target is servo controlled to achieve a desired pattern of ventilation, which may either be based on a target pressure waveform, a target flow waveform, or dual control (breaths which switch between control target modalities). | device malfunction | device operates differently than expected | SAD
FAB-SAD-010 | Theory of Operation | The calibration of the oxygen sensor is performed at two points. The sensor is calibrated at 21% (room air) and at 100% oxygen. It is also possible to activate the automatic calibration of only 21% or 21% and 100% oxygen. The automatic O2 Sensor calibration will be accomplished at start-up and every 24 hours. | fail to activate calibration at multiple points | Failure to Calibrate | SAD
FAB-SAD-011 | Theory of Operation | The Fabian HFO Ventilator is an innovative, compact, microprocessor-based device for accurate control and monitoring of respiration. The pneumatic circuits are based on a novel control concept using two fast micro controllers for monitoring and control. | low control accuracy | imprecision | SAD
FAB-SAD-012 | Theory of Operation | The user interface is controlled by a single board computer based on a RISC-Processor that also supervises the pneumatic controller circuits. For monitoring flow the Fabian HFO uses a small, low dead space plastic sensor that is adapted to the ET tube close to the patient. | processor overheat | overheating of device or device component | SAD
FAB-SAD-013 | Theory of Operation | The output pressure of the blender delivered into the tube system is monitored. There are two security functions realized by this sensor. 1. Inspiratory pressure sensor – blender status 2. Inspiratory pressure sensor – airway pressure compare | loss of sensor data | loss of data | SAD
FAB-SAD-014 | Theory of Operation | The airway pressure sensor monitors the pressure close to the patient at the y-connector. Its value is used to regulate pressure conditions through the exhalation valve by a closed loop regulation circuit. | fail to regulate pressure condition | pressure issue | SAD
FAB-SAD-015 | Theory of Operation | The total outlet flow target (i.e. Vout,target) is servo-controlled to achieve a desired pattern of ventilation, which may be based on the desired pressure waveform. The Ventilator is programmed to achieve a target pressure waveform by controlling the opening pressure of the Exhalation Valve. | incorrect ventilation pattern | improper device output | SAD
FAB-SAD-016 | Theory of Operation | Pressure breaths have a controlled pressure waveform, which may be comprised of components that are preset by the operator (e.g. pressure rise time and plateau pressure) and empirically-derived values (e.g. leakage of air at the patient mask or nasal prongs in NIV mode). | inappropriate preset value | pressure issue | SAD
FAB-SAD-017 | Theory of Operation | In the Fabian HFO Ventilator, pressure of delivered breaths are controlled by the Exhalation Valve. | fail to control delivered pressure | pressure issue | SAD
FAB-SAD-018 | Ventilator | The Fabian HFO ventilator will deliver gas to the patient. The Fabian HFO ventilates with excess pressure based on the continuous-flow principle. (Time cycled, pressure/volume limited or volume guaranteed). Oxygen is metered by the integrated Air/O2 blender. The oxygen concentration is measured internally with a galvanic oxygen sensor. | fail to compensate pressure | pressure issue | SAD
FAB-SAD-019 | Ventilator | The ventilator incorporates connections to high pressure gas inlet; the patient circuit (inspiratory and expiratory ports); a shared connector port; and dedicated ports for the capnometry, (optional) pulse oximetry, and nebulizer accessory devices. | component detached | detachment of device component | SAD
FAB-SAD-020 | Ventilator | The shared connector port provides isolated DC power and signal interface for: • Hot wire flow sensor • TPP module (Esophageal pressure measurement system) | signal interference | radiofrequency interference (RFI) | SAD
FAB-SAD-021 | Ventilator | The capnometer port provides isolated DC power and signal interface for a mainstream or sides-tream capnometer module. | electrical leakage | electrical issue | SAD
FAB-SAD-022 | Ventilator | The nebulizer port provides isolated drive signals for an electronic nebulizer. | signal interference | radiofrequency interference (RFI) | SAD
FAB-SAD-023 | Ventilator | The ventilator incorporates an external alarm connector for a remote nurse call. | incompatible accessory | accessory incompatible | SAD
FAB-SAD-024 | Ventilator | The ventilator incorporates RS232 port for RS232 port for PDMS to ICU monitors. Connectivity via this port is defined in the Connectivity section of the PRS and includes VOXP protocol. | incompatible accessory | accessory incompatible | SAD
FAB-SAD-025 | Ventilator | The ventilator incorporates a USB port for data output, Software update and SpO2 Module (MASIMO) | fail to incorporate | material disintegration | SAD
FAB-SAD-026 | Ventilator | The ventilator incorporates a non-isolated HDMI output for connection to an external display monitor or a projector. This functionality is labeled as not for clinical use. | incompatible accessory | accessory incompatible | SAD
FAB-SAD-027 | Ventilator | The ventilator incorporates an internal rechargeable battery pack. | incompatible accessory | material disintegration | SAD
FAB-SAD-028 | Ventilator | The ventilator incorporates a variable speed fan to provide ambient air flow to inside the ventilator for cooling ventilator components. There will be a gross particulate filter in front of the fan. This filter can be accessed and cleaned or replaced by the end-user without the need to remove any ventilator covers. Note: the primary protection against oxygen buildup within the device enclosure will be dual seals on all oxygen-carrying parts. | pore size out of range | Filtration Issue | SAD
FAB-SAD-029 | Ventilator | The ventilator mobile stand (cart) provides a secure, wheeled mounting for the ventilator and its accessories. The mobile stand incorporates a mechanical attachment point for the ventilator, a circuit support arm, a mechanical attachment point for the optional humidifier, a basket or tray for storing at least one patient circuit and patient interface set and a holder for two ‘E” size gas cylinders. | unsecured stand | migration of device or device component | SAD
FAB-SAD-030 | Ventilator | The ventilator will have an AC power connection. The ventilator will support the use of an IEC 60320 terminated power cordset for connection to wall AC power. | inappropriate power cordset | component incompatible | SAD
FAB-SAD-031 | Ventilator | The ventilator will include a graphical user interface to allow access to all normal operating controls and functions. The graphical user interface will include a color TFT-LCD display assembly, touch screen panel, ambient light sensor. LCD screen provides real-time graphic visualization of pressure, flow, and volume on the LCD screen | fail to access functions | computer software issue | SAD
FAB-SAD-032 | Ventilator | The LCD panel with touchscreen will have a size 26 cm (10.4”) will have a resolution of 800x600 pixels connect the LCD panel and backlight cables to the Mainboard. | low resolution | Image Resolution Poor | SAD
FAB-SAD-033 | Ventilator | The LCD display incorporates a display backlight. The backlight brightness is programmable via the GUI. | fail to adjust brightness | optical issue | SAD
FAB-SAD-034 | Ventilator | The ventilator incorporates two high-pressure gas inlets, for connecting to a high-pressure oxygen/air source and a high-pressure medical air both at nominal 2.0 – 6.5 bar. The connector adaptor plate incorporates magnets to provide a code that identifies the gas type. The magnetic code is read through Hall-effect switches. | fail to read code | failure to read input signal | SAD
FAB-SAD-035 | Ventilator | In Fabian HFO ventilator, the high-pressure gas inlets will each use a gas-specific connector in one of the following types: male DISS, female DISS, SIS, NIST to prevent connection to an incorrect gas pipeline. | incompatible connection | component or accessory incompatibility | SAD
FAB-SAD-036 | Ventilator | In Fabian HFO ventilator, the high-pressure gas inlets can identify, at least, the following gas types: • Oxygen • Air (note: Air is normal use, but identification is required to prevent foreseeable misuse) | fail to identify gas type | failure to sense | SAD
FAB-SAD-037 | Ventilator | For use with high-pressure air, oxygen, a gas specific hose and inlet manifold will be required. This gas hose will be color coded with gas specific connectors at the supply end, to match the encoded ID. | missing color code | device markings issue | SAD
FAB-SAD-038 | Ventilator | This gas hose will be available in various configurations applicable to country of sale. | inappropriate color code | device markings issue | SAD
FAB-SAD-039 | Ventilator | The ventilator will include a filter at the oxygen/air inlet to minimize the possibility of contaminants getting into the breath delivery system and causing failures of the flow control valve and flow sensors. | filter broken or damaged | contamination during use | SAD
FAB-SAD-040 | Ventilator | The ventilator will include a filter at the air inlet to minimize the possibility of contaminants getting into the breath delivery system and causing failures of the flow control valve and flow sensors. | filter broken or damaged | break | SAD
FAB-SAD-041 | Ventilator | The high-pressure gas inlet filters will be capable of replacement as a planned maintenance. This filter (not shown in Figure 1) resides inside the Flow Blender subsystem and the ventilator back cover needs to be removed in order to access this filter. NOTE: A BFE/VFE or HEPA filter is not required here since the gases are from a piped or bottled source | difficult to replace | device maintenance issue | SAD
FAB-SAD-042 | Ventilator | The ventilator will include a coalescing water trap before the air inlet to remove the liquid water from the compressed air. | water splashed | fluid leak | SAD
FAB-SAD-043 | Ventilator | The ventilator incorporates a HEPA filter and water trap at the return gas (expiratory) port of the ventilator. This water trap and exhalation filter will be a replaceable accessory; it can be accessed and replaced by the end-user without the need to remove any ventilator covers. | difficult to replace | difficult to remove; device maintenance issue | SAD
FAB-SAD-044 | Ventilator | The ventilator incorporates a heated exhalation manifold, downstream of the exhalation filter and water trap. This manifold is upstream of the exhalation flow sensor, to ensure that the flow sensor is not subjected to liquid water (condensate). | improper sensor position | Malposition of Device | SAD
FAB-SAD-045 | Ventilator | The exhalation manifold, flow sensor and exposed parts of the exhalation valve are captured within thermal insulation. | improper sensor position | Malposition of Device | SAD
FAB-SAD-046 | Ventilator | To maximize runtime with the battery, the exhalation manifold heater is operational only when the ventilator is running off AC mains power. | inappropriate operation | device handling issue | SAD
FAB-SAD-047 | Ventilator | The ventilator incorporates a solid state (non-depleting) oxygen sensor. The oxygen sensor will be used to monitor the fraction of oxygen in the gas mixture output from the ventilator. Note: the oxygen sensor is for monitoring only and is not used for servo control. | inappropriate sensor type | component incompatible | SAD
FAB-SAD-048 | Ventilator | The oxygen sensor will be used for both air/oxygen mixture. | invalid sensor | component incompatible | SAD
FAB-SAD-049 | Ventilator | A single use external inspiratory bacterial filter will be installed between the inspiratory outlet of the ventilator and the humidifier inlet or circuit if no humidifier is used. | single use violated | use of device issue | SAD
FAB-SAD-050 | Ventilator | The HWFS will be of the constant temperature (resistance) type where the temperature (and hence the resistance) of the heated wire is maintained at a constant value. The temperature of the heated wire changes initially as the flow rate varies, but the electrical power supplied by the ventilator to the heated wire will be automatically adjusted to compensate for the change in temperature and bring the temperature of the wire back to its original value. The change in the supplied electrical power to the heated wire is a measure of flow rate. | inadeqaute temperature resistance | temperature issue | SAD
FAB-SAD-051 | Ventilator | The ventilator is required to connect to an external esophageal pressure module. The ventilator will be designed to provide DC power and serial data signals to enable this function to be added. | fail to transmit signal | failure to transmit record | SAD
FAB-SAD-052 | Ventilator | The ventilator is required to connect to an internal OEM pulse oximeter module. The ventilator will be designed to provide isolated DC power and serial data signals to enable this function to be added. | fail to transmit signal | failure to transmit record | SAD
FAB-SAD-053 | Ventilator | The ventilator may be normally used with a heater / humidifier; in this use scenario, the ventilator mobile stand will provide for secure mounting and transport of the humidifier together with the ventilator. The humidifier will interface mechanically to the patient breathing circuit. The humidifier will connect to AC power directly to the wall receptacle. | fail to interface with other device | device-device incompatibility | SAD
FAB-SAD-054 | Ventilator | The ventilator will be required to connect to an electronic nebulizer, for delivery of aerosolized medications during inspiration. The ventilator will be designed to provide the drive signals for the nebulizer. | fail to interface with other device | device-device incompatibility | SAD
FAB-SAD-055 | Ventilator | The ventilator will be required to connect to an optional external OEM capnometer module. The ventilator will be designed to provide DC power and serial data signals to enable this function to be added. | fail to interface with other device | component or accessory incompatibility | SAD
FAB-SAD-056 | Ventilator | The ventilator incorporates a proximal pressure sensor. The ventilator will be designed to provide a proximal pressure port for monitoring airway pressure at the wye for infant flow circuit only. | fail to interface with other device | component or accessory incompatibility | SAD
FAB-SAD-057 | Pneumatic Architectural View | The pneumatic architecture of Fabian HFO ventilator is shown in Figure 2. The Fabian HFO ventilator is an electromechanical device, with all pneumatic functions implemented within the ventilator module, or within the breathing circuit and associated accessories. | inadequate architecture design | component missing | SAD
FAB-SAD-058 | Pneumatic Architectural View | The pneumatic architecture will comprise of three functional subsystems, the Gas Delivery, Safety, and Patient Monitoring. While all three of these subsystems interact to perform their intended functions, they provide distinctly different operations for care of the patient. | inadequate system design | component missing | SAD
FAB-SAD-059 | Pneumatic Architectural View | The Gas Delivery subsystem will provide a high pressure gas delivery for delivery of compressed ambient air, a high pressure gas delivery for delivering oxygen , a gas mixing subsystem, and an exhalation subsystem. | failure of gas delivery system | Delivery System Failure | SAD
FAB-SAD-060 | Pneumatic Architectural View | The primary component of the Safety subsystem will be 1) the exhalation valve that maintains and controls circuit pressure in conjunction with the gas delivery subsystem and provides an exit path for patient exhaled gas when the ventilator declares a Safety or Fail State condition. | fail to provide exit path | fail-safe design failure | SAD
|  | The primary component of the Safety subsystem will be 2) the pressure discharge solenoid valve that acts as backup for the exhalation valve and opens under software control in order to relieve the circuit pressure | lack of backup valve | fail-safe design failure | SAD
|  | The primary component of the Safety subsystem will be 3) the sub-ambient valve that provides patient access to ambient air when the airway pressure falls below the ambient pressure | fail to access ambient air | fail-safe design failure | SAD
|  | The primary component of the Safety subsystem will be 4) a passive pressure relief valve that opens when the outlet pressure exceeds a certain value in order to reduce the gas pressure applied to the patient. | fail to open valve to reduce pressure | fail-safe design failure | SAD
FAB-SAD-061 | Pneumatic Architectural View | The Patient Monitoring subsystem will include sensors for measuring barometric pressure, inspiratory pressure (via inspiratory pressure sensor), proximal airway pressure (via proximal pressure sensor only for infants flow), delivered oxygen, and airway flow (only for neo/infants in invasive ventilation using a hot wire flow sensor). | inaccurate barometric pressure measurement | incorrect measurement | SAD
FAB-SAD-062 | Pneumatic Architectural View | The Gas Delivery System will have the capability to provide the required flow rates and mixtures of air, oxygen to enable the ventilator to deliver a target pressure or target flow at the patient interface as required by control system based on the user settings. The Gas Delivery system consists of the two high pressure gas cassettes (Flow Cassette subsystem, one for oxygen and the other for air), inspiratory pressure sensor, and the exhalation valve. | fail to deliver required flow rate | Delivery System Failure | SAD
FAB-SAD-063 | Pneumatic Architectural View | A Flow Cassette subsystem will control the gas flow from a high-pressure gas source and deliver a controlled flow based on a commanded value (30) sent every millisecond from the main breath control servo residing on the main PCBA (VCM). | fail to control flow | infusion or flow issue | SAD
FAB-SAD-064 | Pneumatic Architectural View | The high-pressure gas supply will be a piped medical oxygen supply (100% oxygen), a high-pressure gas cylinder of pure oxygen and a piped/compressed medical air. | failure of medical gas supply | medical gas supply problem | SAD
FAB-SAD-065 | Pneumatic Architectural View | The Flow Cassette delivery system consists of a high-pressure gas inlet, gas mixture identification, a check valve, a filtration system, a gas temperature sensor, a pressure transducer, a flow sensor, and a flow control valve. | fail to mix gas | Delivery System Failure | SAD
FAB-SAD-066 | Pneumatic Architectural View | The Flow Cassette subsystem incorporates an ISO 5359 high pressure gas inlet connector and is responsible for identification (1) of the gas mixture (see XLC-SAD-1080). | fail to identify gas mixture | failure to sense | SAD
FAB-SAD-067 | Pneumatic Architectural View | The Flow Cassette is provided with coefficients for gas corrections by the VCM based on the identified gas mixture. | inappropriate gas correction | Overcorrection | SAD
FAB-SAD-068 | Pneumatic Architectural View | A check valve will be used to prevent air or other gas contaminating the hospital gas pipeline or venting to atmosphere in single point failure (no pressure at inlet). An inlet filter will be used (see XLC-SAD-1086). | fail to prevent gas contaminating | protective measure issue | SAD
FAB-SAD-069 | Pneumatic Architectural View | A gas temperature sensor, placed in the gas path, will be used to allow appropriate compensation for the flow sensor. A pressure transducer will measure the inlet pressure of the gas provided to the Flow Cassette. The pressure will be used for gas corrections and as an input to the flow servo for closed-loop control of the flow control valve. | inappropriate gas correction | incorrect measurement | SAD
FAB-SAD-070 | Pneumatic Architectural View | The inlet pressure transducer will also be part of the safety system and be used to detect the loss of the high pressure gas source. Upon loss of high pressure gas the ventilator will annunciate an alarm to notify the clinician to take action, and also reverts to operation on the alternative source alone. | fail to identify loss of pressure gas | failure to sense | SAD
FAB-SAD-071 | Pneumatic Architectural View | An internal flow sensor will be provided to monitor, and provide control for, the flow of gas to the patient outlet. This flow sensor will consist of a flow restrictor, with a differential pressure transducer across the restrictor element. | fail to restrict flow | free or unrestricted flow | SAD
FAB-SAD-072.1 | Pneumatic Architectural View | The pressure transducer incorporates a die (substrate) temperature monitoring capability to allow for appropriate compensation of flow and pressure values. | fail to compensate flow | insufficient flow or underinfusion | SAD
FAB-SAD-072.2 | Pneumatic Architectural View | A micro-controller on the Flow Cassette PCBA will use the die temperature, the raw pressure measurements, and the coefficients stored on an EEPROM on the pressure transducer to calculate corrected differential pressure and absolute pressure values, which will be then accurate over temperature and changing common mode pressure. | fail to calculate differential pressure | parameter calculation error due to software problem | SAD
FAB-SAD-072.3 | Pneumatic Architectural View | The manufacturing calibration process of the flow sensor will allow for operation with oxygen, air. | fail to calibrate oxygen | Failure to Calibrate | SAD
FAB-SAD-073 | Pneumatic Architectural View | The inlet pressure transducer reading will be used to correct the gas flow from measured conditions to SLPM conditions. The internal flow servo loop will modulate the flow control valve and maintain accurate volumetric gas delivery. The delivered flow from the Flow Cassette (3, 8) SLPM, the gas temperature (2, 7), and the inlet pressure (4, 9) will be communicated every millisecond to the upper level system (the main processor on VCM board). These measurements plus other sensor inputs such as the barometric (14) pressure will be used on by the VCM processor to convert the measured flow in SLPM to corrected conditions (BTPS) for the patient. | fail to compensate flow | insufficient flow or underinfusion | SAD
FAB-SAD-074 | Pneumatic Architectural View | The outlet flow from the air Flow Cassette and the outlet flow from the oxygen Flow Cassette are connected to a manifold incorporating a sub-ambient valve, connections for a safety valve (comprised of a pressure discharge solenoid valve and a passive pressure relief valve), connections for an inspiratory pressure sensor (including a inspiratory autozero solenoid valve), connections for a proximal pressure sensor (including a proximal autozero solenoid valve), an oxygen sensor, and then to the ventilator outlet (inspiratory port). | Connection Issue | connection issue | SAD
FAB-SAD-075 | Pneumatic Architectural View | The flows from the air Flow Cassette and from the oxygen Flow Cassette flow are flow-metrically proportioned to deliver the set oxygen concentration and real time combined output flow value (see Section 6.3). | fail to combine flow rates proportioned | improper flow or infusion | SAD
FAB-SAD-076 | Pneumatic Architectural View | The pressure relief and safety subsystem consists of a pressure discharge solenoid valve, a passive pressure relief valve, a sub-ambient valve, and the exhalation valve. | fail to open valve to reduce pressure | fail-safe design failure | SAD
FAB-SAD-077 | Pneumatic Architectural View | The ventilator will have an exhalation valve at its expiratory port. The exhalation valve will be opened under software control to relieve circuit pressure. | fail to fail-safe open | fail-safe design failure | SAD
FAB-SAD-078 | Pneumatic Architectural View | The exhalation valve will ‘fail-safe’ open in case of loss of electrical power, software control, or loss of all inlet gases, and whenever the ventilator is turned off. | fail to fail-safe open | fail-safe design failure | SAD
FAB-SAD-079 | Pneumatic Architectural View | The ventilator will have a pressure discharge solenoid valve at the outlet that will be opened under software control to relieve the circuit pressure. The pressure discharge solenoid valve is intended as a backup to the exhalation valve and it will be activated in cases when opening the exhalation valve alone is not able or sufficient to relieve the circuit pressure. | fail to switch backup valve | fail-safe design failure | SAD
FAB-SAD-080 | Pneumatic Architectural View | To prevent applying excessive gas pressures to the patient, the gas outlet incorporates a passive pressure relief valve that will open if the outlet pressure exceeds 100 cmH2O and will prevent the outlet pressure from exceeding 125 cmH2O. | fail to trigger valve to reduce pressure | fail-safe design failure | SAD
FAB-SAD-081 | Pneumatic Architectural View | The gas outlet incorporates a passive sub-ambient valve that will open if the outlet pressure falls below the ambient pressure in order to allow ambient air to be available to patient. | fail to trigger valve to reduce pressure | fail-safe design failure | SAD
FAB-SAD-082 | Pneumatic Architectural View | During a ventilator Fail-Safe state condition, the exhalation valve, the sub-ambient valve, and the mask sub-ambient relief port (if applicable) will work in conjunction to: • Relieve pressure from the circuit down to ambient pressure conditions • Allow ambient air to be available to the patient for breathing • Minimize re-breathing of gases (function of patient circuit and application) | fail to trigger port to reduce pressure | fail-safe design failure | SAD
FAB-SAD-083 | Pneumatic Architectural View | The exhalation subsystem consists of a heated manifold incorporating an exhalation valve and an X-Flow Sensor, and an unheated filter and water trap. The filter and water trap are as described in XLC-SAD-1092. | inadequate system design | component missing | SAD
FAB-SAD-084 | Pneumatic Architectural View | The exhalation valve is a software-controlled valve that is used to adjust the flow of gas passed through the expiratory port of the ventilator to outside environment. | fail to adjust gas flow | improper flow or infusion | SAD
|  | The exhalation valve is controlled by a software control signal (34) and works in conjunction with the gas delivery subsystems to maintain user set pressure control levels. In exhalation, in CPAP therapy, in TCPL ventilation, and in APRV the exhalation valve is maintained at a set pressure, and outlet flow is servo controlled to a specified target bias flow rate. Also, additional (demand) flow is provided to maintain the pressure in the event of patient inspiratory flow exceeding the bias flow. | fail to maintain target flow rate | improper flow or infusion | SAD
FAB-SAD-085 | Pneumatic Architectural View | The exhalation valve operates on the principle of a force balance across the control diaphragm, which is a disposable valve membrane. A linear magneto-mechanical actuator controls the force on the diaphragm, which in turn controls the circuit pressure. The force generated by the actuator is based on a command (34) from the software closed-loop controller. The exhalation valve poppet position is measured and provided as an output (10). | inappropriate force generated | mechanical issue | SAD
FAB-SAD-086 | Pneumatic Architectural View | When a proximal flow sensor is not in use, airway flow is estimated from measurement of flow in the exhalation subsystem, circuit compliance volume change, and the total outlet flow (minus the known flow to the oxygen sensor). Gas flows in the exhalation subsystem are measured under conditions which vary depending on the breathing circuit configuration. With an active humidifier, gas reaching the exhalation flow sensor remains saturated; the BTPS correction is only required for gas temperature, which is measured within the X-Flow Sensor. | fail to measure flow | insufficient flow or underinfusion | SAD
FAB-SAD-087 | Pneumatic Architectural View | The exhalation valve diaphragm will be capable of replacement as a planned maintenance. This exhalation valve diaphragm can be accessed after removing the covers. | difficult to replace | device maintenance issue | SAD
FAB-SAD-088 | Pneumatic Architectural View | The exhalation X-Flow Sensor is primarily used to determine the patient flow, which is calculated by subtraction of exhalation flow, leak flow, and circuit compliance volume change from the outlet flow. | fail to determine flow | parameter calculation error due to software problem | SAD
FAB-SAD-089 | Pneumatic Architectural View | The Exhalation X-Flow Sensor measures the gas temperature, the pressure, and the flow at the exhalation valve inlet. Effects of back pressure between the exhalation inlet and the exhalation X-Flow Sensor will be considered in the measurement. | fail to measure temperature | incorrect measurement | SAD
FAB-SAD-090 | Pneumatic Architectural View | The Exhalation X-Flow Sensor is provided with coefficients for gas corrections by the VCM, based on the current output mixture of air and high pressure inlet gas and the gas type for the high pressure inlet (air and Oxygen mixture). | fail to compensate flow | improper flow or infusion | SAD
FAB-SAD-091 | Pneumatic Architectural View | The STPD21 corrected flow (14), the gas temperature (15), and the exhalation pressure (13) will be reported from the X-Flow Sensor, as requested by the upper level system (the main processor on VCM board). | fail to compensate flow | improper flow or infusion | SAD
FAB-SAD-092 | Pneumatic Architectural View | A heater is provided downstream of the water trap and filter, and upstream of the exhalation flow sensor. The heater power (35) is controlled to warm the exhalation gas in order to prevent condensation within the flow sensor or exhalation valve. | fail to warm gas | improper flow or infusion | SAD
FAB-SAD-093 | Pneumatic Architectural View | The Patient Monitoring system will include sensors for measuring barometric pressure, machine outlet and exhalation port pressure, exhalation port flow, airway flow (with infant hot-wire flow sensor only), and the delivered oxygen percent. | inadequate system design | component missing | SAD
FAB-SAD-094 | Pneumatic Architectural View | A barometric pressure sensor (12) will provide measurements to the main breath control system or any other device components requiring this data. Both pressure measurements are used to correct gases to the required patient conditions. | inaccurate barometric pressure measurement | incorrect measurement | SAD
FAB-SAD-095 | Pneumatic Architectural View | The proximal and Machine pressure sensor (13) is integral to the operation of the exhalation valve and the safety systems. | fail to interface with other device | device-device incompatibility | SAD
FAB-SAD-096 | Pneumatic Architectural View | A hot wire proximal flow sensor (17) is used only for neo/infant patients for invasive ventilation. The hot wire flow sensor (HWFS) resides at the interface to the patient and measures flow bi-directionally. | fail to interface with other device | device-device incompatibility | SAD
FAB-SAD-097 | Pneumatic Architectural View | Inspiratory pressure transducer in the inspiratory manifold and the exhalation X-Flow Sensor provide for redundancy of pressure monitoring, and detection of circuit occlusion faults. | fail to detect occlusion | occlusion within device | SAD
FAB-SAD-098 | Pneumatic Architectural View | The oxygen sensor will be used to monitor the fraction of oxygen (11) in the gas mixture output from the ventilator and it is used for both air/oxygen mixture. | fail to monitor fraction of oxygen | incorrect measurement | SAD
FAB-SAD-099 | Pneumatic Architectural View | The oxygen sensor measured oxygen values should be processed in software to correct for pressure and for vibration. | fail to compensate for vibration | imprecision | SAD
FAB-SAD-100 | Pneumatic Architectural View | In order to maintain the performance of the oxygen sensor within specification, this paramagnetic oxygen sensor requires calibration using the known supply gas and air. At appropriate intervals, the ventilator software will ask the operator to run an extended self test (EST), which will include the calibration of the oxygen sensor. The ‘calibration’ defined here is the process of obtaining the zero (21% O2) and span (100% O2) values of the oxygen sensor and storing these values in the ventilator for use in normal operation. | fail to provide calibration gas and air | Failure to Calibrate | SAD
FAB-SAD-101 | Pneumatic Architectural View | The patient circuit forms an integral part of the pneumatic architecture, incorporating components for humidification, pressure relief, and pressure and flow measurement of gas to and from the patient. | incorrect circuit type | component incompatible | SAD
FAB-SAD-102 | Pneumatic Architectural View | The breathing circuit consists of the gas-wetted components of an optional dry gas conduit, an optional humidifier chamber, a gas conduit to convey breathing gas from the ventilator outlet to the patient airway interface, a wye connector (in the case of dual-limb circuits), a patient interface (invasive or non-invasive), and a gas conduit to return the exhaled gas from the patient to the ventilator (in the case of dual-limb circuits). | component incompatible | fitting problem | SAD
FAB-SAD-103 | Pneumatic Architectural View | When used with an active humidification system, the breathing circuit incorporates a dry gas conduit (to convey breathing gas from the ventilator outlet to the humidifier), a humidifier chamber (to attach to the active humidifier for addition of heat and water vapor, the inspiratory gas conduit incorporating a heater wire, and a remote temperature probe port included in the inspiratory tubing no more than 18” from the patient interface. | integration failed | incompatibility problem | SAD
FAB-SAD-104 | Pneumatic Architectural View | Passive humidification (HME) can be used instead of active humidification. When used with a Heat and Moisture Exchanger (HME) the HME will fit between the wye and the tracheal tube connector (ETT or tracheostomy tube). | inadequate humidification | Moisture or Humidity Problem | SAD
FAB-SAD-105 | Pneumatic Architectural View | The breathing circuit incorporates a 22mm ISO 5356-1 taper connection to the patient mask. Some of the typical breathing circuit configurations for invasive and non-invasive ventilations are shown in the following sections. | inappropriate circuit configuration | component incompatible | SAD
FAB-SAD-106 | Electrical Architecture View | The Main Board contains the control software for the ventilation procedure and provides connectivity to the PCBAs that are dedicated to major functionalities required for the ventilation procedure. It determines the necessary controls for a programmed ventilation procedure. | incompatible connection | fitting problem | SAD
FAB-SAD-107 | Electrical Architecture View | The Main Board is responsible for distributing electrical power to other system components, integrating electrical signals (e.g. transducer signal, user input, etc.) from them, and providing control signal for the ventilation procedure. | false output | False Device Output | SAD
FAB-SAD-108 | Electrical Architecture View | The Main Board incorporates connectivity to NetDCU. | incompatible connection | fitting problem | SAD
FAB-SAD-109 | Electrical Architecture View | The Fabian HFO Ventilator utilizes both analog and digital components and circuitry. The Main Board incorporates a watchdog circuit for monitoring the programmable electrical components. In case of a failure, the watchdog opens the pressure relief valves and stops power to the Blender module to prevent and/or relieve pressure in the Patient Circuit. | fail to activate protective actions | protective measure issue | SAD
|  | The watchdog circuit monitors for failures of the following components, these are also integral parts of the main board. | fail to detect error | protective measure issue | SAD
FAB-SAD-110 | Electrical Architecture View | The monitoring microprocessor communicates with the User Interface Microprocessor using a SPI connection. It receives the ventilator settings from the UIM and transfers all graphic data (proximal pressure, flow, volume) values and the monitored ventilator state back to the UIM. | loss of communication | Communication or Transmission Issue | SAD
FAB-SAD-111 | Electrical Architecture View | Diversity and redundant monitoring of the most important parameters is achieved by using a different data path and measurements independent of the control microprocessor. Measurements:  air input pressure  oxygen input pressure  blender output pressure  patient proximal pressure  flow at the y-Piece  oxygen sensor 1  oxygen sensor 2  supply voltage | fail to switch redundant processor | protective measure issue | SAD
FAB-SAD-112 | Electrical Architecture View | The Monitor PIC incorporates a SPI for communication with the NetDCU. The Monitor PIC incorporates an interface for receiving tachometer signal from the chassis fan. The Monitor PIC incorporates an analog interface for connection to the O2 Sensor. The Monitor PIC incorporates an analog interface for connection to the HWFS Connector. | loss of connection | connection issue | SAD
FAB-SAD-113 | Electrical Architecture View | The Monitor PIC incorporates an analog interface to 2 x calibration potentiometers, which is located on the Manifold Board, for the Inspiratory Pressure Transducer. Each potentiometer is used for calibration of the transducer. | loss of connection | connection issue | SAD
FAB-SAD-114 | Electrical Architecture View | The Monitor PIC incorporates an I2C interface for connection to a FRAM. The FRAM stores calibration data, which persists through power cycles, for the following components:  Inspiratory Pressure Transducer  Proximal Pressure Transducer  Blender Pressure Regulators  Exhalation valve  HWFS | fail to support accessory | accessory incompatible | SAD
FAB-SAD-115 | Electrical Architecture View | The Monitor PIC incorporates Tacho connection with the fan module. It regulates the speed of the fan to prohibit the high heat inside the ventilator body. It also prevent ignition due to high concentration of oxygen inside the device. | ignition | fire | SAD
FAB-SAD-116 | Electrical Architecture View | The control microprocessor PIC2 receives the ventilator settings via RS232 and controls the blender, bleed valve, HFO Module and the exhalation valve. Measurements:  air input pressure  oxygen input pressure  blender output pressure  patient proximal pressure | fail to trigger valve to reduce pressure | protective measure issue | SAD
FAB-SAD-117 | Electrical Architecture View | The Control PIC incorporates a PWM input to the Exhalation Valve via a valve driver connection, which is integrated into the Manifold Board. | no input | failure to read input signal | SAD
FAB-SAD-118 | Electrical Architecture View | The Control PIC incorporates an analog interface for connection to the Proximal Pressure Transducer. The Control PIC incorporates an analog interface for connection to the Inspiratory Pressure Transducer. | loss of connection | connection issue | SAD
FAB-SAD-119 | Electrical Architecture View | The Control PIC incorporates an analog interface for connection to the Inspiratory Solenoid Valve via a valve driver connection, which is integrated into the Manifold Board. | fail to support accessory | accessory incompatible | SAD
FAB-SAD-120 | Electrical Architecture View | The Control PIC incorporates a digital IO output for control of the Electrical Safety Valve. The Control PIC has connection with HFO Module via SPI synchronous serial communication interface. | connection detached | disconnection | SAD
FAB-SAD-121 | Electrical Architecture View | The Control PIC incorporates an analog communication interface with a Pressure Transducer which again integrated with Blender Board. The Control PIC incorporates a hardware communication protocol UART with NetDCU to control different peripherals. | incompatible connection | fitting problem | SAD
FAB-SAD-122 | Electrical Architecture View | The Net DCU is a master processor single board computer based on a RISC-Processor (CPU: Intel PXA270, Architecture ARM9) produced by F&S, Stuttgart, Germany, for industrial use in the visualization and controlling of complex systems. | processor overheat | overheating of device or device component | SAD
FAB-SAD-123 | Electrical Architecture View | The major functional features of the NetDCU used in the design of the Fabian HFO are:  Windows CE 6.0  CPU: Intel PXA270, Architecture ARM9  Frequency: 520  Video optional SM502 or CPU, Video in optional  Flash 64MB, RAM 64MB  Ethernet 10/100M o 3 x RS232, USB, I2C, SPI, Parallel Interface, Digital IO, Matrix Keyboard  Max.Resolution : SVGA ( optional SXGA) | low calculation speed | electronic property issue | SAD
FAB-SAD-124 | Electrical Architecture View | NetDCU is directly mounted on the Main Board, it is an SOM. NetDCU incorporates a digital IO input connection for communication to the UDP. NetDCU incorporates a digital IO input connection for communication to the Hardware Keys. NetDCU incorporates a USB connection for communication to an external USB device (i.e. nebulizer or for SW update). NetDCU incorporates a SPI for communication with the Monitor PIC. NetDCU provides PWM signal for control of the chassis fan. NetDCU incorporates a UART for communication to the EtCO2 Sensor Connector. | fail to bidirectional communication | Communication or Transmission Issue | SAD
FAB-SAD-125 | Electrical Architecture View | NetDCU incorporates a serial connection for communication to the SpO2 Sensor Connector via a USB-to-Serial Converter, which is integrated into the Sensor Connector Board. NetDCU incorporates a HDMI-Out connection NetDCU incorporates a RS-232 interface for the PDMS. NetDCU incorporates a UART for communication with the Control PIC. NetDCU incorporates a DIO interface for communication with the Alarm PIC. NetDCU incorporates an I2C input to the Power Supply Board. NetDCU incorporates a PWM input to the Display Backlight. | fail to unidirectional communication | Communication or Transmission Issue | SAD
FAB-SAD-126 | Electrical Architecture View | NetDCU incorporates an integrated touch controller for analog touch signal from the touchscreen. NetDCU incorporates a RGB interface to provide drive signal to the LCD display. NetDCU incorporates connectivity to an exchangeable SD memory card, which provides storage for trend & log files and special font characters. The log and trend files may be extracted via a WinCE-compatible USB device. NetDCU incorporates a real-time clock. | fail to interface with other device | component or accessory incompatibility | SAD
FAB-SAD-127 | Electrical Architecture View | The Alarm PIC incorporates a I2C interface for connection to the Speaker Driver, which connects to a downstream Loudspeaker via an analog interface. | incompatible accessory | Communication or Transmission Issue | SAD
FAB-SAD-128 | Electrical Architecture View | The Alarm PIC incorporates a digital IO output for control of the Alarm LED. The Alarm PIC incorporates a DIO interface for communication with NetDCU. | incompatible accessory | Communication or Transmission Issue | SAD
FAB-SAD-129 | Electrical Architecture View | Monitor PIC updates every 100 ms the status of flow sensor, oxygen sensor and gas supply pressure. The Alarm is triggered by the master processor. | long response time | decreased sensitivity | SAD
FAB-SAD-130 | Electrical Architecture View | In case of a Sensor defect or problem, an alarm is triggered and the calibration screen opens automatically. A recalibration can be done, if it fails the sensor could be changed or deactivated. | fail to activate alarm | device alarm system issue | SAD
FAB-SAD-131 | Electrical Architecture View | With a deactivated flow sensor, only ventilation modes without trigger function can be used. | incorrect ventilation mode | failure to cycle | SAD
FAB-SAD-132 | Electrical Architecture View | In case of a valve or flow sensor defect in the blender unit, a wrong gas mixture will be detected by the oxygen sensor. A defect or not right calibrated oxygen sensor will also be detected. In both cases an alarm is triggered and the calibration screen opens automatically. | fail to activate alarm | device alarm system issue | SAD
FAB-SAD-133 | Electrical Architecture View | The gas supply (air and oxygen) is supervised by pressure sensors. If the pressure is lower than 0.8bar, an alarm is triggered by the pressure sensors. | failure of medical gas supply | medical gas supply problem | SAD
FAB-SAD-134 | Electrical Architecture View | In case of a defect pressure sensor the loss of one gas supply is secondary detected by the O2 Sensor because of the differ between set and monitored gas mixture. | fail to identify loss of pressure gas | failure to sense | SAD
FAB-SAD-135 | Electrical Architecture View | The loss of both gas supplies is secondary detected by the alarm settings because of a decreasing pressure and flow | inappropriate alarm settings | device alarm system issue | SAD
FAB-SAD-136 | Electrical Architecture View | When mains power fails when the device is switched on, the battery delivers sufficient energy for an acoustical and optical power fail alarm for a minimum time of 10 minutes. | reduced battery capacity | Low Battery | SAD
FAB-SAD-137 | Electrical Architecture View | The fan is an important component to prevent ignition because of high oxygen concentration in the device, in case of a defect in the pneumatic system. It is supervised by the Monitor PIC by a speed signal with an upper limit (speed > 190Hz) and a lower limit (speed < 80Hz). The signal is measured every 100ms, the filter for alarm triggering is 5seconds. | low fan speed | improper device output | SAD
FAB-SAD-138 | Electrical Architecture View | The Sensor Connection Board incorporates an isolated pass-through communications interface from the below-listed accessory connections to the Monitor PIC/NetDCU. | signal interference | ambient noise issue | SAD
FAB-SAD-139 | Electrical Architecture View | The Sensor Connection Boards provides electrical circuitry and isolated DC power (5 V and 3.33 V) for the following accessory connections: SpO2 module- | noise introduced | Ambient Noise Issue | SAD
FAB-SAD-140 | Electrical Architecture View | To enable the SpO2 measurement, the module must be enabled. This could be done in the menu of the SpO2 module. For the SpO2 module, the Sensor Connection Board incorporates a USB-to-Serial converter which receives USB communication from NetDCU and outputs UART for interface to the SpO2 Sensor Connector. | fail to enable module | use of incorrect control settings | SAD
FAB-SAD-141 | Electrical Architecture View | The Fabian HFO supports two different CO2 module types to provide CO2 monitoring: 1. Microstream, MicroPod™ External etCO2 Module from Oridion® (Covidien/Medtronic) 2. LoFlo Sidestream etCO2 Sensor from Respironics (Philips). | incompatible module | component incompatible | SAD
FAB-SAD-142 | Electrical Architecture View | The main board contains the analog circuits of the hot-wire sensor interface; the 12-bit analog-to-digital converter of the monitor PIC is used. | inappropriate converter designed | component or accessory incompatibility | SAD
FAB-SAD-143 | Electrical Architecture View | The Fabian HFO flow sensor is based on a hot wire anemometer with two platinum wires to determine flow and direction. | fail to determine direction | failure to sense | SAD
FAB-SAD-144 | Electrical Architecture View | Each hot wire is arranged in its own Wheatstone bridge circuit. The bridge supplies a sensor heating current - that varies with the flow velocity - and maintains constant sensor resistance and, thus, constant sensor temperature. | fail to maintain resistance | physical resistance | SAD
FAB-SAD-145 | Electrical Architecture View | Flow applied to the hot wire cools the sensor and unbalances the bridge. As a result, the electronic circuits increase the bridge voltage until the nominal resistance (and temperature) is reached again. The controlling microprocessor can switch off the bridge supply when overload is detected. | sensor temperature dropped | temperature issue | SAD
FAB-SAD-146 | Electrical Architecture View | The Blender Board provides the electrical circuitry required for the control of the Air and O2 Proportional Valves. | inappropriate oxygen fraction | medical gas supply problem | SAD
FAB-SAD-147 | Electrical Architecture View | For air and oxygen, 1 electronically controlled valve (5) each is used in the blender (4) to allow for sufficiently exact flow values. The valves are proportional type and can be controlled from zero to the maximum flow. So, the exact oxygen concentration can be achieved. Each valve is driven by a switch mode regulator. | low oxygen fraction accuracy | medical gas supply problem | SAD
FAB-SAD-148 | Electrical Architecture View | The actual flow of each path is being measured by a flow sensor similar to the patient flow sensor. A local control loop regulates the actual flow to the desired level. | fail to regulate flow to sesired level | improper flow or infusion | SAD
FAB-SAD-149 | Electrical Architecture View | There are two more solenoid valves on the blender module: For calibration of the oxygen sensor, a pneumatic bypass can be activated by opening either the air- or the oxygen calibration valve. A step up converter delivers 24 V to the calibration valves. | fail to trigger calibration | Failure to Calibrate | SAD
FAB-SAD-150 | Electrical Architecture View | The Blender Board also incorporates the Blender PIC, which is described below. The PIC microprocessor sets the two blender valves for delivery of the desired oxygen concentration using a SPI interface to the blender module with the signals MCS, MDATA, and MCLK. | fail to interface with other component | component or accessory incompatibility | SAD
FAB-SAD-151 | Electrical Architecture View | The Blender PIC incorporates a SPI interface for connection to the Control PIC within the Main Board. The Blender PIC incorporates an analog interface for connection to the Air Flow Sensor The Blender PIC incorporates a PWM input for the control of the O2 Proportional Valve. The Blender PIC incorporates a PWM input for the control of the Air Proportional Valve. The Blender PIC incorporates an analog interface for connection to the O2 Flow Sensor. The Blender PIC incorporates a PWM input for the control of the O2 Calibration Valve. The Blender PIC incorporates a PWM input for the control of the Air Calibration Valve. | fail to interface with other component | component or accessory incompatibility | SAD
FAB-SAD-152 | Electrical Architecture View | The Blender PIC incorporates a I2C interface for connection to a FRAM. The FRAM stores persistent calibration data. The Blender Board incorporates 2 x calibration potentiometers for the calibration of the Air Proportional Valve & Flow Sensor (21%) and the O2 Proportional Valve & Flow Sensor (100%). | fail to interface with other component | component or accessory incompatibility | SAD
FAB-SAD-153 | Electrical Architecture View | The Manifold Board incorporates signals from the Proximal Pressure Transducer and the Inspiratory Pressure Transducer and relays the signals to the Control PIC within the Main Board. The Manifold incorporates valve drivers for the Inspiratory Solenoid Valve and the Exhalation Valve. The Manifold Board incorporates 2 x calibration potentiometers for the calibration of the Inspiratory Pressure Transducer. The Manifold incorporates an outlet switching relay connected with HFO module. | fail to interface with other component | component or accessory incompatibility | SAD
FAB-SAD-154 | Electrical Architecture View | The device can be operated from an external power source but features an internal battery (NiMH Power pack) also. Power module is the sum of the following components: - Power PCB - Powerpack NiMH/ Inner battery - Power supply “XP” - Power inlet module The Power Board comprises of Power PIC, which is described below. | missing part | Component Missing | SAD
FAB-SAD-155 | Electrical Architecture View | The Power PIC incorporates circuitry for power selection, DC / DC conversion, battery charge control, and all power supply-related monitoring and control mechanisms. The Power PIC incorporates a digital IO input from the On/Off hardware key. The Power PIC defaults to AC input power when available. When it is not available, the Power PIC can switch to the battery power without user-action and without disruption to the ventilation procedure The Power PIC incorporates a digital IO output to the Main LED within the User Control Interface. The Power PIC incorporates a digital IO output to the Battery Status LED within the User Control Interface. | fail to interface with other component | component or accessory incompatibility | SAD
FAB-SAD-156 | Electrical Architecture View | The Ventilator incorporates an internal battery (NiMH Power pack) that provides DC power to the Power Board (16.8 VDC nominal). The battery connects to the Power Board via an AKKU cable. | fail to interface with other component | component or accessory incompatibility | SAD
FAB-SAD-157 | Electrical Architecture View | The AC/DC Power Supply transforms the input from an external AC power supply (100~240 VAC) to 24 VDC. The Power Supply Board provides DC power to the Power Board (24 VDC). | voltage out of range | electrical issue | SAD
FAB-SAD-158 | Electrical Architecture View | The Potential Equalization Pin is for additional safety and can be connected to an equi-potential zone. Adhere to local guidelines when using this PIN. The guidelines may vary between countries, localities and power companies. Always keep the PIN for Potential equalization accessible. | pin detached | protective measure issue | SAD
FAB-SAD-159 | Electrical Architecture View | The HFO Module’s pneumatic consists of an oscillator and an HFO open valve. An electromagnetic drive together with a flexible membrane forms the oscillator. It delivers the sinus oscillation and is via the HFO open valve connected to the patient circuit. A CLASS-D amplifier is used to drive the oscillator. From the electrical point of view HFO module is connected with main board and outlet switching relay of the manifold. Controller PIC controls the HFO modules to generate required high frequency oscillation. Controller PIC Data is sent via a local SPI to the HFO-Module. Data include: HFO Frequency, I / E-ratio, Amplitude The measured amplitude is compared with the setting value and a correcting value is delivered to the HFO-PIC. In this way it makes a control loop. | loss of function | no device output | SAD
FAB-SAD-160 | Electrical Architecture View | Video In (VGA Port) cable to be connected to the Mainboard. The video option is not enabled yet. The digitalization of analog VideoIn signals results from Philips plug-in SAA7114. The SAA7114 is a video capture device for applications at the image port of Video Graphics Array (VGA) controllers. | loose connection | loose or intermittent connection | SAD
FAB-SAD-161 | Electrical Architecture View | The SAA7114 is a combination of a two-channel analog preprocessing circuit including source selection, anti-aliasing filter and Analog-to-Digital Converter (ADC), an automatic clamp and gain control, a Clock Generation Circuit (CGC), a digital multi-standard decoder containing two-dimensional chrominance/luminance separation by an adaptive comb filter and a high performance scaler, including variable horizontal and vertical up and downscaling and a brightness, contrast and saturation control circuit. | noise in channel | radio signal problem | SAD
FAB-SAD-162 | Electrical Architecture View | The fan is an important component to prevent ignition because of high oxygen concentration in the device, in case of a defect in the pneumatic system. | low fan speed | fire | SAD
FAB-SAD-163 | Electrical Architecture View | It is supervised by the Monitor PIC by a speed signal with an upper limit (speed > 190Hz) and a lower limit (speed < 80Hz). The signal is measured every 100ms, the filter for alarm triggering is 5seconds. | speed signal out of range | electronic property issue | SAD
FAB-SAD-164 | Electrical Architecture View | An external supply with a voltage of 12 – 24 volts must be connected to the back of the device in this case. VDC Connection to EXT +/-, which connects with the Power PCB. This is the connection for 24V DC external power supply. It is not for battery charging, for battery charging there is a separate connection form main board. It is for the ventilator operation. | misconnection | misconnection | SAD
FAB-SAD-165 | Electrical Architecture View | An independent thread toggles through via I2C line every 40ms. If due to an error the pneumatic control software spends excessive time in an outside routine or collapses completely, these pulses will fail, and therefore a system failure alarm will be triggered by hardware. | fail to trigger alarm by hardware | not audible alarm | SAD
FAB-SAD-166 | Electrical Architecture View | The following signals are monitored by hardware: • WD0, WD1, WD2, WD2 : square wave signals from the micro controllers a upper and lower frequency limit ensures proper function Tolerances: approx. ± 20% • WD 0: Frequency: minimum 10 Hz, maximum 450 Hz, T_abfall = 5 s • WD 1,2,3: Frequency: minimum 50 Hz, maximum 450 Hz, T_abfall = 1 | fail to monitor by hardware | failure to read input signal | SAD
FAB-SAD-167 | Electrical Architecture View | These signals are monitored by a dedicated hardware circuit. If one or more fail, the following actions take place:  System alarm  Drop of safety relay  Opening of the bleed valve  Shut down of the blender | fail to monitor by hardware | failure to read input signal | SAD
FAB-SAD-168 | Electrical Architecture View | The sound generator uses a magnetic speaker and a full bridge amplifier. It generates the sound frequency and sets the adjustable amplifier to the desired level. It also activates the system fail-LED in case of a system fail caused by the Hardware watchdog circuit. The different sound sequences get activated from the IU via parallel bus and. | incorrect sound pattern | use of incorrect control settings; electronic property issue | SAD
FAB-SAD-169 | Electrical Architecture View | It is expected that a Class I rated supply will be used, thus PE is required to allow for 2 MOP on primary. The ventilator system shall be Class I – internally powered. Any PE conductive parts meet Class I PE requirements. Electric isolation circuit | loose protective earth conductor | loose or intermittent connection | SAD
FAB-SAD-170 | Electrical Architecture View | One means of patient protection between applied parts as per Clause 8.5.2.1 for a working voltage equal to the maximum mains voltage. | loss of electrical isolation | electrical shorting | SAD
FAB-SAD-171 | Electrical Architecture View | Multiplication factor applied to CL distances. 10,000ft --> 3,048m --> 1.29x for MOOP, 1.14x for MOPP. | CL distance out of range | environmental control or utility issue | SAD
FAB-SAD-172 | Electrical Architecture View | Where a working voltage for a circuit is below 24V, 24V is used in the analysis (conservative). | voltage out of range | electrical issue | SAD
FAB-SAD-173 | Electrical Architecture View | IFU contraindicates the use of SIP/SOPs while the ventilator is not protectively earthed. | electrical shock | Inappropriate Shock | SAD
FAB-SAD-174 | Electrical Architecture View | All applied parts, except SpO2, meet this requirement by means of solid insulation and physical separation of the patient from electrical circuitry inherent to the way the applied parts are used. SpO2 meets this requirement by means of separation of electrical circuitry to the chassis within the system. | fail to solid insulation of patient | electrical shorting | SAD
FAB-SAD-175 | Mechanical Packaging Architectural View | The ventilator is packaged in a mechanical enclosure to meet the environmental requirements specified in the Fabian HFO PRS. The Ventilator provides capability for mechanical interface to the Fabian mobile cart or for shelf use. | Hazardous materials released into environment | environmental control or utility issue | SAD
FAB-SAD-176 | Mechanical Packaging Architectural View | The Ventilator incorporates a sheet metal chassis that provides a mechanical interface for all components of the ventilator assembly, including the PCB assemblies, electronic components, and the pneumatics components. Assemblies/components that are mounted on the metal chassis are oriented in a way that optimizes serviceability of the Ventilator. | weak mechanical interface strength | mechanical issue | SAD
FAB-SAD-177 | Mechanical Packaging Architectural View | Following PCBAs/components are accessible on removing the housing as follows: Top cover: Mainboard, NetDCU; Right cover (looking the Fabian from behind): Proportional gas blender with O2 sensor (O2 sensor cell, Proportional valves, Calibration valves, Input pressure regulators, Blender PCB, Inner flow sensor), Flow sensor connector PCB; Left cover (looking the Fabian from behind): Power PCB, Powerpack/battery pack; Bottom cover: HFO module; Rear cover: Loudspeaker, Fan, Power inlet module, Power Supply ‘XP’, Video connector Board | difficult to remove | difficult to remove | SAD
FAB-SAD-178 | Mechanical Packaging Architectural View | The housing allows access to expiratory membrane, proximal connector, user knob and main fuses, for replacement without the need to remove housing. | fail to replace component | device maintenance issue | SAD
FAB-SAD-179 | Mechanical Packaging Architectural View | The chassis also provides mechanical interface for the integrated user control interface components. The user control interface is designed to be installed with a vertical tilt for ease of viewing. The tilt angle is not adjustable. | weak mechanical interface strength | mechanical issue | SAD
|  | The external housing of the Ventilator incorporates plastic panels that augment the protection of internal components from environmental exposure. An opening on the rear side of the Ventilator allows for ambient air to flow into the fan and subsequently into the Ventilator. A separate exhaust opening (on the bottom side) also allows for the relief of circulating air from the fan and for the gas passing through the exhalation valve into the atmosphere. | insufficient space | inadequate storage | SAD
FAB-SAD-180 | Mechanical Packaging Architectural View | The external packaging includes hand-hold area(s) to help with lifting the ventilator. | inappropriate package design | human-device interface issue | SAD
FAB-SAD-181 | Mechanical Packaging Architectural View | The front side of the housing incorporates the following user interfaces: • Manifold connectors o Inspiratory limb port/HFO port o Expiratory limb port o Proximal pressure port o External bias flow port (Fresh Gas) and port for nCPAP system based on flow generators (single limb systems) | inappropriate hosing design | human-device interface issue | SAD
FAB-SAD-182 | Mechanical Packaging Architectural View | • User control interface o Graphical user Interface  Touchscreen  LED Indicators o Key/Control Panel  Hardware Keys • Home • Graphics • Manual Breath • Alarm Limits • Alarm Silence • Menu/Calibration • Nebulizer • O2 flush • Start/Stop • ON/OFF  Rotary Push Knob | confused interface | inadequate user interface | SAD
| Market | Pre-op planning is a critical step to ensure safety, efficacy and efficiency of resection procedures. | inadequate pre-op planning | application program issue | Auris
| Market | Pre-op planning is intended to inform 3 clinical decisions as framed by surgeons: “where to enter”, “how to move” and the “amount of resection”. | incorrect clinical decision | Poor Quality Image | Auris
| Market | Specifically, when reviewing pre-operative images, surgeons are looking for: 1. Size, location and vascularity of the tumor(s) | poor tumor image | Poor Quality Image | Auris
| Market | Larger tumors are more concerning for spread. Anything above 6 or 7 cm might mean a full removal and how it sits on the kidney in relationship to other organs. | tumor removed inappropriately | Poor Quality Image | Auris
| Market | Where is the cancer? Has it invaded the outside wall? Where has it invaded? What structures have been invaded? | fail to visualize cancer location | poor quality image | Auris
| Market | My focus is to see the proper anatomy of the mass and its vascular anatomy | poor anatomy image | optical issue | Auris
| Market | Specifically, when reviewing pre-operative images, surgeons are looking for: 2. Tumor relationship to surrounding critical structures (i.e. vasculature, ureter, chest wall, etc.) | fail to visualize surrounding critical structures | Poor Quality Image | Auris
| Market | I need to avoid critical structures like the ureter. Knowing where it is can speed things up for us. | fail to visualize surrounding critical structures | Poor Quality Image | Auris
| Market | I’m looking for the relationship of tumor to surrounding structures, particularly the vasculature. The hepatic veins and vena cava are the bigger ones for liver | fail to visualize surrounding critical structures | Poor Quality Image | Auris
| Market | Specifically, when reviewing pre-operative images, surgeons are looking for: 3. Patient-specific anatomical variations or aberrant anatomy. | fail to visualize anatomical variation | poor quality image | Auris
| Market | There is a great deal of variety within the anatomy. This is what we're looking for in the CT | fail to visualize anatomy variety | poor quality image | Auris
| Market | Because of the anatomical variability, we are often operating blindly. We don't know how many bronchi we will come into.” | fail to visualize anatomy variety | poor quality image | Auris
| Market | By integrating 3D models and patient simulation from Visible Patient into J&J ecosystem, we’ll provide the following value to the clinicians: | inadequate 3D model | application program issue | Auris
| Market | we’ll provide the following value to the clinicians: Highly detailed, virtual 3D model, using patients’ pre-op images (CT, MRI) | fail to develop detailed 3D model | application program issue | Auris
| Market | we’ll provide the following value to the clinicians: Color-coded, interactive experience that allows you to explore and visualize spatial relationships between tumors, critical structures and surrounding anatomy of your patient | fail to visualize spatial relationship | image display error | Auris
| Market | we’ll provide the following value to the clinicians: Easy-to-use interface that is accessible on your team’s PC/Mac, tablet, and mobile devices for reference in the OR, at home, and on-the-go | difficult to use | human-device interface issue | Auris
| Market | we’ll provide the following value to the clinicians: Allows you to explore anatomy on your schedule, without reliance on interpretation by your radiology team | fail to explore anatomy independently | human-device interface issue | Auris
| Market | we’ll provide the following value to the clinicians: A virtual clip-applying tool that pre-operatively identifies where tumors are located within the anatomy and provides a real-time assessment of post-operative organ volume | fail to provide virtual clip-applying tool | human-device interface issue | Auris
| Market | we’ll provide the following value to the clinicians: A virtual clip-applying tool that pre-operatively identifies where tumors are located within the anatomy and provides a real-time assessment of post-operative organ volume | fail to real time assess organ | poor quality image | Auris
| Market | Surgeons use 2D images to “envision a 3D model” in their mind when planning their approach. More experienced surgeons are more likely to be confident in this ability. However, even experienced surgeons acknowledge inefficiency and communication challenges with the surgical team when developing a surgical plan, e.g. reliance on the radiologist to develop a surgical plan. | fail to develop surgical plan independently | poor quality image | Auris
| Market | In addition to 3D models, 2D visualization, with color, of a 3D model (Anatomical Atlas) is also a valuable tool for clinicians as it is “what I’m used to and know how to interpret, but can see things better”. | inappropriate color code | use of device issue | Auris
| Market | The patient simulation (Clip Applying) tool provides significant value when planning painstaking, tissue-sparing procedures. For many, the patient simulation tool is considered a “breakthrough” because of high-quality image and instant volume calculations that allow for simulation of different surgical plans pre-operatively (e.g. segmentectomies vs a lobectomy). | fail to provide real time organ volume calculation | poor quality image | Auris
| Market | In addition to improving efficiency of pre-op planning with other clinicians, 3D modeling and patient simulation are seen as valuable tools for communicating the procedure to a patient or a trainee. | inefficacy pre-op planning | application program issue | Auris
| Market | In addition to improving efficiency of pre-op planning with other clinicians, 3D modeling and patient simulation are seen as valuable tools for communicating the procedure to a patient or a trainee. | unclear simulation | poor quality image | Auris
| Market | Surgeons see the flexibility in access across devices and locations (at home, at the office, on the go, in the OR) as a critical attribute of the solution | fail to access across devices and locations | human-device interface issue | Auris
| Market | All surgeons for the target specialties (thoracic, HBP, colorectal and urology) are procedurally required to review pre-operative patient images before the procedure. However, only a subset of surgeons refers back to pre-operative patient images intra-operatively, mostly due to an unexpected finding or uncertainty. Hence, the primary use case for pre-operative planning is before the procedure begins; intraoperative reference of the preoperative planning is secondary. | missing review of pre-operative patient images | use of device issue | Auris
| Market | The primary clinical goals of this project are to design, develop and deploy a solution that will enable Digital Solution to ● launch surgical decision support, especially medical imaging software application products | incorrect surgical decision | application program issue | Auris
| Market | enable surgeons to order, manage, and review patient-specific 2D and 3D anatomical models | fail to manage 3D model | application program issue | Auris
| Market | enable surgeons to view pre-operative images and 2D and 3D anatomical models online | fail to manage 2D images online | application program issue | Auris
| Market | enable surgeons to perform a simulated resection by applying virtual clips to approximated segments in order to simulate post-operative organ volume | fail to provide real time organ volume calculation | application program issue | Auris
| Market | enable surgeons to save a viewing and simulation session and continue at a later time | fail to continue the paused session | application program issue | Auris
| Market | enable surgeons to continue their viewing and simulation session in the OR using the Digital Hub hardware device after the commercial release of the Digital Hub (viewing of data sets on the Digital Hub will be outlined in a future Charter), | insufficient memory size | data back-up problem | Auris
| Market | enable surgeons to continue viewing their images and simulation in the OR on the web portal | fail to access through web portal | human-device interface issue | Auris
| Market | Integrated solution should be comprised of the following major components: surgeon portal, digital solutions viewer, digital surgery enablement platform. | fail to access through web portal | application program issue | Auris
| Market | Surgeon Portal - a web portal for surgeons to log into, select the appropriate language, order the modeling service, review the status of the modeling service, upload DICOM images, start the Digital Solutions Viewer. | insecure login | unauthorized access to computer system | Auris
| Market | Digital Solutions Viewer - a DICOM viewer that allows for interactions with the anatomical model, allow for simulation of resection by placing a virtual clip. | fail to provide simulation resection | application program issue | Auris
| Market | Digital Solutions Viewer - a DICOM viewer that allows for viewing a DICOM image and/or the 3D anatomical model selected by the surgeon through the surgeon portal. | fail to view images through web portal | application program issue | Auris
| Market | Digital Solutions Viewer - a DICOM viewer allows for placement of annotations and measurements, show the anatomical model, allow for interactions with the anatomical model | fail to place measurements | application program issue | Auris
| Market | Digital Surgery Enablement Platform (DSEP) - a cloud infrastructure that enables user management, data storage | fail to manage users | application program issue | Auris
| Market | Digital Surgery Enablement Platform (DSEP) - a cloud infrastructure that enables data storage, and provides DICOM studies to the external 3D model creation service as well as to the Digital Solutions Viewer. | fail to provide external 3D model creation | application program issue | Auris
| Market | Bandit Phase 2 will deliver an integrated experiences for the customer, which can be categorized in three groups: ● 3D Model Submission and Delivery ● Digital Solution Viewer ● Patient Modeling and Outcome Simulation | fail to deliver 3D model | application program issue | Auris
| Market | AI-enhanced visualization and personalized insights to help surgeons get a clearer roadmap for surgery | incorrect surgical decision | poor quality image | Auris
| Market | 3D model submission and delivery: 3D model order submission by clinical user; 3D model order management by order admin; 3D model order billing; PACS integration | inappropriate user | application program issue | Auris
| Market | Image viewer: image viewer; meaurement tools | lack of measurement tool | application program issue | Auris
| Market | patient modeling and outcome simulation: integrated 3D models; patient specific simulation of resection | fail to provide simulation resection | application program issue | Auris
| Market | Target for the Digital Solutions Viewer is the EU. The Digital Solutions Viewer will be the first CE marked product and first product under the new EU Medical Device Regulation (MDR) for Auris. | fail to comply with CE Mark | compliance issue | Auris
| Market | The Digital Solutions Viewer is considered software as a medical device (SaMD) and will need to conform to the requirements of the Medical Device Regulation (EU) 2017/745. The Viewer will need to provide Technical Documentation (TD) covering the design, development, and state of the art with respect to safety and performance and demonstrate compliance to the General safety and Performance requirements in Annex I. | IEC 62304 violated | computer software issue | Auris
| Market | The proposed classification of the Digital Solutions Viewer is a Class IIa Rule 11 (Software intended to provide information which is used to make decisions with diagnosis or therapeutic purposes). | inappropriate medical device classification | compliance issue | Auris
| Market | As the Digital Solutions viewer will be the first MDR product for Auris, the Auris QMS must be MDR ready to support certification. The proposed certificate scope will be for image visualization and planning software. | inappropriate certificate scope | compliance issue | Auris
| Market | The MDR stipulates that the manufacturer must provide sufficient clinical evidence (Article 61) to justify the level of clinical evidence necessary to demonstrate conformity with the relevant GSPR. | insufficient clinical evidence | data issue | Auris
| Market | The MDR provides a possibility to use clinical data related to a marketed equivalent device (found in the clinical evaluation required for a device) under conformity assessment. | lack of clinical data from equivalent device | data issue | Auris
| Market | Visible Patient Planner will be used as a comparator for the Digital Solutions Viewer as the Visible Patient Planner is CE marked and the technical characteristics, functional principle of software algorithms and clinical performance and intended purposes of the software are intended to be similar for the Auris Viewer. | inconsistent algorithm | programming issue | Auris
| Market | The equivalence requirement to demonstrate sufficient levels of access to comparator data should be met as Visible Patient is collaborating with J&J and the Auris Viewer is intended to interface with the Visible Patient Lab 3D recreated models. | fail to access across devices and locations | human-device interface issue | Auris
| Market | Primary User Targets: Surgeons in Thoracic, Urology, Colorectal, and Hepato-biliary | inappropriate user | user used incorrect product for intended use | Auris
| Market | Secondary User Targets (secondary expansion within primary user hospitals): Lower volume surgeons in Primary User Target specialties, Surgeons in GYN | inappropriate user | user used incorrect product for intended use | Auris
| Market | Builds Loyal Digital Surgery Customer Base: Increases the value of the online Ethicon Digital Solutions video, image & case library, creating network effects & switching costs | lack of case library | inadequate instructions for healthcare professional | Auris
| Market | Provides Critical Data Pipeline: Adds pre-op image and 3D model data to feed the Ethicon AI Insights platform | incompatible accessory | component or accessory incompatibility | Auris
| Market | Pre-operative 3D planning software is in-market today, but most focus on orthopaedics and neurosurgery. Bandit has three main competitors in the US (Fujifilm Synapse, TeraRecon iNtuition, and Intuitive Surgical IRIS™), but only Fujifilm and TeraRecon have a significant presence in EMEA, Bandit’s initial regional foothold market. Only Fujifilm is used today in APAC. | inadequate pre-op planning | application program issue | Auris
| Market | Fujifilm and TeraRecon require radiologist creation of 3D models and involvement with interpretation and dictation, whereas Bandit eliminates this step, giving surgeons a streamlined pre-op planning sequence that puts the control in their hands on their schedule. | fail to develop surgical plan independently | application program issue | Auris
| Market | Intuitive IRIS™ is currently cleared for use with CT only, whereas the others, including Bandit, are applicable for MRI as well. | fail to use 2D MRI images to envision a 3D model | application program issue | Auris
| Market | The major project milestones and estimated timing are listed below. This schedule assumes a fully staffed development team and approved program budget by the end of Phase 1. An updated schedule and list of major milestones along with program risks will be provided in the Design and Development Plan (DDP) prior to exiting the Planning and Definition Phase. | fail to identify risk | market issue | Auris
| Market | The following assumptions were made as basis of this Project Authorization Charter: ● CE Mark is prioritized (EMEA launch) and 510K Approval is performed afterwards | inappropriate authorization lists | market issue | Auris
| Market | Expedited Review (MDR) has minimal questions and no major revisions needed | prolonged Eu MDR | compliance issue | Auris
| Market | FDA acceptance of substantial equivalence with minimal questions | prolonged FDR approval | compliance issue | Auris
| Market | Code the complete the Digital Solutions Viewer is solid with few bugs/issues by mid-March 2021 | major bug existed | computer software issue | Auris
| Market | Web Portal & DSEP are not part of MDR submission | MDR submission delayed | compliance issue | Auris
| Market | QMS audit can completed by regulatory body by end of May (earlier than planned) | audit observations | compliance issue | Auris
| Market | QMS documents can be reviewed and released in Agile within 2 weeks (documents complete to released) | inadequate QMS document lists | compliance issue | Auris
| Clinical_Workflow | The intended use of the device is an aid in pre-surgical planning for procedures requiring spatial understanding of complex patient anatomy. | fail to provide spatial relationship | application program issue | Auris
| Clinical_Workflow | The medical device can be used to help the surgeon understand size and location of tumors, abnormalities, and/or defects, and create the surgical plan. | inadequate surgical plan | improper or incorrect procedure or method | Auris
| Clinical_Workflow | The medical device can be used to help the surgeon understand size and location of tumors, abnormalities, and/or defects, and create the surgical plan. | fail to detect tumor defects | application program issue | Auris
| Clinical_Workflow | The medical device can be used to help the surgeon understand size and location of tumors, abnormalities, and/or defects, and create the surgical plan. | fail to detect tumor abnormalities | application program issue | Auris
| Clinical_Workflow | Patients undergoing procedures in the general, bariatric, HPN, gynecology, thoracic, and urology specialties, who require medical images to be taken to plan for their surgical procedure. | poor medical image | Poor Quality Image | Auris
| Clinical_Workflow | If the medical device fails while being used to reference medical images or 3D models during a procedure, the user may need to access the patient images on another PACS viewer to proceed. | lack of backup viewer | fail-safe mechanism issue | Auris
| Clinical_Workflow | Measurements are taken by the medical device. The medical device can take linear measurements between two points displayed on the DICOM or 3D model | inappropriate measurement algorithm | measurement system incompatibility | Auris
| Clinical_Workflow | Measurements are taken by the medical device. The medical device can calculate a designated area shown on a 2D DICOM slice | fail to calculate slice area | incorrect software programming calculations | Auris
| Clinical_Workflow | Measurements are taken by the medical device. The medical device can through the clip apply feature, calculate volume of a designated portion of an organ to be resected | fail to calculate volume of resected organ | incorrect software programming calculations | Auris
| Clinical_Workflow | Measurements taken on a 2D device should be accurate for images acquired with high resolution. | low resolution image | Image Resolution Poor | Auris
| Clinical_Workflow | The 3D measurements and/or volumes should be considered estimates, as the anatomic models are reconstructions based on the segmentation algorithm and spacing of the 2D slices. | fail to reconstruct 3D model | application program issue | Auris
| Clinical_Workflow | The 3D measurements and/or volumes should be considered estimates, as the anatomic models are reconstructions based on the segmentation algorithm and spacing of the 2D slices. | inappropriate segnmentation alogorithm applied | application program issue | Auris
| Clinical_Workflow | The 3D measurements and/or volumes should be considered estimates, as the anatomic models are reconstructions based on the segmentation algorithm and spacing of the 2D slices. | incorrect volume calculation | parameter calculation error due to software problem | Auris
| Clinical_Workflow | The medical device may be used with a radiology monitor | fail to compatible with radiology monitor | Incompatibility Problem | Auris
| Clinical_Workflow | The medical device contains software. The software is intended to be installed by the client rep, user, or IT specialist. Periodic updates may be required to install latest version of the software. | software installation problem | problem with software installation | Auris
| Clinical_Workflow | The medical device contains software. The software is intended to be installed by the client rep, user, or IT specialist. Periodic updates may be required to install latest version of the software. | obsolete software used | application program version or upgrade problem | Auris
| Clinical_Workflow | The lifetime of the medical device may be determined by the software lifecycle and ability to maintain compatibility with common operating systems. The medical device must run on computer hardware, so it inherits that risk | fail to compatible with operating system | Incompatibility Problem | Auris
| Clinical | The lifetime of the medical device may be determined by the software lifecycle and ability to maintain compatibility with common operating systems. The medical device must run on computer hardware, so it inherits that risk | short software life cycle | computer software issue | Auris
| Clinical_Workflow | When using a tablet to interface with the device during a procedure, there is a risk that the procedure would coincide with battery depletion. | premature battery depletion | Premature Discharge of Battery | Auris
| Clinical_Workflow | successful application of the medical device is critically dependent on human factors such as the user interface. The user interface is fundamental to interacting with and displaying measurements of the patient images and anatomic models. | inappropriate interface design | inadequate user interface | Auris
| Clinical_Workflow | use errors can arise from not understanding or properly navigating the user interface. Symbols for the user interface have been based on those common in other PACS viewer products, but may be arranged differently or applied slightly differently for this medical device. | unclear symbol | device markings issue | Auris
| Clinical_Workflow | In general, the medical device is used in the office environment for planning prior to a case. When using the medical device in the operating room, the user may be more distracted and could make use errors such as loading the incorrect patient’s data, or ignoring the orientation indicators. | improper or incorrect use of device | improper or incorrect procedure or method | Auris
| Clinical_Workflow | The medical device displays information. The display relies on text, colors, sizing, etc to relay spatial information and distinguish tissue types. These can be adjusted for accessibility but must take colorblind users into account | fail to consider colorblind user | incorrect display | Auris
| Clinical_Workflow | the medical device must be used with a display large enough and with high enough resolution so that the user can clearly see details of the patient images and anatomic models. | low resolution | Image Resolution Poor | Auris
| Clinical_Workflow | The device must render the patient models taking all these factors into account. | inappropriate model rendered | Poor Quality Image | Auris
| Clinical_Workflow | the medical device is controlled by a menu. The device requires simple menu navigation for loading the correct patient information. Otherwise, there is not a specific sequence required and the user will access the different functions as desired. Most settings are accessible at all times on a sidebar. | difficult to navigate functions | inadequate user interface | Auris
| Clinical_Workflow | the medical device is controlled by a menu. The device requires simple menu navigation for loading the correct patient information. Otherwise, there is not a specific sequence required and the user will access the different functions as desired. Most settings are accessible at all times on a sidebar. | difficult to access settings | inadequate user interface | Auris
| Clinical_Workflow | User shall launch program by openning Aurora viewer during the selecting an order phase | fail to launch program | application program issue | Auris
| Clinical_Workflow | User shall launch program by navigation to "OVERVIEW" tab which is shown at the top of the screen during the selecting an order phase | fail to navigate tab | human-device interface issue | Auris
| Clinical_Workflow | User shall launch program by navigation to "OVERVIEW" tab which is shown at the top of the screen during the selecting an order phase | fail to select order | human-device interface issue | Auris
| Clinical_Workflow | User shall find patient(s) of interest by searching for patient by name, MRN, surgeon during the selecting an order phase | fail to search patient name | application program issue | Auris
| Clinical_Workflow | User shall find patient(s) of interest through sorting lists by specialty, patient, surgeon, delivery date, or status to find patient of interest during the selecting an order phase | fail to sort patient interest list | application program issue | Auris
| Clinical_Workflow | User shall find patient(s) of interest through filters list using "FILTER BY" tool in order to find patient during the selecting an order phase | fail to filter patient interest list | application program issue | Auris
| Clinical_Workflow | User shall select patient and launch image viewer by touching/clicking on row and row will highlight when selected during the selecting an order phase | unresponsive touch | failure to deploy | Auris
| Clinical_Workflow | User shall select patient and launch image viewer by clicking "NEXT" to load patient images and 3D model during the selecting an order phase | unresponsive click | failure to deploy | Auris
| Clinical_Workflow | User shall verify images are correct by double checking patient informaiton during the selecting an order phase | fail to verify image | application program issue | Auris
| Clinical_Workflow | User shall view 2D image stacks by monitoring loading of the 2D images on the "Viewing" panel during the load and view desired image stacks phase | fail to monitor loading images | application program issue | Auris
| Clinical_Workflow | After the images are loaded, the option to toggle "2D Source Image" appears. User toggles "2D Source Image" to view the 2D image stacks. | fail to view image stack | application program issue | Auris
| Clinical_Workflow | User shall select which of the following views they would like to see, by using the checkboxes that now appear under "2D Source Image" menu: Axial, Sagittal, Coronal, Anatomical Atlas during setting screen layout phase | inadequate image types | application program issue | Auris
| Clinical_Workflow | User shall select whether they would like the segmented anatomic regions to be displayed by infill or contour when turning on and off segmentation layers | fail to display anatomic contour | application program issue | Auris
| Clinical_Workflow | User shall adjust view of desired image stacks by operating pan, rotate, zoom, and scroll to select stack they would like to manipulate during the adjust view of desired image stacks phase | unable to adjust view of image stack | application program issue | Auris
| Clinical_Workflow | User shall use touchscreen, mouse, or side menu to select "Pan, Rotate, Zoom or Scroll" | insensitive touch screen | decreased sensitivity | Auris
| Clinical_Workflow | User shall undo a single adjustment with the "Undo" button during the adjust view of desired image stacks phase | fail to undo an adjustment | application program issue | Auris
| Clinical_Workflow | User shall reset the view altogether with the "Re-Center" button during the adjust view of desired image stacks phase | fail to reset view | application program issue | Auris
| Clinical_Workflow | User shall navigate to the Windowing menu on the sidebar to adjust the contrast or brightness of the image during the adjust view of desired image stacks phase | fail to adjust image contrast | application program issue | Auris
| Clinical_Workflow | User shall select from presets to auto-adjust window width and level based on target anatomy the image was acquired to show during the adjust view of desired image stacks phase | inappropriate window width | application program issue | Auris
| Clinical_Workflow | User shall choose to select "Width/Level" button to make customize adjustments to the brightness and contrast during the adjust view of desired image stacks phase | fail to adjust image brightness | application program issue | Auris
| Clinical_Workflow | User shall drag slider(s) to desire width and level percentages when adjust windowing during the adjust view of desired image stacks phase | fail to drag slider | human-device interface issue | Auris
| Clinical_Workflow | User shall reset the view using the "Reset" button when adjust windowing during the adjust view of desired image stacks phase | fail to reset view | application program issue | Auris
| Clinical_Workflow | User shall select the annotation menu by navigate to tools from the sidebar during the annotation tool phase | inappropriate annotation tool | human-device interface issue | Auris
| Clinical_Workflow | User shall select "distance" tool to measure the straight-line distance between two points on the image during the annotation tool phase | incorrect distance measurement | incorrect measurement | Auris
| Clinical_Workflow | User shall click on the image where measured distance should start, and then clicks again where measured distance should end. Linear distance will appear onscreen during the annotation tool phase | fail to pick measurement point | human-device interface issue | Auris
| Clinical_Workflow | User shall drag the endpoints to resize the line, or drag the line itself to reposition the line on the image by using distance tool during the annotation tool phase | fail to drag and drop line | human-device interface issue | Auris
| Clinical_Workflow | User shall drag the endpoints to resize the line, or drag the line itself to reposition the line on the image by using distance tool during the annotation tool phase | fail to resize line | human-device interface issue | Auris
| Clinical_Workflow | User shall delete an added line by clicking the trash icon next to the line name in the Annotation sidebar menu by using distance tool during the annotation tool phase | fail to add line | application program issue | Auris
| Clinical_Workflow | User shall add and delete arrows in a similar fashion after clicking the "Arrow" menu option. Arrows added to the screen will also appear in the Annotation sidebar menu list by arrow tool during the annotation tool phase | fail to delete arrow | human-device interface issue | Auris
| Clinical_Workflow | User shall add a note by selecting the "note" annotation option and clicking anywhere on the image during the annotation tool phase | fail to add notes at anywhere | application program issue | Auris
| Clinical_Workflow | when a textbox appears, user may type in any notes about the case or image. The note will also appear in the Annotation sidebar menu during the annotation tool phase | fail to add any notes | application program issue | Auris
| Clinical_Workflow | User shall return to a previously added note by selecting it from the Annotation menu, and clicking the "Edit" button on the textbox during the annotation tool phase | fail to edit notes | application program issue | Auris
| Clinical_Workflow | User shall remove any added note using the trash icon in the Annotation menu or on the note itself during the annotation tool phase | fail to remove any notes | application program issue | Auris
| Clinical_Workflow | User shall select any of the annotation distances or arrows from the Annotation menu list and rename them by rename annotation during the annotation tool phase | fail to rename arrow | human-device interface issue | Auris
| Clinical_Workflow | User shall select any of the annotation distances or arrows from the Annotation menu list and rename them by rename annotation during the annotation tool phase | fail to rename distance | human-device interface issue | Auris
| Clinical_Workflow | User shall reset all annotations using the "Reset" button during the annotation tool phase | fail to reset annotations | application program issue | Auris
| Clinical_Workflow | User shall toggle off "2D Source Image" and toggle on "3D Model" in the Viewing menu when openning 3D Model View during the visualizing and manipulating 3D model phase | fail to toggle on 3D model | application program issue | Auris
| Clinical_Workflow | User shall toggle off "2D Source Image" and toggle on "3D Model" in the Viewing menu when openning 3D Model View during the visualizing and manipulating 3D model phase | fail to toggle off 2D images | application program issue | Auris
| Clinical_Workflow | User shall confirm they are viewing the correct 3D model by checking the information shown onscreen. They toggle on "Patient Info" if not already on when openning 3D Model View during the visualizing and manipulating 3D model phase | incorrect 3D model | application program issue | Auris
| Clinical_Workflow | User shall select "3D Model" menu to see which anatomic structures are available when selecting 3D models to display during the visualizing and manipulating 3D model phase | incorrect 3D model | application program issue | Auris
| Clinical_Workflow | User shall select/deselect the checkboxes next to the structure they want to see, or "ALL" if they would like to see all that are available when selecting 3D models to display during the visualizing and manipulating 3D model phase | fail to deselect 3D model | application program issue | Auris
| Clinical_Workflow | User shall scan the volumes of the different models to check against expectations when selecting 3D models to display during the visualizing and manipulating 3D model phase | fail to scan model volume | application program issue | Auris
| Clinical_Workflow | User shall select "Pan, Rotate, Zoom" and/or uses the touchscreen and mouse controls to achieve the view they desire that lets them see the relevant anatomical spatial relationships when setting desired orientation during the visualizing and manipulating 3D model phase | inappropriate model orientation | image orientation incorrect | Auris
| Clinical_Workflow | Alternatively, user shall select from anatomical plane icons or cube at the bottom right of the viewing window to auto-orient the model to the desired view when setting desired orientation during the visualizing and manipulating 3D model phase | fail to auto-orient model | image orientation incorrect | Auris
| Clinical_Workflow | User shall enter fullscreen at times for a better view of the model when setting desired orientation during the visualizing and manipulating 3D model phase | fail to enter fullscreen | application program issue | Auris
| Clinical_Workflow | User shall select an individual structure when adjusting visibility of individual structures during the visualizing and manipulating 3D model phase | fail to adjust structure visibility | optical issue | Auris
| Clinical_Workflow | User shall adjust the opacity of that structure using the slider (unknown if color can be adjusted in current iteration) when adjusting visibility of individual structures during the visualizing and manipulating 3D model phase | fail to adjust structure opacity | optical issue | Auris
| Clinical_Workflow | User shall select "Done" button to complete the adjustment when adjusting visibility of individual structures during the visualizing and manipulating 3D model phase | fail to done an adjustment | failure to deploy | Auris
| Clinical_Workflow | User shall select "Clip Apply Model" menu and allows module to load when navigating to Clip Apply feature during applying clips to simulate resection volume phase | fail to apply clips to simulate resection volume | human-device interface issue | Auris
| Clinical_Workflow | User shall manipulate the model as necessary to establish the correct orientation to apply a clip when adding and removing Clips during applying clips to simulate resection volume phase | fail to remove clips | application program issue | Auris
| Clinical_Workflow | User shall manipulate the model as necessary to establish the correct orientation to apply a clip when adding and removing Clips during applying clips to simulate resection volume phase | inappropriate model orientation | image orientation incorrect | Auris
| Clinical_Workflow | User shall right click on the model the apply the clip when adding and removing Clips during applying clips to simulate resection volume phase | wrong mouse click | use of device issue | Auris
| Clinical_Workflow | User shall observe the resected organ volume and healthy tissue volume. If they are happy with the clip placement they may move on to place another clip when adding and removing Clips during applying clips to simulate resection volume phase | fail to place next clip | application program issue | Auris
| Clinical_Workflow | If the user is unhappy with the clip placement they chose, they shall either remove a single clip, or remove all clips when adding and removing Clips during applying clips to simulate resection volume phase | fail to remove clips | application program issue | Auris
| Clinical_Workflow | At any time, the user shall select any of the clips in the menu in order to review the volume simulation details for that clip when adding and removing Clips when saving the plan during applying clips to simulate resection volume phase | fail to apply clips to simulate resection volume | application program issue | Auris
| Clinical_Workflow | When clicking to a different module, user will first be prompted whether to save the plan. They may choose to save or not using the two buttons when saving the plan during applying clips to simulate resection volume phase | fail to save surgical plan | data back-up problem | Auris
| Clinical_Workflow | If they choose to save the plan, user names the plans and confirms. (Need to check) when saving the plan during applying clips to simulate resection volume phase | fail to name surgical plan | data back-up problem | Auris
| Clinical_Workflow | User shall navigate to "My Views" menu when saving current view for later access during saving current view phase | fail to save current view | data back-up problem | Auris
| Clinical_Workflow | User shall select "SAVE NEW VIEW" and pop-up appears when saving current view for later access during saving current view phase | fail to save current view | data back-up problem | Auris
| Clinical_Workflow | User may select to either save and replace a current view, or save it as a new view (and enter a new name), then press "CONFIRM" when saving current view for later access during saving current view phase | fail to replace current view | data back-up problem | Auris
| Clinical_Workflow | User shall press "CONFIRM" a second time and the view is saved when saving current view for later access during saving current view phase | fail to confirm current view | application program issue | Auris
| DIR | Use of software via a non-purpose built computer: the system shall provide a method for the user to interact with Image Viewer via a non-purpose built computer. | fail to compatible with regular computer | Incompatibility Problem | Auris
| DIR | The system shall provide a method to open a pre-operative patient image that is compliant with DICOM standard | fail to open images | image display error | Auris
| DIR | The system shall provide a method to open a pre-operative patient image that is compliant with DICOM standard | inappropriate image format | data issue | Auris
| DIR | The system shall have the capability to render 2D views of the preoperative patient images. | unable to render 2D views | application program issue | Auris
| DIR | The system shall have the capability to render 3D views of standard 3D reconstructions from an MRI/CT study. | unable to render 3D views | application program issue | Auris
| DIR | The system shall have the capability to render a view of 3D model segmentations | unable to render 3D views | application program issue | Auris
| DIR | The system shall have the capability to render a colored 2D overlay of a 3D segmentation on top of an MPR view | unable to render 2D overlay on 3D segmentation | application program issue | Auris
| DIR | The system shall provide a method to zoom in and out on the views. | unable to zoom in views | application program issue | Auris
| DIR | The system shall provide a method to pan a viewing panel. | unable to pan viewing panel | application program issue | Auris
| DIR | The system shall provide a method to rotate a 2D viewing panel. | unable to rotate 2D viewing panel | application program issue | Auris
| DIR | The system shall provide a method to rotate a 3D viewing panel. | unable to rotate 3D viewing panel | application program issue | Auris
| DIR | The system shall provide a method to toggle visibility of pre-operative image data viewed in the 3D panel. | fail to toggle image visibility | application program issue | Auris
| DIR | The system shall provide a method to rotate the cut plane within an MPR view. | unable to rotate cut plane | application program issue | Auris
| DIR | The system shall provide a method to undo and redo user interactions. | fail to undo user interaction | use of device issue | Auris
| DIR | The system shall provide a method to undo and redo user interactions. | fail to redo user interaction | use of device issue | Auris
| DIR | The system shall provide a method for resizing a viewing panel by maximizing, minimizing, or restoring to default size. | unable to resize viewing panel | application program issue | Auris
| DIR | The system shall provide a method for hiding a viewing panel. | unable to hide viewing panel | application program issue | Auris
| DIR | The system shall provide a method to adjust the Window Center and Window Width adjustments using presets. | unable to adjust window center | application program issue | Auris
| DIR | The system shall provide a method for manually adjusting window width and center values for the intensity range | unable to adjust window width | application program issue | Auris
| DIR | The system shall provide a method for a reset all viewing panels simultaneously | unable to reset viewing panel | application program issue | Auris
| DIR | The system shall provide a method to select the layout for the viewing panels | unable to select layout for viewing panel | application program issue | Auris
| DIR | The system shall provide a method to advance the multi-planer resection (MPR) cut planes through the image volume. | unable to advance MPR cut plane | application program issue | Auris
| DIR | The system shall be able to display metadata contained in the DICOM data. | fail to display DICOM data | image display error | Auris
| DIR | The system shall list all objects contained in the Visible Patient model container (3D model) and allow for user selection of objects. | fail to list 3D models | application program issue | Auris
| DIR | The system shall display the volumetric sizing information provided by Visible Patient for each object (segmentation) which is part of the 3D model. | fail to display volumetric sizing information | no display or display failure | Auris
| DIR | The system shall be able to display metadata contained in the 3D model container | fail to display metadata | no display or display failure | Auris
| DIR | The system shall display a semi-transparent overlay of the 3D model in the 2D viewing panels | fail to display semi-transparent overlay | application program issue | Auris
| DIR | The system shall use SI units and their derivatives | inappropriate measurement unit | incorrect measurement | Auris
| DIR | The system shall provide a method for performing one or more measurements within one of the viewing panels | unable to perform multiple measurements | application program issue | Auris
| DIR | The system shall label and list measurements automatically | fail to label measurement automatically | application program issue | Auris
| DIR | The system shall provide the ability to place an annotation with its associated marker on a displayed image or 3D object. | fail to place annotation on 3D model | application program issue | Auris
| DIR | The system shall provide the ability to edit the placement and content of annotations placed on a viewing panel. | fail to edit notes | application program issue | Auris
| DIR | The system shall allow for simulation of resection by applying a clip to approximated segments. | fail to simulate resection | application program issue | Auris
| DIR | The system shall display the volumetric sizing information provided by Visible Patient for each simulated clip. | fail to display volumetric sizing information | no display or display failure | Auris
| DIR | The system shall provide the ability to support individual language packs. | inadequate language pack | inadequate user interface | Auris
| DIR | The system display shall provide the ability to display U/I's, messages, and other language specific content in more than one language. | inadequate language pack | inadequate user interface | Auris
| DIR | The system shall, when appropriate and available, use standard internationally recognized symbols. | unstandardized symbol | device markings issue | Auris
| DIR | The system shall, when appropriate and available, use standard internationally recognized icons. | unstandardized icon | device markings issue | Auris
| DIR | The System shall be capable of recording all preoperative, intraoperative, and postoperative functions performed by the System. | fail to record operative data | data back-up problem | Auris
| DIR | The System shall record all functions performed by the System pre, intra, and post operatively. | fail to record operative data | loss of data | Auris
| DIR | The System should record all functions performed by the System that are relevant to monitoring, diagnosis and repair of the System by service personnel. | fail to record monitoring data | data back-up problem | Auris
| DIR | The System shall record system functions and operational events generated by the System. | fail to record operational events | data back-up problem | Auris
| CIR | The purpose of this document is to define Customer Input Requirements (CIR), including clinical and business requirements, for Digital Solutions Viewer that enables pre-op planning. | fail to create pre-op planning | application program issue | Auris
| CIR | The scope of this document covers the CIR for Release 1 of the Digital Solutions Viewer, which includes clinical pre-operative planning software used by the clinical user to create a pre-operative plan based on the completed 3D models and underlying pre-op patient images. | incomplete 3D model | application program issue | Auris
| CIR | All references to specific policies, procedures, or work instructions in this document are to the latest revision of that document unless otherwise noted. All references to personnel in this document are to that person or his/her designated representative unless otherwise stated. | incorrect revision | device markings issue | Auris
| CIR | 3D Model: Anatomy segmentations, related measurements, and metadata created based on pre-op patient images by the 3D model creator. | fail to create 3D Model | poor quality image | Auris
| CIR | 3D Model Order: A request submitted by a clinical user to create a 3D model, which contains patient images and order metadata, e.g. selected model type, patient diagnosis information and any additional order-specific request from the surgeon. | fail to provide patient diagnosis info | use of device issue | Auris
| CIR | Application: The system comprised of 3D model ordering software, clinical pre-op planning software and underlying infrastructure components. | software corrupted | computer software issue | Auris
| CIR | DICOM: Digital Imaging and Communications in Medicine is the standard for the communication and management of medical imaging information and related data. | image not in DICOM format | image display error | Auris
| CIR | PACS: Picture Archiving and Communication System, a system used in medical imaging to store, retrieve, distribute, analyze, and digitally process medical images. | fail to process medical imaging | application program issue | Auris
| CIR | Pre-Op Patient Images: Pre-operative patient images such as CT and MRI studies in a DICOM format. | image not in DICOM format | data issue | Auris
| CIR | Clinical User: A surgeon of any experience-level within the target specialties (e.g. THX, HPB, URO, COL) or surgeon’s assistant. This user intends to perform the following tasks: • create a 3D model order request • monitor 3D model status and address comments/problems • access/ download completed 3D models and report • develop pre-op plan • reference pre-op plan intraoperatively | fail to refer pre-op plan intraoperatively | application program issue | Auris
| CIR | 3D Model Service Provider: An internal Visible Patient representatives (order admin) that manages 3D model orders and facilitates communication between the clinical user and the 3D model creator (annotator). 3D model service provider (order admin) intends to perform the following tasks: • update 3D model order status, problem indicator and comments • download de-identified 3D model orders • upload completed 3D models | fail to upload 3D model | application program issue | Auris
| CIR | Healthcare Provider: A collective term to refer to clinical and non-clinical user associated with a hospital. | inadequate IFU | inadequate instructions for healthcare professional | Auris
| CIR | Healthcare Provider: A collective term to refer to clinical and non-clinical user associated with a hospital. | lack of video | nadequate instructions for non-healthcare professional | Auris
| CIR | Marketing and Sales Representative: Internal users responsible for marketing and selling the product and monitoring key performance indicators related to product use to improve the product, price, place and promotion strategies. | fail to monitor product use | Market Issue | Auris
| CIR | View Pre-Op Images: As a clinical user I need to manipulate visualization parameters associated with pre-op patient images to improve visualization of the anatomy and pathology of interest so that I can develop a patient-specific surgical plan and anticipate intra-op findings. | fail to manipulate visualization parameters | application program issue | Auris
| CIR | Manipulate Pre-Op Images: As a clinical user I need to take point to point measurements using a 2D view of the pre-op patient images and a 2D view of a 3D model, if the 3D model is available so that I can develop a patient-specific surgical plan and anticipate intra-operative findings. | unable to measure point to point | application program issue | Auris
| CIR | 2D Measurements: As a clinical user I need to take point to point measurements using a 2D view of the pre-op patient images and a 2D view of a 3D model, if the 3D model is available so that I can develop a patient-specific surgical plan and anticipate intra-operative findings | unable to develop surgical plan | application program issue | Auris
| CIR | View 3D Models and Associated Volumes: As a clinical user I need to view 3D models of anatomy and pathology and associated volumes so that I can develop a patient-specific surgical plan and anticipate intra-op findings. | fail to review 3D model of pathology | application program issue | Auris
| CIR | Manipulate 3D Model: As a clinical user I need to manipulate visualization parameters associated with 3D models to improve visualization of the target anatomy and pathology so that I can develop a patient-specific surgical plan and anticipate intra-op findings | fail to manipulate visualization parameters | application program issue | Auris
| CIR | Volume Calculations and Simulation of Resection: As a clinical user I need to access calculated volumes of the target anatomy and pathology so that I can simulate resection to determine the surgical approach. | fail to provide simulation resection | application program issue | Auris
| CIR | Volume Calculations and Simulation of Resection: As a clinical user I need to access calculated volumes of the target anatomy and pathology so that I can simulate resection to determine the surgical approach. | incorrect volume calculation | parameter calculation error due to software problem | Auris
| CIR | Configuring Display of Pre-op Images: As a clinical user I need to access and configure display of pre-op patient images and associated 3D models so that a clinical user can reference pre-op patient images and associated information as need arises before or during the procedure. | unable to configure image display | application program issue | Auris
| CIR | Configuring Display of Pre-op Images: As a clinical user I need to access and configure display of pre-op patient images and associated 3D models so that a clinical user can reference pre-op patient images and associated information as need arises before or during the procedure. | unable to configure 3D model display | application program issue | Auris
| CIR | Compatibility with OR Displays: As a clinical user I need the application to enable viewing of pre-op patient images and associated 3D models from existing OR displays so I can see all relevant patient information at the same time and minimizes interruptions/delays of having to reference multiple screens. | fail to display 3D model at OR monitor | incompatibility problem | Auris
| CIR | Compatibility with OR Displays: As a clinical user I need the application to enable viewing of pre-op patient images and associated 3D models from existing OR displays so I can see all relevant patient information at the same time and minimizes interruptions/delays of having to reference multiple screens. | fail to display patient image at OR monitor | incompatibility problem | Auris
| CIR | Compatibility with Sterile and Non-Sterile Environment: As a clinical user I need the application to enable control of visualization parameters for pre-op patient images and 3D models from sterile and non-sterile environment so that I can minimize interruptions to my clinical workflow and reduce reliance on other personnel to control the device. | clinical workflow interrupted | disinfection or sterilization issue at user location | Auris
| CIR | Secure access to patient data: As a healthcare provider, I need to ensure a secure and compliant access to personal and patient data, so I can meet cyber-security and data privacy requirements of the region, country and state where I practice. | patient data breached | unauthorized access to computer system | Auris
| CIR | Localization of Software: As a clinical user, I need to access the solution (clinical user-facing software) in the language I understand, so that I use the system correctly and efficiently. | lack of local language | human-device interface issue | Auris
| CIR | Indications for Use for 3D Model Viewing and Manipulation: As Company marketing representative, I need the application to support indications for use of the 3D model service provider so that I can ensure that the indication for use is in alignment with indications for 3D models and related functionality supplied by the third party. | Indications for Use misalignment | failure to align | Auris
| CIR | Countries Supported for 3D Model Viewing and Manipulation: As Company marketing representative, I need the application to be available in countries and geographic regions where 3D model service provider has obtained regulatory clearance and registration so that I can maximize customer and business value in according to the business plan/model. | fail to obtain regulatory clerance | compliance issue | Auris
| CIR | Countries Supported for 3D Model Viewing and Manipulation: As Company marketing representative, I need the application to be available in countries and geographic regions where 3D model service provider has obtained regulatory clearance and registration so that I can maximize customer and business value in according to the business plan/model. | fail to obtain registration | compliance issue | Auris
| CIR | Capture and Storage of System Data: As company marketing representative, I need the application to capture and store application configuration/settings, user entered data and logs in a long term storage in accordance with data privacy regulations and agreements, so that I can enable use of the data for product improvement, development. | fail to capture data | data issue | Auris
| CIR | Capture and Storage of System Data: As company marketing representative, I need the application to capture and store application configuration/settings, user entered data and logs in a long term storage in accordance with data privacy regulations and agreements, so that I can enable use of the data for product improvement, development. | fail to store data | data issue | Auris
| CIR | Capture and Storage of System Data: As company marketing representative, I need the application to capture and store application configuration/settings, user entered data and logs in a long term storage in accordance with data privacy regulations and agreements, so that I can enable use of the data for product improvement, development. | fail to follow privacy regulation | compliance issue | Auris
| CIR | Branding: As company marketing representative, I need the application to follow company defined branding, so that I can maximize consistency in the look and feel of company digital products that may be used by the same user base and reduce cognitive load on the users when using company digital products. | branding misaligned | compliance issue | Auris
| CIR | Demo Capabilities: As company marketing and sales representative, I need to be able to demonstrate the application to hospital stakeholders, so that I can communicate the value and expected workflow of the application effectively. | unclear workflow | data issue | Auris
| Clinical_HA | Major hepatic hemorrhage is an understood risk of performing a minimally invasive Hepatectomy procedure. Major hemorrhage of hepatic vessels during dissection or parenchymal division. | incorrect procedure | improper or incorrect procedure or method | Auris
| Clinical_HA | Liver failure following surgery: Postoperative liver failure is an understood risk of performing a minimally invasive Hepatectomy procedure. | incorrect procedure | improper or incorrect procedure or method | Auris
| Clinical_HA | Central Nervous System Event (CNS): Referring to any incidents that may cause damage to the CNS. CNS event is an understood risk of performing a laparoscopic colorectal procedure. | incorrect procedure | improper or incorrect procedure or method | Auris
| Clinical_HA | Systemic Sepsis: A complication of an infection where chemicals released into the bloodstream to fight the infection trigger inflammatory responses throughout the body. | incorrect procedure | bacterial contamination of device | Auris
| Clinical_HA | Nerve Injury: Nerve injury – leading to incontinence or sexual dysfunction. Nerve Injury is an understood risk of performing a laparoscopic urologic procedure. | incorrect procedure | improper or incorrect procedure or method | Auris
| Clinical_HA | Peripheral Neuropathy: Nerve damage (most often related to obturator nerve) cause by mechanical trauma to nerve fibers. Peripheral neuropathy is an understood risk of performing a laparoscopic urologic procedure. | incorrect procedure | improper or incorrect procedure or method | Auris
| Clinical_HA | Acute Bleeding: Acute bleeding can lead to a hematoma or pooling of blood creating the appearance of bruising if visible. Acute bleeding is an understood risk of performing a laparoscopic urologic procedure. | incorrect procedure | improper or incorrect procedure or method | Auris
| Clinical_HA | Major Vascular Injury can occur due to surgical manipulation of the pulmonary vessels. Aberrant pulmonary vasculature is well-documented in the literature. Major damage is a catastrophic complication and can lead to swift death due to blood loss. Major Vascular Injury is an understood risk of performing a minimally invasive Pulmonary Resection procedure. | incorrect procedure | improper or incorrect procedure or method | Auris
| Clinical_HA | Bronchial Injury: Bronchial Injury can occur due to surgical manipulation of the bronchus. Bronchial Injury is an understood risk of performing a minimally invasive Pulmonary Resection procedure. | incorrect procedure | improper or incorrect procedure or method | Auris
| Clinical_HA | Acute respiratory failure is the failure of pulmonary capillaries to perform air exchange, resulting in fluid buildup in the alveoli, systemic hypoxia and hypercapnia. Pulmonary edema caused by surgical manipulation can be responsible for acute respiratory failure. Acute Respiratory Failure is an understood risk of performing a minimally invasive Pulmonary Resection procedure. | incorrect procedure | improper or incorrect procedure or method | Auris
connection issue; detachment of device component
connection issue
connection issue
connection issue

## PRS
ID | Category | Design Requirement | Failure_Mode | Hazard
FAB-H-REQ-18 | Fabian cart | A cart shall be available to allow the fabian ventilator to be used for intra-facility patient transport. Note: the fabian cart is a separate product, intended for use with the Fabian HFO and other models of fabian ventilator, and has its own requirements documentation | -inappropriate transportation configuration -unable to use -fail to secure to cart | -user used incorrect product for intended use -medication error -incompatibility problem
FAB-H-REQ-19 | Cart Handles | Any cart indicated as appropriate for use with the fabian ventilator shall incorporate handles | -fail to incorporate -incorrect configuration -inappropriate interface design | -inadequate user interface -incorrect display -defective component
FAB-H-REQ-20 | Cart – means of attachment | Any cart indicated as appropriate for use with the fabian ventilator shall incorporate secure means of attachment for the ventilator. | -loss of indication -loss of connection -inappropriate protective measure | -failure to conduct -no display or display failure -component or accessory incompatibility
FAB-H-REQ-21 | Nebulizer connection | The fabian ventilator shall support connection of a third party electrically-powered nebulizer (Aerogen AeroNeb Solo). Note: the ventilator provides power for the AeroNeb, however the AeroNeb provides all control and monitoring of nebulization. | -incompatible connection -loss of connection -fail to supply power | -use of device issue -loss of power -electronic property issue
FAB-H-REQ-22 | Nebulizer shared USB port | The nebulizer shall be connected using the USB port (shared between this use, and support for USB 'thumb-drive' devices). | -fail to connect -fail to support interface -loss of connection | -device difficult to setup or prepare -compliance issue -fitting problem
FAB-H-REQ-23 | Dedicated capnometry port | The fabian ventilator shall provide a uniquely-keyed Lemo-type connector port for connection of a third party capnometry device | -connector mismatched -fail to connect -device marked incorrectly | -Disconnection -device unsafe to use in environment -accessory incompatible
FAB-H-REQ-24 | Sidestream Capnometry option | The fabian ventilator shall support the Oridion MicroPod sidestream capnometry module (153207 Micropod LEMO, Microstream EtCO2 Unit) | -fail to support accessory -loss of support -support failed | -use of device issue -application program version or upgrade problem -defective component
FAB-H-REQ-25 | Pulse Oximetry connection | The fabian ventilator shall provide a uniquely-keyed Lemo-type connector port for connection of a third party pulse oximetry device | -port mismatched -device marked incorrectly -fail to connect | -accessory incompatible -Disconnection -loss of power
FAB-H-REQ-26 | Pulse Oximeter supported device | The fabian ventilator shall support the Masimo pulse oximetry module (152406 Masimo SET, SpO2 LEMO-HPLPM Module, Set). | -inadequate pulse amplitude -loss of support -fail to deliver pulse | -use of device issue -failure to deploy -electronic property issue
FAB-H-REQ-533 | Gas cylinders | The fabian ventilator shall be labelled as compatible with medical air and oxygen cylinders that comply with applicable CGA / ISO standards. | -fail to compatible with oxygen supply -incorrect labeling -oxygen out of specification | -compliance issue -device unsafe to use in environment -component incompatible
FAB-H-REQ-5 | Gas hoses | The fabian ventilator shall be provided with gas hoses for air and oxygen complying with ISO 5359. | -failure of gas delivery system -fail to comply with standards or regulation -fail to control ventilator | -compliance issue -incompatibility problem -computer software issue
FAB-H-REQ-780 | Power cord | The fabian ventilator shall be provided with a power cord with IEC 60320 connector to the ventilator, and as applicable to the wall outlet in the country to which the ventilator is sold. | -wrong connector -connector mismatched -insufficient electrical contact | -compliance issue -device-device incompatibility -fitting problem
FAB-H-REQ-194 | Qualified breathing circuits | The fabian ventilator shall be qualified for use with specified Vyaire breathing circuit configurations: | -incorrect circuit type -fail to setup configuration -configuration confused | -calibration issue -failure to calibrate -Calibration Issue
FAB-H-REQ-195 | Vyaire Dual limb breathing circuit AH165 | AirLife dual limb heated breathing circuit, infant/pediatric; Dual limb circuit with heater wires for use with MR850 heater base, with patient wye with ISO 5356-1 connector; This circuit is not indicated for < 4 lpm. | -insulation material heated -circuit short -tubing deflated | -therapeutic or diagnostic output failure -failure to sense -sticking
FAB-H-REQ-196 | Vyaire Infant Flow Circuit AH132 | AirLife single limb heated breathing circuit for Infant Flow LP. Notes: Single limb circuit with heater wires for use with MR850 heater base and LP NCPAP patient interfaces. | -short circuit -circuit short -insulation material heated | -application program issue -component or accessory incompatibility -therapeutic or diagnostic output failure
FAB-H-REQ-197 | Fresh gas line adaptor | For use in nCPAP, DuoPAP and O2 Therapy modes, an adaptor shall be provided to connect the FG outlet to the humidifier chamber inlet. | -inadequate connection -fail to connect -incompatible connection | -detachment of device or device component -Burst -device markings issue
FAB-H-REQ-198 | Generic circuit rating | The fabian ventilator shall be qualified for use with generic breathing circuits that meet ISO 5367 and satisfy the following rated specification; The fabian ventilator shall be qualified for use with generic breathing circuits that meet ISO 5367 and satisfy the following rated specification. Note: The neonatal resistance specifications, the pediatric compliance specification, and the leak specifications, are aligned with ISO 5367:2014. The pediatric resistance specifications, and the neonatal compliance specifications, exceed the values specified in ISO 5367, and are intended to allow use of a wider range of commercially available circuit configurations. | -incorrect circuit type -lack of protection circuit -invalid input | -use of device issue -Failure of device to Self-Test -inadequate user interface
FAB-H-REQ-199 | Inspiratory Resistance | Inspiratory resistance shall be: NEO <= 0.74 mbar / L/min at 2.5 L/min; PED <= 0.40 mbar / L/min at 15 L/min | -high resistance -resistance inadequate -low resistancy | -air leak -Output Issue -excess flow or over-infusion
FAB-H-REQ-200 | Expiratory Resistance | Expiratory resistance shall be: NEO <= 0.74 mbar / L/min at 2.5 L/min; PED <= 0.40 mbar / L/min at 15 L/min | -high resistance -resistance inadequate -inappropriate flammability level | -air leak -Output Issue -physical resistance / sticking
FAB-H-REQ-201 | Compliance | Compliance shall be: NEO <= 2.0 mL / mbar at 60 mbar; PED <= 4.0 mL / mbar at 60 mbar | -inappropriate flow volume -fail to be deburred -inappropriate reuse | -air leak -fluid leak -image resolution poor
FAB-H-REQ-202 | Leak | Leak shall be: NEO <= 30 mL / min at 60 mbar; PED <= 40 mL / min at 60 mbar | -out of range -fail to clear tidal volume -reduced capacity | -air leak -fluid leak -Leak
FAB-H-REQ-53 | “Infant Flow LP” nasal CPAP interface | The fabian ventilator shall be qualified for use with the “Infant Flow LP” nasal CPAP interface and associated prongs and nasal masks (Vyaire SKU 7772011) | -inadequate flow -fail to deliver -inadequate specification | -compliance issue -Device Packaging Compromised -use of device issue
FAB-H-REQ-63 | “Infant Flow” nasal CPAP interface | The fabian ventilator shall be qualified for use with the “Infant Flow” (legacy) nasal CPAP interface and associated prongs and nasal masks (Inspiration Healthcare product) | -inadequate flow -unqualified use -fail to deliver | -Device Inoperable -compliance issue -Device Packaging Compromised
FAB-H-REQ-102 | “Medin Medijet” interface | The fabian ventilator shall be qualified for use with the “Medin Medijet” nasal CPAP interface and associated prongs and nasal masks (Medin product) | -inadequate interface specification -inappropriate interface design -unqualified use | -compliance issue -Device Inoperable -Device Packaging Compromised
FAB-H-REQ-1 | F&P Humidifier | The fabian ventilator shall be qualified for use with the Fisher and Paykel MR 850 heater base, and with the AH290 (F&P compatible) humidification chambers. | -unqualified use -inappropriate user -inadequate service | -application security issue -component or accessory incompatibility -user used incorrect product for intended use
FAB-H-REQ-105 | Wilamed Humidifier | The fabian ventilator shall be qualified for use with the Wilamed AIRcon heater base, and with Wilamed humidification chambers. | -fail to inline with humidification device -insufficient cooling -fail to deliver | -incompatibility problem -component or accessory incompatibility -use of device issue
FAB-H-REQ-141 | Mustang Humidifier | The fabian ventilator shall be qualified for use with the Vyaire Mustang heater base, and with Mustang humidification chambers. | -unqualified use -fail to deliver -fail to setup configuration | -application security issue -incompatibility problem -component or accessory incompatibility
FAB-H-REQ-27 | Proximal Flow Sensors | The fabian ventilator shall be qualified and labelled for use with the following proprietary flow sensors and cables: | -lack of labeling -incorrect labeling -inadequate labeling | -use of device issue -user used incorrect product for intended use -programming issue
FAB-H-REQ-28 | Infant flow sensor (SPU) | Hot wire flow sensor, single patient use, infant (151111 Flow Sensor, SPU, Infant, box #10) | -single use violated -sensor data breached -loss of sensor data | -insufficient flow or underinfusion -user used incorrect product for intended use -compliance issue
FAB-H-REQ-770 | Infant flow sensor (reusable) | Hot wire flow sensor, reusable, infant (151120 Flow Sensor, reusable, Infant) | -sensor damaged -wire bent -lack of safety wire | -Leak -Deployment Issue -difficult to open or close
FAB-H-REQ-29 | Paediatric flow sensor | Hot wire flow sensor, reusable, paediatric (151121 Flow Sensor, reusable, Pediatric) | -sensor damaged -wire bent -lack of safety wire | -Deployment Issue -insufficient flow or underinfusion -difficult to open or close
FAB-H-REQ-30 | Flow Sensor Cable | Hot wire flow sensor cable, re-usable (151100.01 Flow Sensor Cable, reusable) | -cable broken or damaged -wire torn -cable short | -Pressure Issue -imprecision -incorrect measurement
FAB-H-REQ-771 | Exhalation valve diaphragm holder | The fabian ventilator shall be qualified for use with a reusable exhalation valve diaphragm holder, which incorporates the ISO 5356-1 conical connection for the breathing set expiratory tube. | -fail to deliver -fail to display accuracy -fail to setup configuration | -compliance issue -component incompatible -fitting problem
FAB-H-REQ-772 | Exhalation valve diaphragm | The fabian ventilator shall be qualified for use with a reusable exhalation valve diaphragm. | -unqualified use -faulty valve system -fail to deliver | -defective component -device operates differently than expected -component or accessory incompatibility
FAB-H-REQ-2 | Closed suction system | The fabian ventilator shall be qualified for use with the AirLife Verso closed suction catheters and adapters. | -unqualified use -fail to allow air embolism -carrier fails to retaining set during activation | -use of device issue -user used incorrect product for intended use -component incompatible
FAB-H-REQ-3 | Breathing system filters | The fabian ventilator shall be qualified and labelled for use with the AirLife 001851 non-conductive bacterial/viral breathing system filter at the ventilator outlet | -inadequate labeling -incorrect labeling -lack of labeling | -use of device issue -defective component -device-device incompatibility
FAB-H-REQ-4 | iNOVent Nitric Oxide delivery system | The fabian ventilator shall be qualified for use with the iNOvent nitric oxide delivery system in all ventilation modes. | -fail to conjunct with nitro oxide -fail to display accuracy -inadequate service | -medical gas supply problem -infusion or flow problem -failure to deploy
FAB-H-REQ-774 | iNOMax Nitric Oxide delivery system | The fabian ventilator shall be qualified for use with the iNOMax nitric oxide delivery system in all ventilation modes. | -fail to conjunct with nitro oxide -inappropriate user -incorrect configuration | -medical gas supply problem -infusion or flow problem -failure to deploy
FAB-H-REQ-31 | Ventilation mode selection | The fabian HFO ventilator shall provide means for the operator to select between a range of modes for invasive and non-invasive ventilation, including nasal pressure controlled ventilation and high flow oxygen therapy. Invasive modes: IPPV (reset default), SIPPV, SIMV, SIMV+PSV, PSV, CPAP. Non-invasive modes: nCPAP, DuoPAP / S-DuoPAP, O2 Therapy, High Frequency Ventilation: HFO (not model “fabian HFO Light”) | -fail to enter mode -inadequate markings -inappropriate user | -use of device issue -medication error -improper device output
FAB-H-REQ-160 | Conventional mode ventilation breathing circuit configuration | Conventional ventilation modes shall be provided using a dual limb breathing circuit, with gas flow provided at the INSP outlet. Note 1: This covers IPPV, SIPPV, SIMV, SIMV+PSV, PSV and CPAP. Note 2: the INSP outlet incorporates an active safety valve, and over-pressure relief compliant to ISO 80601-2-12. | -incorrect ventilation mode -insufficient flow -inappropriate ventilation support | -compliance issue -use of device issue -human factors issue
FAB-H-REQ-178 | Conventional ventilation breath type | In all conventional modes, mandatory breaths shall have TCPL breath type, with a fixed inspiratory phase flow (I-Flow), and a fixed pressure limit (dependent on mode and trigger type), fixed duration (determined by the set I-Time). Note: for full specification of the mandatory breath profile see section 7.1.8.1 | -incorrect breath type -inadequate breath type -incorrect breath type setting | -programming issue -dose calculation error due to software problem -therapeutic or diagnostic output failure
FAB-H-REQ-203 | IPPV mode breath delivery | In IPPV ventilation mode, the ventilator shall provide mandatory breaths with properties determined by the ventilator controls, at a fixed rate independent of any patient spontaneous respiratory effort. Note: Depending on the configuration specified as FAB-H-REQ-645, the rate of mandatory breaths is determined either by the “Freq” control (as rate per minute) or the “E-Time” control (rate = 60s / (I-Time + E-Time)). | -fail to initiate breath -incorrect breath type -inadequate breath type | -calibration issue -excess flow or over-infusion -inaccurate flowrate
FAB-H-REQ-204 | IPPV mode frequency control | The “Freq” control shall have accuracy ± 1 bpm. Note: for all parameters related to mandatory breaths refer to section 7.1.8.1 below. For parameters related to gas composition and control of PEEP refer to section 7.1.5. | -accuracy out of specification -incorrect breath type -low accuracy | -unexpected therapeutic results -device difficult to setup or prepare -unauthorized access to computer system
FAB-H-REQ-775 | IPPV mode E-Time control | The “E-Time” control shall have accuracy such that the resulting breath rate is accurate to ± 1 bpm. Note: for all parameters related to mandatory breaths refer to section 7.1.8.1 below. For parameters related to gas composition and control of PEEP refer to section 7.1.5. | -incorrect breath type -low accuracy -incorrect breath period | -device difficult to setup or prepare -inaccurate delivery -use of device issue
FAB-H-REQ-205 | SIPPV mode breath delivery | In SIPPV ventilation mode, the ventilator shall provide mandatory breaths with properties determined by the ventilator controls at a rate determined by patient trigger events, or at the set rate, whichever is greater.Note: Depending on the configuration specified as FAB-H-REQ-645, the rate of mandatory breaths is determined either by the “Freq” control (as rate per minute) or the “E-Time” control (rate = 60s / (I-Time + E-Time)). | -fail to initiate breath -incorrect breath type -inadequate breath type | -inaccurate delivery -output below specifications -excess flow or over-infusion
FAB-H-REQ-206 | SIPPV mode breath rate | The delivered breath rate shall have accuracy ± 1 bpm. Note: for all parameters related to mandatory breaths refer to section 7.1.8.1 below. For parameters related to gas composition and control of PEEP refer to section 7.1.5. | -inappropriate rate calculation -incorrect breath period -low rate | -use of device issue -inaccurate delivery -computer software issue
FAB-H-REQ-776 | SIPPV mode E-Time control | The “E-Time” control shall have accuracy such that the resulting breath rate is accurate to ± 1 bpm. Note: for all parameters related to mandatory breaths refer to section 7.1.8.1 below. For parameters related to gas composition and control of PEEP refer to section 7.1.5. | -incorrect breath type -low accuracy -incorrect breath period | -device difficult to setup or prepare -inaccurate delivery -use of device issue
FAB-H-REQ-207 | SIMV mode breath delivery | In SIMV ventilation mode, the ventilator shall provide mandatory breaths with properties determined by the ventilator controls at the set rate, synchronized to patient trigger events if they occur within the appropriate triggering window. Note: Depending on the configuration specified as FAB-H-REQ-645, the rate of mandatory breaths is determined either by the “Freq” control (as rate per minute) or the “E-Time” control (rate = 60s / (I-Time + E-Time)). | -incorrect breath type setting -incorrect breath type -fail to initiate breath | -output below specifications -device alarm system issue -inaccurate delivery
FAB-H-REQ-208 | SIMV mode breath rate | The delivered breath rate shall have accuracy ± 1 bpm. Note: for all parameters related to mandatory breaths refer to section 7.1.8.1 below. For parameters related to gas composition and control of PEEP refer to section 7.1.5. | -inappropriate rate calculation -incorrect breath period -low rate | -use of device issue -inaccurate delivery -computer software issue
FAB-H-REQ-777 | SIMV mode E-Time control | The “E-Time” control shall have accuracy such that the resulting breath rate is accurate to ± 1 bpm. Note: for all parameters related to mandatory breaths refer to section 7.1.8.1 below. For parameters related to gas composition and control of PEEP refer to section 7.1.5. | -incorrect breath type -low accuracy -incorrect breath period | -device difficult to setup or prepare -inaccurate delivery -use of device issue
FAB-H-REQ-209 | SIMV+PSV mode breath delivery | In SIMV+PSV ventilation mode, the ventilator shall provide mandatory breaths with properties determined by the ventilator controls at the set rate, synchronized to patient trigger events where possible. Note: Depending on the configuration specified as FAB-H-REQ-645, the rate of mandatory breaths is determined either by the “Freq” control (as rate per minute) or the “E-Time” control (rate = 60s / (I-Time + E-Time)). | -incorrect breath type setting -incorrect breath type -inadequate breath type | -inaccurate delivery -output below specifications -excess flow or over-infusion
FAB-H-REQ-210 | SIMV+PSV mode patient triggered breaths | In SIMV+PSV ventilation mode, any patient trigger events that occur at greater than the set rate shall be supported by a PSV breath (defined in section 7.1.8.2 below) | -incorrect breath type -request trigger failed -component position incorrect | -inaccurate delivery -output below specifications -improper device output
FAB-H-REQ-211 | SIMV+PSV mode breath rate | The delivered breath rate (for mandatory breaths) shall have accuracy ± 1 bpm. Note: for all parameters related to mandatory breaths refer to section 7.1.8.1 below, and for parameters related to PSV breaths refer to section 7.1.8.2. For parameters related to gas composition and control of PEEP refer to section 7.1.5. | -inappropriate rate calculation -incorrect breath type -incorrect breath period | -use of device issue -improper device output -parameter calculation error due to software problem
FAB-H-REQ-778 | SIMV+PSV mode E-Time control | The “E-Time” control shall have accuracy such that the resulting breath rate is accurate to ± 1 bpm. Note: for all parameters related to mandatory breaths refer to section 7.1.8.1 below. For parameters related to gas composition and control of PEEP refer to section 7.1.5. | -incorrect breath type -low accuracy -incorrect breath period | -device difficult to setup or prepare -inaccurate delivery -use of device issue
FAB-H-REQ-32 | PSV mode breath delivery | In PSV ventilation mode, any patient trigger events that occur shall be supported by a PSV breath. Note: for all parameters related to PSV breaths refer to section 7.1.8.2. For parameters related to gas composition and control of PEEP refer to section 7.1.5. | -incorrect breath type -fail to support accessory -lack of hazard related scenario | -failure to deliver -improper device output -programming issue
FAB-H-REQ-33 | CPAP Mode | In CPAP ventilation mode, the pressure at the patient airway shall be maintained at the set pressure level. No mandatory or PSV breaths are provided. Note: for all parameters related to gas composition and control of PEEP refer to section 7.1.5. | -incorrect ventilation mode -inappropriate ventilation support -loss of pressure | -pressure issue -device difficult to setup or prepare -use of incorrect control settings
FAB-H-REQ-212 | Nasal Ventilation (nCPAP and DuoPAP Modes) | Non-invasive ventilation shall be provided using a single limb breathing circuit and a nasal CPAP variable-flow interface (e.g. Infant Flow), with gas flow to the breathing circuit provided at the FG outlet. This circuit configuration provides for respiratory support by nasal prongs or nasal mask only. Note: the FG outlet does not implement an active safety valve function, and the over-pressure relief valve is not set to the ISO 80601-2-12 limit. This higher limited pressure is required to allow for the use of relatively high flow with narrow-bore respiratory therapy tubing, and with the back pressure from a nasal CPAP generator interface. | -inadequate flow -fail to control flow -incorrect ventilation mode | -calibration issue -use of device issue -medical gas supply problem
FAB-H-REQ-281 | Nasal ventilation interface selection | A control shall be provided to allow the operator to specify the nCPAP interface (“tubeset”); Available modes: nCPAP, DuoPAP / S-DuoPAP; Range: “Infant Flow LP”, “Infant Flow”, “Medin Medijet”; Reset Default: “Infant Flow” | -fail to implement control measure -inappropriate interface design -inadequate interface specification | -insufficient flow or underinfusion -use of device issue -device cleaning issue
FAB-H-REQ-213 | Nasal ventilation leak compensation | A control shall be provided to allow the operator to specify the leak compensation mode, from following options: Leak Compensation OFF; Leak Compensation LOW (≤ 10% leak compensation flow); Leak Compensation MEDIUM (≤ 20% leak compensation flow); Leak Compensation HIGH (≤ 30% leak compensation flow) | -fail to update leak flow -fail to compensate -fail to monitor contamination | -leak -fluid leak -Leak
FAB-H-REQ-216 | Nasal ventilation interface selection indication | A visual indicator shall be provided during nCPAP or DuoPAP modes to show the selected interface. | -fail to indicate priority -fail to generate indication -invisible indication | -no display or display failure -defective item -medication error
FAB-H-REQ-534 | Exhalation valve control in single-limb ventilation | The exhalation valve control shall be set to zero (valve fully open) in single limb ventilation modes (nCPAP, DuoPAP, O2 Therapy) | -fail to zero -fail to close valve -fail to zero encoder count | -deployment issue -defective item -defective component
FAB-H-REQ-425 | Nasal CPAP mode availability | Nasal CPAP mode shall be available with a single-limb breathing circuit configuration and following nCPAP generator types: Infant Flow LP (Vyaire), Infant Flow (Inspiration Healthcare), Medin Medijet (Medin Inc) | -incorrect circuit type -incorrect breath type -incorrect breath type setting | -device operates differently than expected -use of device issue -device cleaning issue
FAB-H-REQ-426 | Nasal CPAP pressure delivery | In Nasal CPAP mode, the pressure at the patient airway shall be maintained at the set pressure level (no mandatory or PSV breaths are provided). | -inappropriate pressure level -fail to hold pressure -low pressure | -pressure issue -fluid leak -device difficult to setup or prepare
FAB-H-REQ-427 | DuoPAP mode availability | DuoPAP mode shall be available with a single-limb breathing circuit configuration and following nCPAP generator types: Infant Flow LP (Vyaire), Infant Flow (Inspiration Healthcare), Medin Medijet (Medin Inc) | -incorrect circuit type -incorrect breath type -incorrect breath type setting | -device operates differently than expected -use of device issue -inaccurate flowrate
FAB-H-REQ-428 | DuoPAP mode pressure delivery | In DuoPAP mode, mandatory pressure control breaths (bilevel pressure) shall be provided at a fixed rate independent of any patient spontaneous respiratory effort. Note: Depending on the configuration specified as FAB-H-REQ-645, the rate of mandatory breaths is determined either by the “Freq” control (as rate per minute) or the “E-Time” control (rate = 60s / (I-Time + E-Time)). | -pressure control failed -fail to adjust pressure -incorrect breath type | -failure to deploy -Pressure Issue -programming issue
FAB-H-REQ-429 | DuoPAP mode rate | The delivered breath rate shall have accuracy ± 1 bpm. | -low rate -inappropriate rate calculation -incorrect breath type | -parameter calculation error due to software problem -inaccurate delivery -medical gas supply problem
FAB-H-REQ-779 | DuoPAP mode E-Time control | The “E-Time” control shall have accuracy such that the resulting breath rate is accurate to ± 1 bpm. Note: for all parameters related to mandatory breaths refer to section 7.1.8.1 below. For parameters related to gas composition and control of PEEP refer to section 7.1.5. | -incorrect breath type -low accuracy -incorrect breath period | -device difficult to setup or prepare -inaccurate delivery -use of device issue
FAB-H-REQ-214 | Triggered DuoPAP mode | Triggered ventilation shall be available in DuoPAP mode using a pressure trigger when the interface setting is “Medin Medijet”. | -fail to maintain ventilation mode -incorrect ventilation mode -incorrect settings | -use of incorrect control settings -failure to deploy -device difficult to setup or prepare
FAB-H-REQ-215 | Triggered DuoPAP mode with flow sensor | Triggered ventilation shall be available in DuoPAP mode when a flow sensor is connected, using a flow trigger based on exhaust connector flow, when the interface setting is “Infant Flow” or “Infant Flow LP”. | -fail to maintain ventilation mode -fail to control ventilator -inappropriate ventilation support | -device operates differently than expected -fluid leak -use of incorrect control settings
FAB-H-REQ-219 | Triggered breaths in S-DuoPAP mode | When S-DuoPAP is active, the ventilator shall provide mandatory breaths with properties determined by the ventilator controls at a rate determined by patient trigger events, or at the set rate, whichever is greater. Note: breath timing is identical to SIPPV mode. | -incorrect breath type -fail to initiate breath -inadequate breath type | -inaccurate delivery -excess flow or over-infusion -device alarm system issue
FAB-H-REQ-279 | Triggered DuoPAP mode labelling on-screen | When S-DuoPAP is active, the ventilator user interface shall use “S-DuoPAP” as the label for the DuoPAP mode control and status indicators. | -fail to generate indication -fail to activate alert -lack of indication | -inadequate user interface -incorrect display -use of device issue
FAB-H-REQ-535 | O2 Therapy mode availability | O2 Therapy mode shall be available with a single-limb breathing circuit configuration. In this mode, the ventilator provides a flow of gas to nasal cannulae, there is negligible pressure provided to the patient airway. | -fail to provide sufficient pressure -inappropriate flow pattern -fail to control flow | -use of device issue -defective component -calibration issue
FAB-H-REQ-217 | O2 Therapy flow delivery | In O2 Therapy mode, a continuous flow of oxygenated air shall be provided to the nasal cannulae, at a flowrate set by the operator “Flow” control. Note: pressure related alarms are not enabled in this mode. With a high flow nasal cannula configuration, it is not usually possible to equate pressure measured in the breathing circuit to pressure at the patient airway. | -fail to control flow -loss of pressure -inappropriate flow volume | -failure to cycle -Blockage within Device or Device Component -delivery system failure
FAB-H-REQ-218 | High Flow control accuracy | The Flow control shall have accuracy ± (1 lpm + 10% of setting). Note: This mode of therapy is not intended to provide pressure support or any form of mandatory breaths. | -inappropriate control flow -low accuracy -accuracy out of specification | -incorrect measurement -calibration issue -use of device issue
FAB-H-REQ-283 | High Frequency mode not available in Light model | High Frequency Oscillation shall be available only for model “fabian HFO” and not for model “fabian HFO Light” | -inaccurate frequency -low motion resolution -impedance fluctuated | -Failure to Sterilize -intermittent infusion -application program issue
FAB-H-REQ-293 | High Frequency Oscillation availability | High Frequency Oscillation shall be supported using a dual limb breathing circuit. In this mode, the ventilator provides a constant flow of gas from the INSP outlet to the breathing circuit, while also enabling high frequency pressure delivery using the internal HFO module. | -fail to generate periodic notification -incorrect triggered pulse -fail to generate periodic request | -use of device issue -improper device output -electronic property issue
FAB-H-REQ-334 | High Frequency Oscillation mode pressure delivery | In HFO ventilation mode, the pressure at the patient airway shall be oscillating around the set mean pressure level “Pmean” with operator set frequency and pressure amplitude. No mandatory or PSV breaths are provided. | -inappropriate pressure level -low pressure -over pressure | -pressure issue -use of incorrect control settings -fluid leak
FAB-H-REQ-348 | High Frequency Oscillation mode bias flow setting | In the HFO mode, the device shall provide the possibility to route the patient gas flow (bias flow) from the FG port (as alternative to INSP port). | -inappropriate route -fail to access -fail to drive piston | -electromagnetic interference (emi) -device difficult to setup or prepare -No Known Device Problem
FAB-H-REQ-405 | High Frequency Oscillation mode “lung recruitment” maneuver | When “lung recruitment” is enabled, the ventilator shall increase the mean pressure level to a user-specified pressure level “PmeanRec”, for a user-specified amount of time “I-TimeRec” and at a specified frequency “FreqRec”, without changing the amplitude of HFO. | -inappropriate pressure level -fail to raise pressure -low pressure | -pressure issue -calibration issue -delivery system failure
FAB-H-REQ-440 | Volume Guarantee availability | Volume Guarantee shall be available in IPPV, SIPPV, SIMV, SIMV+PSV, PSV and HFO modes, when a flow sensor is connected and functioning to specification. | -inadequate specification -lack of specification -fail to address deliverable | -use of device issue -compliance issue -failure to reset
FAB-H-REQ-441 | Volume Guarantee operation (HFO mode) | When Volume Guarantee is enabled during HFO mode, the ventilator shall adjust the oscillatory pressure amplitude to maintain the expired tidal volume for each oscillation at the set target. | -fail to adjust settings -fail to adjust volume -fail to update tidal volume | -use of device issue -failure to service -device difficult to setup or prepare
FAB-H-REQ-220 | Volume Guarantee operation (conventional modes) | When Volume Guarantee is enabled in conventional dual limb ventilation modes, the ventilator shall adjust the inspiratory pressure for mandatory breaths to maintain the expired tidal volume Vte at the set target.Note: VG applies to mandatory breaths in IPPV, SIPPV, SIMV and SIMV+PSV modes, and to PSV breaths in PSV mode. It does not apply to PSV breaths in SIMV+PSV mode. | -fail to update tidal volume -fail to adjust pressure -fail to adjust volume | -output below specifications -device difficult to setup or prepare -use of device issue
FAB-H-REQ-221 | Volume guarantee volume accuracy | The V-Guarantee control shall have accuracy ± (0.5 mL + 10% setting) | -low measurement accuracy -low accuracy -accuracy out of specification | -failure to convert to back-up -poor quality image -incorrect measurement
FAB-H-REQ-222 | Volume Guarantee pressure limits | The inspiratory pressure for mandatory breaths in Volume Guarantee shall be constrained to the range (PEEP + 2 mbar) to set Pmax (with maximum of 80 mbar). Note: This is a specification for the target pressure. The actual range of pressure is subject to the tolerance for pressure delivery. | -incorrect breath type setting -inappropriate pressure level -incorrect breath type | -use of incorrect control settings -delivery system failure -pressure issue
FAB-H-REQ-442 | Volume Guarantee maximum delivered pressure | The ventilator shall act to prevent delivered pressure from exceeding “Pmax” for a period of greater than 50 ms. | -over pressure -pressure out of range -fail to prevent flow | -overheating of device or device component -delivery system failure -defective component
FAB-H-REQ-443 | Volume Guarantee test breath pressure | The first breath after activating VG (“VG test breath”) or after a setting change shall have a pressure target no greater than maximum of (Pinsp setting, Pmax setting + 3.0 mbar). Note: The test breath pressure target set by software is PEEP + 0.75 ∙ (Pinsp – PEEP), or PEEP + 2.0 mbar if greater. The actual pressure is subject to the tolerance for pressure delivery. | -inappropriate pressure setting -low pressure -fail to initiate breath | -use of incorrect control settings -pressure issue -use of device issue
FAB-H-REQ-223 | Volume Guarantee for patient-triggered and time-triggered breaths | Control of inspiratory pressure shall be maintained separately for patient-triggered and time-triggered mandatory breath | -fail to hold pressure -loss of pressure -incorrect breath type | -defective component -use of incorrect control settings -decreased sensitivity
FAB-H-REQ-224 | Volume Guarantee pressure changes | The inspiratory pressure for mandatory breaths in Volume Guarantee shall be limited to change by no more than ± 3 mbar of previous value from breath to breath (of the same type, patient-triggered or time-triggered). | -incorrect breath type -over pressure -pressure out of range | -pressure issue -improper device output -use of incorrect control settings
FAB-H-REQ-225 | Volume Guarantee disabled on flow sensor disconnect | Volume Guarantee shall be disabled in the event of a flow sensor disconnection or flow sensor fault alarm. | -fail to disable alarm -fail to disable control -fail to disable overlay | -Pressure Issue -incorrect measurement -imprecision
FAB-H-REQ-226 | Volume Guarantee autoset volume limit | During Volume Guarantee ventilation, mandatory breaths shall be terminated if the inspired volume exceeds 130% of the set target volume. Note: this volume is leak-compensated to avoid premature termination in the presence of airway leak. | -inappropriate ventilation support -incorrect breath type -low volume | -use of device issue -parameter calculation error due to software problem -output below specifications
FAB-H-REQ-227 | Volume Guarantee deactivation | On deactivation of Volume Guarantee, the set inspiratory pressure shall be returned to the value of the Pinsp control. | -fail to adjust volume -fail to adjust settings -incorrect settings | -pressure issue -use of incorrect control settings -output below specifications
FAB-H-REQ-228 | Volume Limit operation | When Volume Limit is enabled, the ventilator shall terminate inspiratory phase of a mandatory breath when the delivered volume exceeds V-Limit. | -incorrect breath type -fail to maintain ventilation mode -fail to disable ventilation | -use of device issue -delivery system failure -failure to service
FAB-H-REQ-444 | Volume Limit availability | Volume Limit shall be available in IPPV, SIPPV, SIMV, SIMV+PSV and PSV modes, when a flow sensor is connected and functioning to specification. | -fail to limit flow -inadequate flow -incorrect limit setting | -use of device issue -use of incorrect control settings -programming issue
FAB-H-REQ-229 | Volume Limit accuracy | The V-Limit control shall have accuracy ± (0.5 mL + 10% setting). | -low measurement accuracy -incorrect limit setting -fail to set limit correctly | -failure to convert to back-up -incorrect measurement -poor quality image
FAB-H-REQ-445 | Volume Limit exceeded notification | When a breath is terminated due to the V-Limit being reached a message “Volume Limited” shall be displayed. | -incorrect message display -fail to display message -loss of message | -device alarm system issue -fail-safe mechanism issue -energy output to patient tissue incorrect
FAB-H-REQ-498 | Rise Time Control Breaths | A Rise Time control shall be provided to allow capability like 'pressure controlled' ventilation, in which sufficient flow is provided to ensure a specified pressure rise time. | -fail to provide sufficient pressure -fail to control flow -inadequate flow | -failure to cycle -device cleaning issue -Difficult to Open or Close
FAB-H-REQ-230 | Rise Time breaths – flow control | When Rise Time control is enabled, the flow provided to the outlet shall be determined by the ventilator to ensure that the specified pressure rise time can be achieved. Note: In this use case, the I-Flow control is inactive. | -fail to control flow -fail to raise pressure -fail to update pressure | -failure to deploy -programming issue -component incompatible
FAB-H-REQ-231 | Rise Time breaths – exhalation valve control | When Rise-Time is enabled, the ventilator shall adjust the exhalation valve drive using a ramped increase in drive current to enable the pressure to reach the preset target within the time set by the Rise-Time control. Note: in presence of spontaneous inspiratory effort, or with a high pressure difference from PEEP to target pressure, the measured rise time may be greater than set. | -fail to adjust pressure -fail to adjust driveline pressure -fail to adjust settings | -device difficult to setup or prepare -failure to deploy -use of incorrect control settings
FAB-H-REQ-232 | Rise Time accuracy | In absence of spontaneous inspiratory effort, the Rise Time control shall have accuracy ± (0.05 s). | -low accuracy -accuracy out of specification -incorrect time remain | -optical issue -failure to align -bleed back
FAB-H-REQ-233 | Apnea ventilation in PSV mode with apnea alarm enabled | In PSV mode, apnea ventilation shall be provided at the set “Backup Rate” after the preset “Apnea Time”. | -incorrect ventilation mode -inappropriate ventilation support -inappropriate rate calculation | -use of incorrect control settings -device maintenance issue -deployment issue
FAB-H-REQ-234 | Apnea ventilation in PSV mode with apnea alarm disabled | In PSV mode, with the “Apnea Time” is set Off, backup breaths shall be triggered if the patient breath rate falls below the set “Backup Rate”. | -incorrect breath type setting -inappropriate cycle breath time -inappropriate rate calculation | -use of incorrect control settings -pressure issue -output below specifications
FAB-H-REQ-235 | Apnea ventilation in CPAP mode | In CPAP mode, when an interruption in spontaneous breathing exceeds the “Apnea Period”, from 1 to 5 backup breaths (as determined by the Backup control) shall be triggered using pre-determined settings. | -incorrect breath type -fail to adjust settings -inappropriate pressure setting | -failure to align -inaccurate delivery -pressure issue
FAB-H-REQ-236 | Apnea ventilation reset | The apnea timer shall be reset at the start of inspiration of each detected patient breath. | -fail to start timer -fail to update timer -fail to set or reset alarm automatically | -use of device issue -failure to service -protective measure issue
FAB-H-REQ-499 | PEEP control | In all modes intended for a dual limb breathing circuit configuration other than HFO and CPAP, the pressure target between breaths shall be determined by the PEEP control. | -inaccurate pressure reading -incorrect breath type -incorrect breath type setting | -pressure issue -use of incorrect control settings -decreased sensitivity
FAB-H-REQ-501 | CPAP control | In CPAP, nCPAP and DuoPAP modes the pressure target between breaths shall be determined by the CPAP control. Note: In CPAP, NCPAP and DuoPAP modes, the expiratory pressure target is determined by the CPAP control. In HFO mode the mean pressure target is determined by a Pmean control. | -inaccurate pressure reading -inappropriate pressure setting -inappropriate pressure level | -pressure issue -incorrect software programming calculations -decreased sensitivity
FAB-H-REQ-502 | Pmean control | In HFO mode the pressure target upon which high frequency oscillations are superimposed shall be determined by the Pmean control. | -inaccurate pressure reading -inappropriate pressure level -inappropriate pressure setting | -failure to deploy -device alarm system issue -incorrect software programming calculations
FAB-H-REQ-237 | PEEP / CPAP pressure accuracy | Set CPAP / PEEP shall have accuracy ± (0.5 mbar + 3% of setting). | -low measurement accuracy -accuracy out of specification -low accuracy | -incorrect measurement -Medical Gas Supply Problem -failure to infuse
FAB-H-REQ-516 | Pmean accuracy | Set Pmean shall have accuracy ± (0.5 mbar + 3% of setting). | -low measurement accuracy -rotation accuracy out of range -low accuracy | -incorrect measurement -failure to convert to back-up -failure to infuse
FAB-H-REQ-538 | Exhalation Flow | In modes IPPV, SIPPV, SIMV, SIMV+PSV and PSV, the flow of gas to the breathing circuit between breaths is determined by the E-Flow control. Note: In CPAP mode, this flow is determined by the Flow-Min control. In HFO, nCPAP, DuoPAP and O2-Therapy modes, this flow is determined by the Flow control. | -fail to control flow -fail to produce flow waveform -incorrect breath type | -excess flow or over-infusion -dose calculation error due to software problem -mechanical jam
FAB-H-REQ-238 | Exhalation flow accuracy | The set flow during exhalation shall have accuracy ± (0.5 lpm + 10% setting). | -low displacement accuracy -inaccurate flow rate -low measurement accuracy | -incorrect measurement -Inflation Issue -fluid leak
FAB-H-REQ-239 | FIO2 control | The volumetric percentage of oxygen in the gas delivered to the patient shall be determined by the O2 setting. | -fail to set limit correctly -incorrect measurement -inaccurate measurement | -inaccurate delivery -medical gas supply problem -unable to obtain readings
FAB-H-REQ-240 | Delivered FiO2 accuracy | The oxygen fraction shall have accuracy ± 3 % (provided that the oxygen sensor is correctly calibrated). | -low oxygen concentration -accuracy out of specification -low accuracy | -calibration issue -incorrect measurement -excess flow or over-infusion
FAB-H-REQ-241 | Delivered FiO2 accuracy with non-calibrated / faulty O2 sensor | The oxygen fraction shall have accuracy ± 5 % (when the oxygen sensor is not calibrated or disabled). Note: this ventilator uses a dual oxygen sensor and implements closed loop control based on the oxygen sensor reading to improve accuracy, provided that the oxygen sensor is both calibrated and not in a fault state. | -low accuracy -fail to calibrate manually -accuracy out of specification | -medical gas supply problem -use of device issue -calibration issue
FAB-H-REQ-242 | Oxygen Flush | An oxygen flush control shall be provided to enable the volume fraction of oxygen delivered to the patient to be increased by a preset amount, for a maximum period of 120 seconds. Note: The duration of oxygen flush is configurable, see FAB-H-REQ-154. | -fail to adjust volume -unable to increase -low volume | -inaccurate delivery -defective component -calibration issue
FAB-H-REQ-788 | FiO2 response time | The oxygen fraction as measured at the patient connection port shall have response time (to 90% of step change) not exceeding 15s. Note: Degradation of this response time needs to be quantified for flow rates below 8 L/min, and breath rates below 10 bpm. | -long response time -measurement range out of specification -reduced wearing time | -medical gas supply problem -parameter calculation error due to software problem -excess flow or over-infusion
FAB-H-REQ-541 | PRICO function activation | PRICO functions shall be available only if the PRICO license is activated on the device. | -loss of function -fail to turn on device -fail to activate | -instruction for use issue -failure to deploy -user used incorrect product for intended use
FAB-H-REQ-542 | PRICO availability in all ventilation modes | PRICO function shall be available in all ventilation modes: IPPV, SIPPV, SIMV, SIMV+PSV, PSV, CPAP, nCPAP, DuoPAP, O2 Therapy, HFO | -incorrect mode -fail to allow air in and out -inappropriate user | -failure to deploy -improper device output -device maintenance issue
FAB-H-REQ-543 | PRICO screen | A dedicated screen shall be provided to allow control of PRICO function, which includes at least the following elements: Oximeter pleth waveform, SpO2 monitor, Perfusion Index (PI) monitor, Pulse Rate monitor, SIQ indication (as bargraph, with low SIQ alarm threshold indicated), Controls for SpO2 target range, and FiO2 control range, PRICO on / off control | -lack of power control -loss of control -fail to re-enable control | -no device output -protective measure issue -delivered as unsterile product
FAB-H-REQ-544 | PRICO target range | PRICO shall provide the capability to the user to set the target SpO2 range (as low and high limit values) within overall limits of 80 to 100 %. | -fail to set limit correctly -inappropriate threshold setting -out of limits | -infusion or flow issue -defective item -insufficient flow or underinfusion
FAB-H-REQ-545 | PRICO control of FiO2 | PRICO shall control the FiO2 to maintain the target SpO2 to be in the centre of the set range. | -inappropriate control flow -fail to control flow -out of range | -defective item -pressure issue -device maintenance issue
FAB-H-REQ-546 | PRICO range for adjustment of FiO2 | PRICO shall provide the capability to the user to set the limits of FiO2 within which the PRICO can control, within overall limits of 21 to 100 %. | -incorrect limit setting -residual over limit -inappropriate threshold setting | -pressure issue -device difficult to setup or prepare -failure to align
FAB-H-REQ-547 | PRICO indicator | An indicator shall be provided on screen at all time that the PRICO function is active (PRICO on). | -invisible indication -fail to generate indication -loss of indication | -no display or display failure -no device output -failure to deploy
FAB-H-REQ-548 | PRICO action while SpO2 remains in target range | The PRICO function shall adjust FiO2 by an increment of -1 %, zero, or +1 %, at intervals of 30 seconds, such that FiO2 is reduced is SpO2 is greater than the mid-point of the target range, and increased if SpO2 is less than the mid-point of the target range. | -fail to adjust settings -adjustment failed -replacement time long | -unintended system motion -infusion or flow issue -device alarm system issue
FAB-H-REQ-549 | PRICO action while SpO2 falls outside target range | The PRICO function shall adjust FiO2 by an increment of between -10 to -1 % at intervals of 30 seconds while SpO2 remains above the target range, and between +1 to +10 % if SpO2 remains below the target range. Note: The exact step size is determined by the current value of SpO2, and the trend of SpO2. An adjustment is made immediately when SpO2 first falls outside the target range and then at 30s intervals. | -operation speed out of range -fail to adjust settings -adjustment failed | -device difficult to setup or prepare -mechanical issue -defective item
FAB-H-REQ-550 | Action on O2 Flush will PRICO is on | The “O2 Flush” function shall be able to be activated while PRICO is on and shall have the additional effect of turning PRICO off. Note: The operator would have to manually re-enable PRICO after “O2 Flush” has completed. This is to allow the patient SpO2 to stabilize. | -fail to activate -fail to turn on device -loss of function | -human-device interface issue -component missing -break
FAB-H-REQ-551 | PRICO termination due to fault states | PRICO shall be deactivated and the target FiO2 shall be switched to the set FiO2 (the value of the FiO2 control prior to enabling PRICO) automatically in case any of the following alarm occur: Air supply pressure,Input pressure blender,O2 sensor calibration error, Oxygen sensor defect, Oxygen sensor used up, Oxygen supply pressure, SpO2 module disconnected, SpO2 sensor failure. | -fail to reset alarm automatically -switch to backup mode incorrectly -fail to activate occlusion alarm | -device alarm system issue -break -defective item
FAB-H-REQ-552 | PRICO temporary deactivation during fault states | PRICO shall be deactivated and the target FiO2 shall be switched to the set FiO2 automatically if any of the following alarm is active. In case the alarm condition is resolved PRICO shall be re-activated automatically: Check ET tube, Clean flow sensor, Check SpO2 sensor, Flow sensor defect, Flow sensor not connected, Low SIQ, O2 value out of range, Patient disconnected, Tube Occlusion. | -fail to activate alarm -fail to reset alarm automatically -switch to backup mode incorrectly | -device alarm system issue -break -defective item
FAB-H-REQ-34 | Time-triggered breaths | Breaths shall be triggered by time in IPPV mode, and in SIPPV, SIMV, SIMV+PSV and DuoPAP modes when no patient-trigger event occurs within a time window defined by (60 / Rate) or (I-Time + E-Time). Note: Depending on the configuration specified as FAB-H-REQ-645, the rate of mandatory breaths is determined either by the “Freq” control (as rate per minute) or the “E-Time” control (rate = 60s / (I-Time + E-Time)). | -commission time long -installation time long -fail to trigger | -improper device output -failure to deliver -dose calculation error due to software problem
FAB-H-REQ-243 | Modes that support patient triggered ventilator breaths | Breaths shall be triggered by patient effort in SIPPV, SIMV, SIMV+PSV and PSV modes, and in DuoPAP mode when NIV trigger is enabled. | -fail to associate patient -inappropriate patient type -fail to trigger | -programming issue -decreased sensitivity -inaccurate delivery
FAB-H-REQ-553 | Patient-triggered breaths in conventional ventilation | The Trigger Mode control (applicable to modes SIPPV, SIMV, SIMV+PSV and PSV only) shall determine the mechanism used to detect patient spontaneous inspiratory effort, as flow trigger, volume trigger or pressure trigger as selected by the operator. Note: Flow trigger and volume trigger selections both require the flow sensor to be connected. | -incorrect detection -loss of detection -lack of change control | -use of device issue -computer software issue -excess flow or over-infusion
FAB-H-REQ-554 | NIV trigger - patient-triggered breaths in S-DuoPAP mode | When activated, the NIV Trigger function (applicable only to DuoPAP / S-DuoPAP mode) shall detect patient spontaneous inspiratory effort based on the specified tubeset, and availability of a flow sensor: Medin Medijet – pressure trigger, Infant Flow - flow trigger, Infant Flow LP – flow trigger, Note: flow trigger function requires the flow sensor to be connected. | -incorrect detection -fail to hold inspiratory -fail to detect occlusion | -failure to cycle -device operates differently than expected -operating system version or upgrade problem
FAB-H-REQ-555 | Flow Trigger | When Trigger Mode is set to Flow Trigger, the Trigger control shall determine the threshold for inspiratory flow required to initiate a patient-triggered breath. | -fail to identify threshold -fail to trigger -fail to generate indication | -failure to cycle -excess flow or over-infusion -insufficient flow or underinfusion
FAB-H-REQ-430 | Flow trigger control range | Trigger values of 1 to 10 shall represent a flow threshold of 0.12 to 1.20 lpm (neonatal) or 0.6 to 6.0 lpm (pediatric). | -inappropriate variable value -fail to flow check -incorrect value | -failure to cycle -insufficient flow or underinfusion -Deployment Issue
FAB-H-REQ-244 | Pressure Trigger | When Trigger Mode is set to Pressure Trigger, the Trigger control determines the threshold for inspiratory pressure (as a decrement below baseline pressure) required to initiate a patient-triggered breath. | -use incorrect control settings -pressure control failed -fail to trigger | -delivery system failure -failure to cycle -Increase in Pressure
FAB-H-REQ-556 | Pressure Trigger control range | Trigger values of 1 to 10 shall represent a pressure threshold of 0.12 to 1.20 mbar (neonatal) or 0.6 to 6.0 mbar (pediatric). | -inappropriate variable value -low pressure -inappropriate threshold setting | -failure to cycle -pressure issue -device unsafe to use in environment
FAB-H-REQ-557 | Volume Trigger | When Trigger Mode is set to Volume Trigger, the Trigger control determines the threshold for inspiratory volume required to initiate a patient-triggered breath. | -use incorrect control settings -incorrect limit setting -inappropriate control flow | -use of device issue -output above specifications -delivery system failure
FAB-H-REQ-431 | Volume Trigger control range | Trigger value of 1 to 10 shall represent thresholds of 5% to 15% of the measured exhaled tidal volume (Vte). Note: Trigger value of 1 represents 5% of Vte; 10 represents 15% of Vte. | -fail to deliver tidal volume -fail to clear tidal volume -fail to calculate tidal volume | -incorrect measurement -device unsafe to use in environment -Medical Gas Supply Problem
FAB-H-REQ-558 | NIV flow trigger | When NIV Trigger is enabled and operating in Flow Trigger mode (tubeset Infant Flow or Infant Flow LP; flow sensor active) a patient trigger shall be detected when estimated inspiratory flow exceeds a threshold based on the “Trigger sensitivity” control, and the delta pressure [ΔP = Pduo – CPAP] within a range 0.135 L/min (for trigger sensitivity 1, ΔP = 2.0 mbar) to 0.488 L/min (for trigger sensitivity 10, ΔP = 13.0 mbar). Notes: The flow sensor in this mode is connected to the nCPAP generator exhaust and measures total exhaust flow. During inspiration, this is equivalent to outlet flow minus patient inspiratory flow and leak flow. The exact flow threshold, and the time required for flow to exceed the threshold prior to a trigger being actioned, do not have specified accuracy. | -request trigger failed -incorrect detection -fail to activate | -Pressure Issue -improper flow or infusion -insufficient flow or underinfusion
FAB-H-REQ-559 | NIV flow trigger minimum exhalation | NIV Trigger shall be inhibited until Exhalation time has exceeded 260 ms, and Pprox < [CPAP + 0.125 ∙ ΔP + 0.9 mbar] Note: ΔP represents the difference between the inspiratory pressure and expiratory pressure, this ensures that pressure has fallen to within 12.5% of the expiratory level. | -pressure dropped -reduced wearing time -pressure out of range | -use of incorrect control settings -deployment issue -programming issue
FAB-H-REQ-560 | NIV pressure trigger | When NIV Trigger is enabled and operating in Pressure Trigger mode (tubeset Medin Medijet only) a patient trigger shall be detected when pressure drop below PEEP exceeds a threshold based on the “Trigger sensitivity” control, and the set CPAP level within a range 0.1325 mbar (for trigger sensitivity 1, CPAP = 2.0 mbar) to 0.651 mbar (for trigger sensitivity 10, CPAP = 13.0 mbar). Note: The exact pressure threshold, and the time required for pressure drop below PEEP to exceed the threshold prior to a trigger being actioned, do not have specified accuracy. | -incorrect detection -request trigger failed -detection absence failed | -delivery system failure -inaccurate delivery -programming issue
FAB-H-REQ-35 | Manual Breaths | A manual breath shall be triggered by action of the manual breath control in all modes, provided that a minimum exhalation time of 200ms has been met. | -incorrect breath type -breath delayed -fail to initiate breath | -improper device output -defective item -device cleaning issue
FAB-H-REQ-245 | Mandatory breath duration | Mandatory breaths shall have a set inspiratory duration, configured by the operator using the Inspiratory Time control. | -inappropriate cycle breath time -incorrect breath type setting -incorrect breath type | -use of incorrect control settings -delivery system failure -decreased sensitivity
FAB-H-REQ-246 | Breath duration accuracy | The inspiratory duration shall have accuracy ± (0.05 s). | -low accuracy rate -low measurement accuracy -inappropriate inspiratory time | -inaccurate delivery -component missing -defective component
FAB-H-REQ-247 | Breath termination on high pressure alarm | Mandatory breaths shall be terminated within a minimum of 200 ms on activation of the High PIP alarm. | -fail to terminate process completely -fail to disable alarm -low alarm volume | -device alarm system issue -improper device output -infusion or flow issue
FAB-H-REQ-561 | Mandatory breaths are TCPL | The default mandatory breath type for all modes intended for a dual-limb breathing circuit shall be TCPL. ote: This applies when the Rise-Time control is disabled. | -incorrect breath type -inadequate breath type -confused breath type | -improper device output -deployment issue -application program issue
FAB-H-REQ-432 | TCPL breath flow control | The ventilator shall provide a fixed flow of oxygenated air to the breathing circuit throughout the duration of the inspiratory phase of a TCPL breath at the set I-Flow. | -air flow blocked -over flow -air delivery compromised | -excess flow or over-infusion -failure to cycle -Medical Gas Supply Problem
FAB-H-REQ-433 | TCPL flow control accuracy | The set I-Flow shall have accuracy ± (0.5 lpm + 10% setting). | -low accuracy -low displacement accuracy -accuracy out of specification | -incorrect measurement -failure to convert to back-up -output below specifications
FAB-H-REQ-434 | TCPL pressure control | With Volume Guarantee disabled, the ventilator shall modulate the exhalation valve force in order to set the opening pressure for the exhalation valve to a set level (“P-Insp”). | -fail to disable ventilation -fail to adjust volume -fail to adjust pressure | -pressure issue -defective component -deployment issue
FAB-H-REQ-435 | TCPL pressure accuracy | The set pressure shall have accuracy ± (0.5 mbar + 3% of setting). | -low pressure -inaccurate pressure reading -fail to set value correctly | -incorrect measurement -pressure issue -failure to convert to back-up
FAB-H-REQ-436 | TCPL pressure with volume guarantee active | With Volume Guarantee enabled, the ventilator shall modulate the exhalation valve force in order to set the opening pressure for the exhalation valve to an algorithmically-determined pressure target as specified in section 7.1.3.1. Note: this auto-set target is intended to guarantee the exhaled tidal volume, and is derived from the operator set Volume, and operator set Pressure Limit. | -fail to adjust volume -fail to adjust settings -fail to hold pressure | -pressure issue -device difficult to setup or prepare -defective component
FAB-H-REQ-562 | Rise Time controlled breaths | When the Rise Time control is enabled, the mandatory breath type shall be a modified form of TCPL breath, with auto-set inspiratory flow, and ramped control of the exhalation valve pressure target waveform. | -incorrect breath type -fail to control flow -inadequate breath type | -defective item -compliance issue -pressure issue
FAB-H-REQ-437 | Rise time control breath - autoset flow control | When Rise Time control is enabled, the inspiratory flow shall be auto-set by the ventilator and independent of the I-Flow control. | -fail to control flow -inappropriate flow pattern -fail to activate control | -protective measure issue -defective item -defective component
FAB-H-REQ-438 | Rise time control breath - pressure control | When Rise Time control is enabled, the ventilator shall modulate the exhalation valve force to a target pressure waveform with a ramped increase to the set pressure level over the set Rise Time, followed by a plateau for the remaining I-Time. | -fail to update pressure -fail to adjust pressure -fail to control flow | -pressure issue -fluid leak -improper flow or infusion
FAB-H-REQ-563 | PSV Breaths | Ventilator breaths, provided in response to patient trigger events between mandatory breaths in SIMV+PSV mode, and to all patient trigger events in PSV mode, shall be PSV breaths. | -incorrect breath type -inappropriate patient type -inappropriate ventilation support | -infusion or flow issue -inaccurate delivery -improper device output
FAB-H-REQ-248 | PSV Breaths - Flow Control | The ventilator shall provide a fixed flow of oxygenated air to the breathing circuit throughout the duration of the inspiratory phase of a PSV breath, at the set I-Flow value. | -fail to control flow -oxygen pressure low -over flow | -excess flow or over-infusion -failure to cycle -Medical Gas Supply Problem
FAB-H-REQ-249 | PSV breath termination | PSV breaths shall be terminated when the inspiratory flow falls below a set proportion of the peak inspiratory flow, determined by the Termination-PSV control | -inappropriate control flow -fail to terminate process completely -fail to prevent flow | -failure to cycle -output above specifications -improper flow or infusion
FAB-H-REQ-250 | PSV termination control accuracy | The Termination-PSV control shall have accuracy ± (15 %). | -accuracy out of specification -low accuracy -low measurement accuracy | -defective component -temperature issue -failure to convert to back-up
FAB-H-REQ-251 | PSV breath termination on high pressure alarm | PSV breaths shall be terminated within 200 ms on activation of the High PIP alarm. | -fail to terminate process completely -decommission time long -fail to be radiused | -improper device output -device alarm system issue -infusion or flow issue
FAB-H-REQ-252 | PSV breath pressure control | The ventilator shall modulate the exhalation valve force in order to prevent the airway pressure from exceeding a set level : P-PSV (when configured as absolute value), PEEP + P-PSV (when configured as delta). Note: Configuration as absolute or pressure difference is specified in FAB-H-REQ-643 | -pressure out of range -fail to hold pressure -fail to reduce pressure | -pressure issue -device misassembled during manufacturing or shipping -device difficult to setup or prepare
FAB-H-REQ-253 | PSV breath flow control accuracy | The set flow shall have accuracy ± (0.5 lpm + 10% setting). | -low accuracy -low displacement accuracy -inaccurate flow rate | -incorrect measurement -improper flow or infusion -Inflation Issue
FAB-H-REQ-254 | PSV breath pressure control accuracy | The set pressure shall have accuracy ± (0.5 mbar + 3% of setting). | -low pressure -inaccurate pressure reading -fail to set value correctly | -incorrect measurement -pressure issue -failure to convert to back-up
FAB-H-REQ-564 | DuoPAP Breaths | The DuoPAP breath type shall be applicable only to a single limb breathing circuit with nCPAP generator. | -incorrect breath type -confused breath type -inadequate breath type | -use of device issue -medical gas supply problem -device difficult to setup or prepare
FAB-H-REQ-255 | DuoPAP breath duration | DuoPAP breaths shall have a set inspiratory duration, configured by the operator using the I-Time control. | -incorrect breath type -breath delayed -inappropriate inspiratory time | -use of incorrect control settings -inaccurate delivery -decreased sensitivity
FAB-H-REQ-782 | DuoPAP breath duration accuracy | The inspiratory duration for a DuoPAP breath shall have accuracy ± (0.05 s). | -inappropriate inspiratory time -incorrect breath type -low rate | -component missing -failure to convert to back-up -inaccurate delivery
FAB-H-REQ-256 | DuoPAP breath termination on high pressure alarm | DuoPAP breaths shall be terminated within 200 ms on activation of the High PIP alarm. | -fail to terminate process completely -fail to avoid incorrect connection terminal -incorrect breath type | -improper device output -device alarm system issue -infusion or flow issue
FAB-H-REQ-257 | DuoPAP breath flow control | The ventilator shall provide a fixed flow of oxygenated air to the breathing circuit throughout the duration of the inspiratory phase of a DuoPAP breath, pre-determined to achieve the target inspiratory pressure (P-High). Note: in DuoPAP mode, the inspiratory pressure level is relative to ambient and not additive to PEEP. Note 2: the flow required to achieve a given pressure target is different for the three supported nCPAP interfaces (Infant Flow, Infant Flow LP, Medijet) | -fail to control flow -over flow -inadequate flow | -excess flow or over-infusion -calibration issue -failure to cycle
FAB-H-REQ-258 | DuoPAP breath pressure accuracy | The set pressure shall have accuracy ± (0.5 mbar + 3% of setting). | -low pressure -inaccurate pressure reading -fail to set value correctly | -incorrect measurement -pressure issue -failure to convert to back-up
FAB-H-REQ-259 | Manual breath profile | Manual breaths shall be mandatory breaths in all modes, with inspiratory pressure determined by the Pinsp or Pmanual control; and with inspiratory phase continuing as long as the manual breath control is held, or until the maximum set time (specified in FAB-H-REQ-153) is reached. Note: The manual breath control is edge-triggered, and is disabled until 200ms following the end of previous inspiratory period. | -incorrect breath type -inadequate breath type -fail to initiate breath | -pressure issue -improper device output -device cleaning issue
FAB-H-REQ-783 | Manual breaths in IPPV, SIPPV, SIMV and SIMV+PSV modes | In IPPV, SIPPV, SIMV and SIMV+PSV modes, manual breaths shall be a mandatory breath with pressure limit determined by the “Pinsp” control, and with inspiratory phase continuing as long as the manual breath control is held, or until the maximum set time is reached. | -incorrect breath type -fail to adjust pressure -incorrect breath type setting | -pressure issue -failure to deliver -programming issue
FAB-H-REQ-260 | Manual breaths in PSV and CPAP mode | In CPAP mode, manual breaths shall be TCPL breaths with the pressure limit determined by the “P-Manual” control, and with inspiratory phase continuing as long as the manual breath control is held, or until the maximum set time is reached. | -incorrect breath type -fail to hold pressure -fail to adjust pressure | -pressure issue -failure to deliver -programming issue
FAB-H-REQ-261 | Manual breaths in nCPAP and DuoPAP modes | In nCPAP and DuoPAP modes, the ventilator shall provide a fixed flow of oxygenated air to the breathing circuit throughout the duration of the inspiratory phase of a manual breath, pre-determined to achieve the target pressure (“P-Manual” in nCPAP mode, “P-High” in DuoPAP mode) | -fail to control flow -air delivery compromised -fail to compatible with oxygen supply | -failure to cycle -unintended movement -excess flow or over-infusion
FAB-H-REQ-784 | Manual breaths in HFO mode | In HFO mode, the manual breaths shall be TCPL breaths with the pressure limit determined by the P-Manual control, with oscillation paused and with inspiratory phase continuing as long as the manual breath control is held, or until the maximum set time is reached. Note: The manual breath function can be disabled in HFO mode, see FAB-H-REQ-647. | -incorrect breath type -fail to initiate breath -fail to adjust pressure | -pressure issue -failure to deliver -incorrect software programming calculations
FAB-H-REQ-262 | Manual breath duration | The inspiratory duration of a manual breath shall be determined by the duration of the key-press: the breath continues (to a configured maximum within the range 2 to 30 seconds, see FAB-H-REQ-153) while the manual breath key remains depressed. | -incorrect breath type -fail to initiate breath -inappropriate inspiratory time | -defective item -inaccurate delivery -pressure issue
FAB-H-REQ-263 | Manual breath termination on high pressure alarm | Manual breaths shall be terminated within 200 ms on activation of the High PIP alarm. | -incorrect breath type -fail to terminate process completely -inappropriate inspiratory time | -improper device output -infusion or flow issue -Tip-over
FAB-H-REQ-264 | Manual breath pressure accuracy | The set pressure shall have accuracy ± (0.5 mbar + 3% of setting). | -low pressure -inaccurate pressure reading -fail to set value correctly | -incorrect measurement -pressure issue -failure to convert to back-up
FAB-H-REQ-565 | Standby mode when patient is disconnected | The fabian ventilator shall implement a standby mode, in which flow to the breathing circuit is reduced to a value primarily intended to prevent a connected heated humidifier from over-heating, breath delivery ceases, and all patient alarms are disabled. | -fail to cancel suspend mode -fail to reduce pressure -fail to enter mode | -output above specifications -calibration issue -bleed back
FAB-H-REQ-265 | Standby mode flow | In Standby mode, the flow to the breathing circuit shall be set to a constant 4 lpm. | -incorrect circuit type -fail to set tidal volume limit -inappropriate flow pattern | -defective item -failure to cycle -defective component
FAB-H-REQ-266 | Standby mode pressure control | In Standby mode, the exhalation valve control shall be constant set PEEP of 1.0 mbar (subject to tolerance for PEEP, see FAB-H-REQ-237). | -inaccurate settings -fail to generate periodic notification -fail to update indicator | -defective item -defective component -pressure issue
FAB-H-REQ-267 | Standby mode alarms disabled | In Standby mode, all patient alarms shall be disabled. | -fail to disable alarm -loss of state (activate or deactivate or alarm or warning) -disable pacing | -device alarm system issue -Device Handling Issue -false alarm
FAB-H-REQ-268 | Standby mode indication | During standby mode there shall be continuous visual indication of the standby mode status. | -fail to generate indication -loss of state indication -lack of indication | -optical issue -parameter calculation error due to software problem -no display or display failure
FAB-H-REQ-269 | Standby mode termination – dual-limb modes | With a dual-limb breathing circuit configured, the ventilator shall resume from Standby mode after a maximum of 120 s, or when the operator cancels by pressing the standby key. | -fail to resume -fail to enter mode -unable to repeat | -device alarm system issue -output above specifications -use of device issue
FAB-H-REQ-270 | Standby mode termination – single limb modes | With a single-limb breathing circuit configured, the ventilator shall resume from Standby mode when the operator cancels by pressing the standby key. Note: standby does not time out when configured for single-limb ventilation modes. | -fail to resume -fail to enter mode -fail to activate at start time | -use of device issue -failure to convert to back-up -calibration issue
FAB-H-REQ-566 | Tube Occlusion Alarm Active | Whenever the “Tube Occlusion” alarm is activated the fabian ventilator shall cease outlet gas flow and opens the exhalation valve to relieve pressure. | -fail to activate alarm -fail to limit flow -fail to activate | -defective component -device operates differently than expected -device maintenance issue
FAB-H-REQ-271 | Tube occlusion alarm pressure threshold | The tube occlusion alarm state shall be activated when the inspiratory pressure or the proximal pressure has exceeded the specified threshold (above intended peak inspiratory pressure) for a period of 50 milliseconds. | -fail to activate alarm -false alarm -loss of alarm | -delivery system failure -incorrect error code -pressure issue
FAB-H-REQ-272 | Occlusion alarm threshold – dual limb circuit | With a dual limb breathing circuit configuration, the occlusion alarm pressure threshold shall be 12 mbar. | -incorrect circuit type -inappropriate threshold setting -inappropriate pressure level | -low audible alarm -device alarm system issue -pressure issue
FAB-H-REQ-273 | Occlusion alarm threshold – nCPAP circuit | In ventilation modes nCPAP, DuoPAP, the occlusion alarm pressure threshold shall be 10 mbar. | -incorrect ventilation mode -inappropriate ventilation support -fail to activate alarm | -Pressure Issue -Device Alarm System Issue -Defective Alarm
FAB-H-REQ-274 | Occlusion alarm threshold – High Flow circuit | In ventilation mode O2 Therapy, the occlusion alarm shall be disabled. | -fail to disable ventilation -fail to disable alarm -disable pacing | -break -imprecision -medical gas supply problem
FAB-H-REQ-275 | Occlusion of primary gas outlet | The tube occlusion alarm state shall be activated in presence of an obstruction of the primary gas outlet. | -fail to activate alarm -false alarm -loss of alarm | -use of device issue -data issue -positioning issue
FAB-H-REQ-276 | Exhalation valve control during occlusion | When the tube occlusion alarm state is activated the exhalation valve control shall be set to zero (exhalation valve fully opened). | -fail to activate alarm -fail to activate control -valve stuck open | -defective component -defective item -device maintenance issue
FAB-H-REQ-277 | Flow control during occlusion | When the tube occlusion alarm state is activated the outlet gas flow shall be disabled (outlet flowrate set to zero). | -fail to activate alarm -fail to disable alarm -fail to disable ventilation | -false alarm -device maintenance issue -device alarm system issue
FAB-H-REQ-278 | Occlusion alarm de-activation | The tube occlusion alarm state shall be de-activated after the pressure has fallen below set PEEP, and the time of the next breath trigger has been reached, or after 1s in CPAP and nCPAP modes. | -fail to activate occlusion alarm -fail to activate alarm -loss of alarm | -use of device issue -deployment issue -use of incorrect control settings
FAB-H-REQ-439 | Circuit test for dual limb circuit | The Operators Manual shall include a procedure for testing the dual limb breathing circuit prior to patient use, that can detect a breathing circuit with leak, compliance or resistance such that ventilation performance may be impaired. | -inadequate test procedure -incorrect test procedure -fail to activate test | -failure to sense -failure to calibrate -device sensing issue
FAB-H-REQ-567 | Circuit test for single limb circuit and nCPAP generator | The Operators Manual shall include a procedure for testing the nCPAP breathing circuit and tubeset prior to patient use, that can detect a breathing circuit configuration with properties such that ventilation performance may be impaired. | -inadequate test procedure -lack of test configuration -lack of test tool | -failure to sense -device sensing issue -calibration issue
FAB-H-REQ-539 | Screen lock indication | While the screen is locked, an icon or other graphical representation will be provided to indicate the screen locked state. Note: this is implemented with a ‘Key’ icon. | -loss of state indication -fail to lock -invalid key | -No Display or Display Failure -no display or display failure -incorrect display
FAB-H-REQ-446 | Screen lock deactivation on alarm | The screen shall automatically unlock on activation of any alarm. | -fail to reset alarm automatically -fail to activate alarm -fail to unlock | -loss of data -failure to conduct -failure to deliver energy
FAB-H-REQ-447 | Screen unlock by operator | When the screen is locked, means shall be provided to unlock the screen using the rotary control. | -unable to lock screen -fail to lock -unlocking mechanism failed | -No Display or Display Failure -incorrect display -use of device issue
FAB-H-REQ-282 | Manual Breath control | A Manual Breath control shall be provided to allow the operator to trigger a manual breath. | -fail to initiate breath -fail to activate control -fail to adjust manually | -instruction for use issue -device cleaning issue -failure to deploy
FAB-H-REQ-568 | Membrane button label - Manual Breath | The Manual Breath membrane button shall be marked with the following icon: | -confused button -missing symbol -invisible settings | -inadequate instructions for healthcare professional -instruction for use issue -unauthorized access to computer system
FAB-H-REQ-569 | Oxygen Flush control | An oxygen flush control shall be provided to enable / disable the oxygen flush function | -fail to disable ventilation -fail to update oxygen volume -fail to compatible with oxygen supply | -No Display or Display Failure -defective component -calibration issue
FAB-H-REQ-570 | Membrane button label - O2 Flush | The O2 Flush membrane button shall be marked with the following icon: | -fail to indicate button status -confused button -lack of document | -sticking -no display or display failure -inadequate instructions for healthcare professional
FAB-H-REQ-448 | Standby mode control | A Standby control be provided to allow the ventilator to be activate or deactivate Standby mode. | -fail to activate control -fail to activate alarm -fail to deactivate | -device maintenance issue -device cleaning issue -defective item
FAB-H-REQ-571 | Membrane button label – Standby mode | The Start/Stop membrane button shall be labelled with the following icon: | -inadequate labeling -lack of button -incorrect labeling | -sticking -user used incorrect product for intended use -material rigid or stiff
FAB-H-REQ-449 | Action required to activate Standby | The standby control shall require an extended keypress minimum of 4 s and a GUI dialog requiring a confirmation touch-key press to enter standby mode. | -fail to confirm -fail to clear alarm -fail to hold inspiratory | -device unsafe to use in environment -improper or incorrect procedure or method -defective component
FAB-H-REQ-450 | Power On/Off control | A power control shall be provided to allow the ventilator to be switched on or off. | -power on or off switch failed -fail to monitor power event -loss of power | -installation-related problem -compliance issue -unintended movement
FAB-H-REQ-572 | Membrane button label - On/Off | The On/Off membrane button shall be marked with the following symbols: IEC 60417-5264 symbol for 'ON', IEC 60417-5265 symbol for 'OFF' | -fail to indicate button status -inadequate markings -missing symbol | -device markings issue -component or accessory incompatibility -inadequate user interface
FAB-H-REQ-451 | Action required to trigger power off | The power on/off control shall require to be held for a minimum of 4 s and a GUI dialog requiring a confirmation touch-key press to cause the ventilator to switch off. | -fail to indicate ac or battery power -fail to confirm -fail to detect power status | -device unsafe to use in environment -defective item -device difficult to setup or prepare
FAB-H-REQ-452 | Power off audible alarm | On switching the ventilator off, an audible alarm shall be activated for a minimum of 120 s, capable of being silenced by action of the Alarm Silence control. | -fail to silence alarm -fail to activate alarm -loss of audible alarm | -device alarm system issue -low audible alarm -low battery
FAB-H-REQ-284 | Menu / calibration control | A control shall be provided to allow the operator to access settings and calibration functions. | -fail to access -unable to access -unauthorized access | -human-device interface issue -compliance issue -device cleaning issue
FAB-H-REQ-573 | Membrane button label - Menu/Calibration | The Menu/Calibration membrane button shall be marked with the following icon: | -confused button -fail to indicate button status -lack of button | -sticking -inadequate instructions for healthcare professional -user used incorrect product for intended use
FAB-H-REQ-285 | Waves / Loops control | A control shall be provided to allow the operator to switch between displays including waveform graphs and loop graphs | -fail to manage switch -fail to implement control measure -loss of control | -device cleaning issue -protective measure issue -mechanical issue
FAB-H-REQ-574 | Membrane button label – Waves / Loops | The Waves / Loops membrane button shall be marked with the following icon: | -missing symbol -inadequate message display -loss of indication | -inadequate user interface -sticking -device markings issue
FAB-H-REQ-286 | Home screen control | A control shall be provided to allow the operator to access the main screen. | -fail to access -fail to activate control -unauthorized access | -human-device interface issue -detachment of device component -unauthorized access to computer system
FAB-H-REQ-575 | Membrane button label - Home | The Home membrane button shall be marked with the following icon: | -lack of button -missing symbol -incorrect settings | -inadequate instructions for healthcare professional -inadequate user interface -material rigid or stiff
FAB-H-REQ-287 | Alarm Limits control | A control shall be provided to allow the operator to access alarm threshold settings. | -unable to access -unauthorized access -fail to access | -human-device interface issue -unauthorized access to computer system -unexpected therapeutic results
FAB-H-REQ-576 | Membrane button - Alarm Limits | The Alarm Limits membrane button shall be marked with the following icon: | -incorrect message display -missing symbol -inappropriate symbol | -sticking -device alarm system issue -no display or display failure
FAB-H-REQ-453 | Alarm silence control | A control shall be provided to allow the operator to activate the alarm system Audio Pause function. | -fail to activate alarm -lack of pause function -fail to activate pause | -device alarm system issue -communication or transmission issue -device maintenance issue
FAB-H-REQ-454 | Alarm silence control label | The Alarm Silence control shall be labelled using icon IEC 60417-5576 with cross of broken lines. | -incorrect labeling -fail to clear alarm -fail to unidirectional communication | -electrical issue -defective item -electrical shorting
FAB-H-REQ-280 | Screen Freeze control | A Screen Freeze control shall be provided to allow the screen to be frozen (to allow review of saved waveforms and loops) | -frozen screen -unable to freeze -fail to disable control | -no device output -compliance issue -No Display or Display Failure
FAB-H-REQ-288 | Patient weight control | A control shall be provided to allow the operator to enter the patient weight, used for display of normalized tidal volume (volume per kg). Range: Off; 300 g to 30 kg, Unit of measure: g (300 to 5,000 g); kg (5kg to 30 kg), Resolution: From 300 g to 2,000 g: 50 g, From 2,000 g to 5,000 g: 200 g, From 5.00 kg to 10 kg: 0.2 kg, From 10 kg to 30 kg: 0.5 kg, Reset default: Off, Power on default: Off | -fail to deliver tidal volume -fail to display data -fail to transport patient | -device unsafe to use in environment -defective component -use of device issue
FAB-H-REQ-577 | Patient weight reset on change of patient size | The patient weight control shall be reset to Off when the patient size control (FAB-H-REQ-289) is changed. | -fail to activate control -fail to reset -fail to reset sensor | -loose or intermittent connection -user used incorrect product for intended use -device alarm system issue
FAB-H-REQ-289 | Patient size control | A control shall be provided to allow the operator to specify the patient size as either neonate, or pediatric. Note: this control is used to confirm the type of flow sensor that can be used; paediatric setting disables nCPAP and DuoPAP modes; and the control also limits the ranges of some controls. | -inappropriate patient type -fail to adjust size -fail to support non-spontaneous breathing patient | -user used incorrect product for intended use -compliance issue -device cleaning issue
FAB-H-REQ-290 | Patient size indication | The selected patient size shall be indicated continuously during clinical use. | -fail to indicate priority -fail to adjust size -incorrect sample size | -electrical issue -human-device interface issue -inadequate user interface
FAB-H-REQ-291 | Action on change of patient size | On change of the selected patient size, all controls will be adjusted to fall within the range for the control for the new patient size. Note: patient size acts as a limit to the ranges for many ventilation parameters | -fail to adjust size -incorrect sample size -inappropriate display size | -device difficult to setup or prepare -use of incorrect control settings -reset issue
FAB-H-REQ-292 | Presetting controls on change of mode | Means shall be provided to select a ventilation mode, and then adjust the ventilation parameters for that mode, prior to activating the mode change. | -incorrect ventilation mode -fail to maintain ventilation mode -inadequate ventilation | -use of device issue -device difficult to setup or prepare -failure to deploy
FAB-H-REQ-294 | Visual indication of selected control | Visual highlighting shall be provided to indicate that a control is selected and may be changed. | -fail to generate indication -fail to indicate priority -fail to identify change | -protective measure issue -no display or display failure -failure to sense
FAB-H-REQ-295 | Confirmation of control changes | All changes to ventilation controls shall require a confirmation step. | -fail to confirm -fail to approve change -lack of confirmation | -use of incorrect control settings -device markings issue -compliance issue
FAB-H-REQ-296 | Visual cue for control interaction | If two or more controls interact, a visual cue shall be provided to indicate the parameter requiring modification. The following interactions are implemented. | -fail to indicate priority -confused indication -loss of indication | -no display or display failure -device alarm system issue -device expiration issue
FAB-H-REQ-455 | O2 Flush interaction | The O2-Flush control shall be at least 2% above O2 setting when O2 setting is less than or equal to 98%, and no greater than 100% | -fail to retain settings -inappropriate flow volume -fail to compensate flow | -use of incorrect control settings -device difficult to setup or prepare -infusion or flow issue
FAB-H-REQ-456 | Volume Limit interaction | The Volume Limit control shall be available only if Volume Guarantee is disabled | -fail to disable control -inappropriate flow volume -fail to activate control | -use of device issue -protective measure issue -device alarm system issue
FAB-H-REQ-457 | Volume Guarantee interaction | The Volume Guarantee control shall be available only if Volume Limit is disabled | -fail to disable control -inappropriate flow volume -incorrect volume calculation | -use of device issue -protective measure issue -device alarm system issue
FAB-H-REQ-458 | P-Insp interaction | The P-Insp control shall be limited such that Pinsp ≥ PEEP + 4 mbar | -pin length out of specification -segment out of range -angle out of range | -electrical shorting -compliance issue -material rigid or stiff
FAB-H-REQ-459 | P-PSV interaction | The P-PSV control shall be less than or equal to P-Insp | -loss of power -incorrect settings -over utilization | -material puncture -torn material -defective item
FAB-H-REQ-460 | P-Duo interaction | The P-Duo control shall be no less than CPAP + 2 mbar | -low pressure -low power output -incorrect settings | -defective component -inadequate storage -biocompatibility issue
FAB-H-REQ-461 | P-Manual interaction | The P-Manual control (when mode is nCPAP) shall comply: P-Manual ≥ CPAP + 2 mbar | -inappropriate control flow -fail to comply with standards or regulation -incorrect mode | -electronic property issue -inadequate storage -failure to conduct
FAB-H-REQ-462 | Rise Time interaction | The Rise-Time shall be less than or equal to I-Time | -inappropriate inspiratory time -rise time out of range -incorrect start time | -volume accuracy issue -torn material -infusion or flow issue
FAB-H-REQ-463 | Freq, | The “Freq” control (modes that include mandatory breaths) and “Backup Rate” control (PSV mode only) shall each control the same internally stored parameter (i.e. changes to either control also set the other control to the same value) | -use incorrect control settings -incorrect settings -loss of control | -compliance issue -fitting problem -defective component
FAB-H-REQ-578 | Control Locked Values | Specified controls shall implement a 'locked' behaviour that requires confirmation to set an extreme value. | -fail to confirm -fail to lock -lack of locking mechanism | -improper flow or infusion -output issue -pressure issue
FAB-H-REQ-464 | Locked range for Pinsp | Pinsp control shall have locked range P-Insp > 25 mbar | -fail to lock -loss of control -latch force out of specification | -defective item -defective component -size incorrect for patient
FAB-H-REQ-465 | Locked range for CPAP control | CPAP control shall have locked range CPAP > 10 mbar | -fail to lock -fail to capture threshold -low readings | -electrical issue -device unsafe to use in environment -size incorrect for patient
FAB-H-REQ-466 | Locked range for PEEP control (upper) | PEEP control shall have locked range PEEP> 10 mbar | -fail to lock -low readings -low volume | -defective item -defective component -electrical issue
FAB-H-REQ-467 | Locked range for PEEP control (lower) | PEEP control shall have locked range PEEP< 2 mbar | -fail to lock -unable to lock screen -loss of control | -defective item -decreased sensitivity -defective component
FAB-H-REQ-468 | Locked range for PSV control | PSV control shall have locked range: P-PSV > 25 mbar or PEEP + P-PSV > 25 mbar, depending on configuration for P-PSV as absolute or above PEEP | -force out of range -loss of control -latch force out of specification | -use of incorrect control settings -mechanical issue -decreased sensitivity
FAB-H-REQ-469 | Locked range for V-Guarantee control | V-Guarantee control shall have locked range: V-Guarantee > 30 mL | -fail to lock -force out of range -latch force out of specification | -defective component -size incorrect for patient -Delivery System Failure
FAB-H-REQ-470 | Locked range for V-Limit control | V-Limit control shall have locked range: V-Limit > 30 mL | -fail to lock -force out of range -latch force out of specification | -defective component -size incorrect for patient -Delivery System Failure
FAB-H-REQ-471 | Locked range for P-Max control | P-Max control shall have locked range: P-Max > 25 mbar | -fail to lock -force out of range -latch force out of specification | -defective component -misconnection -size incorrect for patient
FAB-H-REQ-472 | Freq control specification (dual limb modes) | The “Freq” control shall have the following parameters: Units of measure: bpm, Modes applicable: IPPV, SIPPV, SIMV, SIMV+PSV, Range of control: IPPV, SIPPV, NEO: 2 to 200, SIMV, SIMV+PSV, NEO: 2 to 100, IPPV, SIPPV, PED: 2 to 100, SIMV, SIMV+PSV, PED: 2 to 75, esolution: 1, Reset Default: 60. Note 1: The frequency setting in DuoPAP mode and the frequency setting in invasive modes are distinct, and a change to the setting in invasive ventilation does not change the setting in DuoPAP mode, or vice versa. Note 2: This control depends on configuration for use of Freq and I-Time, or I-Time and E-Time, for control of breath timing. | -inadequate measure unit -inappropriate measurement unit -incorrect measurement | -defective item -defective component -device unsafe to use in environment
FAB-H-REQ-494 | Freq control specification (single limb ventilation) | The “Freq” control shall have the following parameters: Units of measure: bpm Modes applicable: DuoPAP, Range of control: 2 to 100, Resolution: 1, Reset Default: 30 Note: This control depends on configuration for use of Freq and I-Time, or I-Time and E-Time, for control of breath timing. | -inadequate measure unit -inappropriate measurement unit -parameter confused | -defective component -device unsafe to use in environment -protective measure issue
FAB-H-REQ-473 | Backup Rate control specification | The “Backup rate” control shall have the following parameters: Units of measure: bpm, Modes applicable: PSV, DuoPAP, Range of control: 2 – 100, Resolution: 1, Reset Default: 60 | -inappropriate measurement unit -inadequate measure unit -incorrect measurement | -image resolution poor -protective measure issue -defective item
FAB-H-REQ-474 | I-Time control specification (dual limb modes) | The “I-Time” control shall have the following parameters: Units of measure: seconds, Modes applicable: IPPV, SIPPV, SIMV, SIMV+PSV, PSV, Range of control: NEO: 0.10 to 2.00, PED: 0.30 to 2.00; Resolution: In range [0.10s;1.00s]: ± 0.01s, In range [1.00s;2.00s]: ± 0.02s, In range [2.00s;5.00s]: ± 0.05s, In range [5.00s;10.0s]: ± 0.10s, In range [10.0s;20.0s]: ± 0.20s, In range [20.0s;30.0s]: ± 0.50s. Reset Default: 0.5 | -inadequate measure unit -inappropriate measurement unit -incorrect measurement | -defective item -device unsafe to use in environment -defective component
FAB-H-REQ-495 | I-Time control specification (single limb ventilation modes) | The “I-Time” control shall have the following parameters: Units of measure: seconds, Modes applicable: DuoPAP, Range of control: 0.15 to 15.00, Resolution: In range [0.15s;1.00s]: ± 0.01s, In range [1.00s;2.00s]: ± 0.02s, In range [2.00s;5.00s]: ± 0.05s, In range [5.00s;15.0s]: ± 0.10s, Reset Default: 1.0, Labelling: When I-Time > 0.60 s the label for this control changes to “T-high” | -inadequate measure unit -inappropriate measurement unit -incorrect measurement | -defective item -unintended system motion -defective component
FAB-H-REQ-475 | E-time control specification (dual limb modes) | The “E-Time” control shall have the following parameters: Units of measure: seconds; Modes applicable: IPPV, SIPPV, SIMV, SIMV+PSV, PSV; Range of control: IPPV, SIPPV, PSV: 0.20 to 30.00, SIPPV, SIMV+PSV: 0.50 to 30.00; Resolution: In range [0.20s;1.00s]: ± 0.01s, In range [1.00s;2.00s]: ± 0.02s, In range [2.00s;5.00s]: ± 0.05s, In range [5.00s;10.0s]: ± 0.10s, In range [10.0s;20.0s]: ± 0.20s, In range [20.0s;30.0s]: ± 0.50s. Reset Default: 0.5; Labelling: When I-Time > 0.60 s the label for this control changes to “T-low” Note: This control depends on configuration for use of Freq and I-Time, or I-Time and E-Time, for control of breath timing | -inadequate measure unit -inappropriate measurement unit -incorrect measurement | -defective item -defective component -device unsafe to use in environment
FAB-H-REQ-496 | E-time control specification (single limb ventilation modes) | The “E-Time” control shall have the following parameters: Units of measure: seconds; Modes applicable: DuoPAP; Range of control: 0.20 to 30.00; Resolution: In range [0.20s;1.00s]: ± 0.01s, In range [1.00s;2.00s]: ± 0.02s, In range [2.00s;5.00s]: ± 0.05s, In range [5.00s;10.0s]: ± 0.10s, In range [10.0s;20.0s]: ± 0.20s, In range [20.0s;30.0s]: ± 0.50s; Reset Default: 1.0; Note: This control depends on configuration for use of Freq and I-Time, or I-Time and E-Time, for control of breath timing | -inappropriate measurement unit -inadequate measure unit -incorrect measurement | -defective item -unintended system motion -defective component
FAB-H-REQ-476 | I-Flow control specification | The “I-Flow” control shall have the following parameters: Units of measure: L/min; Modes applicable: IPPV, SIPPV, SIMV, SIMV+PSV, PSV; Range of control: 1 to 32; Resolution: In range [1.0;5.0]: 0.1 l/min, In range [5.0;10]: 0.5 l/min, In range [10;32]: 1.0 l/min; Reset Default: 8 | -inadequate measure unit -inappropriate measurement unit -incorrect measurement | -defective item -defective component -device unsafe to use in environment
FAB-H-REQ-477 | E-Flow control specification | The “E-Flow” control shall have the following parameters: Units of measure: L/min; Modes applicable: IPPV, SIPPV, SIMV, SIMV+PSV, PSV; Range of control: 1 to 32; Resolution: 0.1; Reset Default: 4 | -inappropriate measurement unit -inadequate measure unit -incorrect measurement | -defective component -device unsafe to use in environment -defective item
FAB-H-REQ-579 | Flow control specification (HFO mode only) | The “Flow” control shall have the following parameters: Units of measure: L/min; Modes applicable: HFO; Range of control: 5 to 20; Resolution: 0.1; Reset Default: 8 | -inappropriate measurement unit -inadequate measure unit -incorrect measurement | -defective component -protective measure issue -device unsafe to use in environment
FAB-H-REQ-478 | Flow-Min control specification (CPAP mode only) | The “Flow-Min” control shall have the following parameters: Units of measure: L/min; Modes applicable: CPAP; Range of control: 4 to 16; Resolution: 0.1; Reset Default: 4 | -inadequate measure unit -inappropriate measurement unit -incorrect measurement | -device unsafe to use in environment -defective component -defective item
FAB-H-REQ-580 | Pressure rise control specification | The “Pressure Rise” control shall have the following parameters: Units of measure: N/A; Modes applicable: IPPV, SIPPV, SIMV. SIMV+PSV, PSV; Range of control: “I-Flow” or “Rise Time”; Resolution: N/A; Reset Default: “I-Flow” | -inappropriate measurement unit -inadequate measure unit -incorrect measurement | -defective item -image resolution poor -defective component
FAB-H-REQ-479 | Rise Time control specification | The “Rise Time” control shall have the following parameters: Units of measure: s; Modes applicable: IPPV, SIPPV, SIMV. SIMV+PSV, PSV; Range of control: NEO: 0.10 to 2.00, PED: 0.30 to 2.00, Resolution: 0.10; Reset Default: 0.10 | -inadequate measure unit -inappropriate measurement unit -incorrect measurement | -defective item -image resolution poor -device unsafe to use in environment
FAB-H-REQ-480 | PEEP control specification | The “PEEP” control shall have the following parameters: Units of measure: mbar (default), cmH2O; Modes applicable: IPPV, SIPPV, SIMV, SIMV+PSV, PSV; Range of control: 1 to 30; Resolution: 0.5; Reset Default: 5 | -inadequate measure unit -inappropriate measurement unit -incorrect measurement | -defective component -defective item -device unsafe to use in environment
FAB-H-REQ-481 | CPAP control specification | The “CPAP” control shall have the following parameters: Units of measure: mbar (default), cmH2O; Modes applicable: CPAP, nCPAP, DuoPAP; Range of control: CPAP: 1 to 30, nCPAP: 2 to 13, DuoPAP: 2 to 13, Resolution: 0.5, Reset Default: 5 Note: CPAP and PEEP are distinct labels for the same internal parameter. | -inadequate measure unit -inappropriate measurement unit -incorrect measurement | -defective item -device unsafe to use in environment -defective component
FAB-H-REQ-482 | P-insp control specification | The “P-insp” control shall have the following parameters: Units of measure: mbar (default), cmH2O; Modes applicable: IPPV, SIPPV, SIMV, SIMV+PSV, PSV; Range of control: 4 to 80; Resolution: 1; Reset Default: 15 Note: In PSV mode, this parameter is labelled as Pbackup. | -inadequate measure unit -inappropriate measurement unit -incorrect measurement | -defective component -device unsafe to use in environment -defective item
FAB-H-REQ-483 | P-max control specification | The “P-max” control shall have the following parameters: Units of measure: mbar (default), cmH2O; Modes applicable: IPPV, SIPPV, SIMV, SIMV+PSV, PSV; Range of control: 4 to 80; Resolution: 1; Reset Default: 15 | -inadequate measure unit -inappropriate measurement unit -incorrect measurement | -defective component -device unsafe to use in environment -defective item
FAB-H-REQ-484 | P-PSV control specification | The “P-PSV” control shall have the following parameters: Units of measure: mbar (default), cmH2O; Modes applicable: SIMV+PSV, PSV; Range of control: 3 to 80; Resolution: 0.5; Reset Default: 7 | -inadequate measure unit -inappropriate measurement unit -incorrect measurement | -defective component -protective measure issue -optical issue
FAB-H-REQ-485 | P-manual control specification | The “P-manual” control shall have the following following parameters: Units of measure: mbar (default), cmH2O; Modes applicable: CPAP, nCPAP, HFO; Range of control: CPAP: 4 to 80, nCPAP: 5 to 15, HFO: 6 to 80; Resolution: 1; Reset Default: 15 Note: P-manual is a re-labelling for P-insp in these modes, and the values should be identical on switching between modes | -inadequate measure unit -inappropriate measurement unit -incorrect measurement | -defective component -device unsafe to use in environment -protective measure issue
FAB-H-REQ-581 | Pmean control specification (HFO ventilation mode) | The “Pmean” control shall have the following parameters: Units of measure: mbar (default), cmH2O; Modes applicable: HFO; Range of control: 5 to 50; Resolution: 0.5; Reset Default: 10 | -inadequate measure unit -inappropriate measurement unit -incorrect measurement | -defective component -protective measure issue -device unsafe to use in environment
FAB-H-REQ-582 | HFO Amp control specification (HFO ventilation mode) | The “HFO Amp” control shall have the following parameters: Units of measure: mbar (default), cmH2O; Modes applicable: HFO; Range of control: 5 to 100; Resolution: 1; Reset Default: 20 | -inadequate measure unit -inappropriate measurement unit -incorrect measurement | -defective component -protective measure issue -device unsafe to use in environment
FAB-H-REQ-583 | HFO Amp,max control specification (HFO ventilation mode with VG enabled) | The “HFO Amp,max” control shall have the following parameters: Units of measure: mbar (default), cmH2O; Modes applicable: HFO; Range of control: 5 to 100; Resolution: 1; Reset Default: 20 | -inadequate measure unit -inappropriate measurement unit -parameter confused | -defective component -protective measure issue -device unsafe to use in environment
FAB-H-REQ-584 | HF Freq control specification (HFO ventilation mode) | The “HF Freq” control shall have the following parameters: Units of measure: Hz; Modes applicable: HFO; Range of control: 5 to 20; Resolution: 1; Reset Default: 10 | -inadequate measure unit -inappropriate measurement unit -loss of measurement | -defective component -protective measure issue -device unsafe to use in environment
FAB-H-REQ-585 | I:E control specification (HFO ventilation mode) | The “I:E” control shall have the following parameters: Units of measure: N/A; Modes applicable: HFO; Range of control: 1:1 to 1:3; Resolution: 1:1; Reset Default: 1:3 | -inadequate measure unit -inappropriate measurement unit -incorrect measurement | -defective component -defective item -device unsafe to use in environment
FAB-H-REQ-586 | Pmean,rec control specification (HFO ventilation mode) | The “Pmean,rec” control shall have the following parameters: Units of measure: mbar (default), cmH2O; Modes applicable: HFO; Range of control: 7 to 50; Resolution: 0.5; Reset Default: 20. Note: This is the mean pressure during the HFO lung recruitment procedure. | -inadequate measure unit -inappropriate measurement unit -incorrect measurement | -device unsafe to use in environment -use of device issue -defective component
FAB-H-REQ-587 | I-time,rec control specification (HFO ventilation mode) | The “I-time,rec” control shall have the following parameters: Units of measure: s; Modes applicable: HFO; Range of control: 2 to 60; Resolution: 1; Reset Default: 5 Note: This is the duration of the period of elevated mean pressure during the HFO lung recruitment procedure. | -inadequate measure unit -inappropriate measurement unit -incorrect measurement | -protective measure issue -electronic property issue -incorrect measurement
FAB-H-REQ-588 | Freq,rec control specification (HFO ventilation mode) | The “Freq,rec” control shall have the following parameters: Units of measure: 1/hr, 1/min; Modes applicable: HFO; Range of control: Off, 1/hour to 59/hour, 1/min to 4/min; Resolution: Range [1/hour to 59/hour]: 1/hour, Range [1/min to 4/min]: 1/min; Reset Default: Off Note: This parameter is the frequency with which the recruitment procedure is repeated. The recruitment procedure is repeated at intervals between 15 s and 60 mins. | -inadequate measure unit -inappropriate measurement unit -incorrect measurement | -defective item -device unsafe to use in environment -unintended system motion
FAB-H-REQ-486 | O2 control specification | The “O2” control shall have the following parameters: Units of measure: %; Modes applicable: ALL; Range of control: 21 to 100; Resolution: 1; Reset Default: 21 | -inadequate measure unit -inappropriate measurement unit -incorrect measurement | -defective component -defective item -device unsafe to use in environment
FAB-H-REQ-487 | O2 Flush control specification | The ”O2 Flush” control shall have the following parameters: Units of measure: %; Modes applicable: ALL; Range of control: 23 to 100; Resolution: 1; Reset Default: 23 | -inadequate measure unit -inappropriate measurement unit -current out of range | -defective component -defective item -device unsafe to use in environment
FAB-H-REQ-488 | V-limit control specification | The “V-limit” control, available only when VG is Off, shall have the following parameters: Units of measure: mL; Modes applicable: IPPV, SIPPV, SIMV, SIMV+PSV, PSV; Range of control: NEO: Off, 1 to 150; PED: Off, 10 to 500; Resolution: 1; Reset Default: Off / 1 Note: The V-Limit control does not have a value “Off” but V-Limit function can be separately set On or Off. V-Limit is only available when V-Guarantee is Off, see FAB-H-REQ-456 | -inadequate measure unit -inappropriate measurement unit -fail to measure | -defective item -unintended system motion -device alarm system issue
FAB-H-REQ-489 | V-guarantee control specification | The “V-guarantee” control, available only when VLimit is Off, shall have the following parameters: Units of measure: mL; Modes applicable: IPPV, SIPPV, SIMV, SIMV+PSV, PSV, HFO; Range of control: NEO: Off, 2 to 60 PED: Off, 10 to 300 HFO: 0.3 to 30; Resolution: 1; Reset Default: Off / 1.2 (Conv.) / 0.8 (HFO) Note: The V-Guarantee control does not have a value “Off” but V-Guarantee function can be separately set On or Off. V-Guarantee is only available when V-Limit is Off, see FAB-H-REQ-457 | -inadequate measure unit -inappropriate measurement unit -fail to measure | -unintended system motion -defective item -failure to convert to back-up
FAB-H-REQ-490 | Trigger mode control specification | The “Trigger Mode” control shall have the following parameters: Units of measure: N/A; Modes applicable: SIPPV, SIMV, SIMV+PSV, PSV; Range of control: Pressure, Flow, Volume; Resolution: N/A; Reset Default: Flow Note: The NIV Trigger function available in DuoPAP mode uses the tube set configuration to determine trigger, “Trigger Mode” is not used. | -inadequate measure unit -inappropriate measurement unit -incorrect parameter | -defective component -protective measure issue -defective item
FAB-H-REQ-491 | Trigger control specification (dual limb ventilation) | The “Trigger” control shall have the following parameters: Units of measure: N/A; Modes applicable: SIPPV, SIMV, SIMV+PSV, PSV; Range of control: 1 to 10; Resolution: 1; Reset Default: 5 Note: Values of 1 to 10 have different physical meaning depending on the trigger mode and patient size, but value 1 is most sensitive (detects smallest amplitude breathing) | -inadequate measure unit -inappropriate measurement unit -incorrect parameter | -protective measure issue -defective component -device unsafe to use in environment
FAB-H-REQ-800 | Trigger control specification (NIV Trigger) | The “Trigger” control active in DuoPAP mode shall have following parameters: Units of measure: N/A; Modes applicable: DuoPAP; Range of control: 1 to 10, Off; Resolution: 1; Reset Default: Off | -inadequate measure unit -inappropriate measurement unit -incorrect parameter | -defective component -defective item -device unsafe to use in environment
FAB-H-REQ-492 | Termination-PSV control specification | The “Termination-PSV” control shall have the following parameters: Units of measure: %; Modes applicable: SIMV+PSV, PSV; Range of control: 1 to 85; Resolution: 1; Reset Default: 15 | -inadequate measure unit -inappropriate measurement unit -incorrect parameter | -defective item -defective component -device unsafe to use in environment
FAB-H-REQ-493 | Backup control specification | The “Backup” control shall have the following parameters: Units of measure: N/A (number of breaths); Modes applicable: CPAP; Range of control: Off, 1 to 5; Resolution: 1; Reset Default: Off | -inadequate measure unit -inappropriate measurement unit -fail to measure | -device unsafe to use in environment -defective component -defective item
FAB-H-REQ-497 | Flow control specification (single limb ventilation modes) | The “Flow” control shall have the following parameters: Units of measure: L/min; Modes applicable: O2 Therapy; Range of control: NEO: 1 to 15, PED: 1 to 30; Resolution: Range 1 to 5: 0.2, Range 5.0 to 30: 1; Reset Default: 4 | -inadequate measure unit -inappropriate measurement unit -incorrect measurement | -defective item -positioning issue -defective component
FAB-H-REQ-500 | P-duo control specification | The “Pduo” control shall have the following parameters: Units of measure: mbar (default), cmJH2O; Modes applicable: DuoPAP; Range of control: 5 to 15; Resolution: 1; Reset Default: 15 | -inadequate measure unit -inappropriate measurement unit -incorrect measurement | -defective component -protective measure issue -defective item
FAB-H-REQ-589 | Leak Compensation control specification | The “Leak Comp” control shall have the following parameters: Units of measure: N/A; Modes applicable: nCPAP, DuoPAP; Range of control: Off, Low, Medium, High; Resolution: N/A; Reset Default: Off | -inadequate measure unit -inappropriate measurement unit -fail to measure | -defective component -device unsafe to use in environment -defective item
FAB-H-REQ-590 | FOT amplitude | The FOT amplitude shall be fixed: 5 mbar | -inaccurate amplitude measurement -density out of specification -frequency out of range | -patient data issue -defective component -programming issue
FAB-H-REQ-591 | FOT frequency of forced oscillation | The FOT forced oscillation frequency shall be fixed: 10 Hz | -frequency out of specification -frequency out of range -inaccurate frequency | -electronic property issue -use of device issue -parameter calculation error due to software problem
FAB-H-REQ-592 | FOT forced oscillation I:E ratio | The FOT forced oscillation I:E ratio shall be fixed: 1:1 | -fail to adjust pressure -unstable threshold -fail to update pressure | -delivered as unsterile product -biocompatibility issue -patient data issue
FAB-H-REQ-593 | FOT period | The FOT period shall be less than or equal to 18 seconds. Note: The FOT measurement period is 3 seconds; the FOT algorithm implements a settling time of up to 15 seconds prior to measurement. | -inappropriate measurement unit -inaccurate measurement -long charge time | -inaccurate delivery -volume accuracy issue -improper device output
FAB-H-REQ-594 | FOT stabilization time | The FOT stabilization period shall be less than or equal to 180 seconds. Note: The FOT stabilization period can be skipped by the user. | -short validation period -inappropriate stability change -time out of range | -inaccurate delivery -device difficult to setup or prepare -installation-related problem
FAB-H-REQ-595 | FOT steps | The FOT algorithm step count control “FOT Steps” shall have following parameters. Units of measure: N/A; Modes applicable: IPPV, SIPPV, SIMV, SIMV+PSV, PSV, CPAP, HFO; Range of control: 1 to 21 (odd values only); Resolution: 2; Reset Default: 11 | -inadequate measure unit -inappropriate measurement unit -incorrect measurement | -defective component -device unsafe to use in environment -defective item
FAB-H-REQ-596 | FOT Pmean titration (FOT during HFO) | During the FOT procedure in HFO mode, “Pmean” shall step from “Pmean,low” to “Pmean,high” and back in “FOT steps” discrete steps. | -inappropriate intensity step -incorrect measurement -inappropriate measurement unit | -electrical issue -mechanical issue -use of device issue
FAB-H-REQ-597 | FOT Pmean,low | “Pmean,low” control shall have following parameters: Units of measure: mbar, cmH2O; Modes applicable: HFO; Range of control: range 5 to 49; Resolution: 1; Reset Default: 10 | -inadequate measure unit -inappropriate measurement unit -measurement low | -defective component -defective item -device unsafe to use in environment
FAB-H-REQ-598 | FOT Pmean,high | “Pmean,high” control shall have following parameters: Units of measure: mbar, cmH2O; Modes applicable: HFO; Range of control: range 6 to 50; Resolution: 1; Reset Default: 20 | -inadequate measure unit -inappropriate measurement unit -incorrect measurement | -defective item -defective component -device unsafe to use in environment
FAB-H-REQ-599 | FOT PEEP titration (FOT during during conventional ventilation) | During the FOT procedure in conventional ventilation “PEEP” shall step from “PEEP,low” to “PEEP,high” and back in “FOT steps” discrete steps. | -inappropriate intensity step -incorrect ventilation mode -incorrect measurement | -mechanical issue -electrical issue -use of device issue
FAB-H-REQ-600 | FOT PEEP,low | “PEEP,low” control shall have following parameters: Units of measure: mbar, cmH2O; Modes applicable: IPPV, SIPPV, SIMV, SIMV+PSV, PSV, CPAP; Range of control: range 3 to 29; Resolution: 1; Reset Default: 10 | -inadequate measure unit -inappropriate measurement unit -measurement low | -defective item -unintended system motion -defective component
FAB-H-REQ-601 | FOT PEEP,high | “PEEP,high” control shall have following parameters: Units of measure: mbar, cmH2O; Modes applicable: PPV, SIPPV, SIMV, SIMV+PSV, PSV, CPAP; Range of control: range 4 to 30; Resolution: 1; Reset Default: 20 | -inadequate measure unit -inappropriate measurement unit -incorrect measurement | -defective item -device unsafe to use in environment -defective component
FAB-H-REQ-297 | Indication of calculated E-Time | The ventilator shall provide indication of expiratory time, calculated from set I-Time and Freq. Note: Applies when configured to set I-Time and Freq | -loss of indication -lack of indication -invisible indication | -pressure issue -calibration issue -infusion or flow issue
FAB-H-REQ-298 | Calculated E-time display range | The calculated “E-Time” display shall cover at least the range 0.2 to 30.0 s. | -incorrect display format -out of range -motion out of range | -incorrect display -device alarm system issue -parameter calculation error due to software problem
FAB-H-REQ-299 | Indication of calculated Freq | The ventilator shall provide indication of frequency, calculated from set I-Time and E-Time. Note: Applies when configured to set I-Time and E-Time | -loss of indication -fail to indicate priority -invisible indication | -pressure issue -overheating of device or device component -infusion or flow issue
FAB-H-REQ-302 | Calculated Freq display range | The calculated “Freq” display shall cover at least the range 1 to 150 bpm. | -incorrect display format -motion out of range -out of range | -incorrect display -output below specifications -device alarm system issue
FAB-H-REQ-300 | Indication of calculated I:E Ratio | The ventilator shall provide indication of I:E Ratio, calculated from set I-Time and Rate. | -loss of indication -invisible indication -fail to indicate priority | -incorrect measurement -pressure issue -calibration issue
FAB-H-REQ-301 | Calculated I:E Ratio format | I:E ratio values shall be displayed as 1:xx.x (inspiratory time less than or equal to expiratory time) or xx.x:1 (inspiratory time greater than expiratory time). | -incorrect display format -incorrect message display -inadequate message display | -incorrect display -use of incorrect control settings -torn material
FAB-H-REQ-602 | Calculated PIP,max | PIP,max shall be calculated as the maximally reached peak pressure expected during the whole FOT recruitment procedure. Notes: PIP,max = (Pinsp – PEEP) + PEEP,high. PIPmax is limited to 100 mbar in HFO mode and to 80 mbar for Conventional modes. | -incorrect bolus calculation -fail to adjust pressure -inaccurate pressure reading | -use of incorrect control settings -incorrect display -parameter calculation error due to software problem
FAB-H-REQ-603 | Calculated FOT step | Step shall be calculated as the pressure difference between two consecutive FOT measurements. It is automatically calculated from the difference between the low and high-pressure settings during FOT. The minimum step is 1 mbar. | -inaccurate pressure reading -incorrect measurement -fail to adjust pressure | -use of incorrect control settings -parameter calculation error due to software problem -device markings issue
FAB-H-REQ-609 | Airway Pressure monitor | Measured values derived from time-varying measurements of airway pressure shall share a common specification for accuracy. Airway Pressure measurements use a proximal line to provide for monitoring of airway pressure at the patient interface. | -low measurement accuracy -orientation accuracy out of specification -incorrect measurement | -Low Battery -incorrect measurement -Incorrect Measurement
FAB-H-REQ-57 | Airway pressure measurement range | Measurements of airway pressure shall cover at least the range -10 to +125 mbar. | -inappropriate measurement unit -incorrect measurement -inadequate measure unit | -pressure issue -incorrect measurement -device difficult to setup or prepare
FAB-H-REQ-58 | Airway pressure measurement accuracy | Airway pressure values shall have accuracy ±(0.5 mbar + 4% of actual value). | -inaccurate pressure reading -accuracy out of specification -incorrect readings | -incorrect measurement -calibration issue -failure to convert to back-up
FAB-H-REQ-36 | Monitor for PIP | The PIP monitor shall be available in following modes: IPPV, SIPPV, SIMV, SIMV+PSV, PSV | -lack of display -incorrect screen display -inadequate interface specification | -protective measure issue -device maintenance issue -patient-device incompatibility
FAB-H-REQ-303 | PIP specification - modes IPPV, SIPPV, SIMV | In modes IPPV, SIPPV and SIMV the value PIP shall represent the peak airway pressure during the period of each mandatory breath, updated prior to the start of the next mandatory breath. Note: in modes IPPV, SIPPV, SIMV and DuoPAP only mandatory breaths are delivered. In mode PSV only PSV breaths are delivered other than during apnea backup ventilation. | -fail to update pressure -incorrect breath type -incorrect value | -incorrect measurement -parameter calculation error due to software problem -infusion or flow issue
FAB-H-REQ-610 | PIP specification - mode PSV | In mode PSV the value PIP shall represent the peak airway pressure during the period of each PSV breath, updated prior to the start of the next PSV breath. | -incorrect breath type -inappropriate flow pattern -loss of indication | -incorrect measurement -parameter calculation error due to software problem -pressure issue
FAB-H-REQ-517 | PIP specification - SIMV+PSV mode | In mode SIMV+PSV, the value PIP shall represent the peak airway pressure during the period of each mandatory breath, updated prior to the start of the next mandatory breath. Note: in mode SIMV+PSV, a mix of mandatory and PSV breaths are delivered, with PIP typically higher for mandatory breaths than for PSV breaths, and it would be confusing for the operator to have this monitor jump between distinct values at each breath. | -fail to update pressure -incorrect value -inappropriate flow pattern | -use of device issue -parameter calculation error due to software problem -incorrect measurement
FAB-H-REQ-304 | PIP display range | PIP shall be displayed within the range 0 to +120 mbar | -inappropriate display size -incorrect message display -inadequate message display | -no display or display failure -device alarm system issue -device difficult to setup or prepare
FAB-H-REQ-611 | Pduo specification - DuoPAP | In mode DuoPAP the value Pduo shall represent the peak airway pressure during the period of each DuoPAP breath, updated prior to the start of the next DuoPAP breath. | -incorrect breath type -incorrect value -inappropriate flow pattern | -incorrect measurement -parameter calculation error due to software problem -failure to service
FAB-H-REQ-612 | Pduo display range | Pduo shall be displayed within the range 0 to +120 mbar | -inappropriate display size -incorrect message display -incorrect display format | -no display or display failure -device difficult to setup or prepare -device alarm system issue
FAB-H-REQ-37 | Monitor for Pmean | The Pmean monitor shall be available in following modes: IPPV, SIPPV, SIMV, SIMV+PSV, PSV, CPAP, nCPAP, DuoPAP, O2 Therapy, HFO | -incorrect display format -incorrect screen display -inadequate user interface | -patient-device incompatibility -device maintenance issue -protective measure issue
FAB-H-REQ-305 | Pmean specification | The value Pmean shall represent the average airway pressure during the whole of the inspiratory and expiratory period of each breath, updated prior to the start of the next mandatory or PSV breath. | -inaccurate pressure reading -incorrect breath type -incorrect value | -pressure issue -no pressure -incorrect measurement
FAB-H-REQ-306 | Pmean display range | Pmean shall be displayed within the range 0 to +120 mbar. | -incorrect display format -incorrect message display -inappropriate display size | -no display or display failure -device difficult to setup or prepare -device alarm system issue
FAB-H-REQ-613 | HFO Amp specification | The value HFO Amp shall represent the peak to peak pressure value during high frequency oscillatory ventilation, averaged over a period of TBD . | -fail to update pressure -inappropriate peak flow -incorrect value | -improper device output -parameter calculation error due to software problem -improper flow or infusion
FAB-H-REQ-614 | HFO Amp display range | In HFO mode, the HFO Amplitude shall be displayed as a peak to peak pressure value with range 0 to +120 mbar. | -incorrect display format -out of range -inadequate message display | -pressure issue -use of incorrect control settings -incorrect display
FAB-H-REQ-38 | Monitor for PEEP | The PEEP monitor shall be available in following modes: IPPV, SIPPV, SIMV, SIMV+PSV, PSV, DuoPAP | -inadequate evaluation -unable to view display -unable to program | -device maintenance issue -protective measure issue -environmental control or utility issue
FAB-H-REQ-307 | PEEP specification | The value PEEP shall represent the airway pressure at the end of the expiratory period of each breath, excluding any period during which airway pressure is reduced by patient spontaneous inspiratory effort, updated prior to the start of the next mandatory or PSV breath. | -incorrect value -incorrect measurement -incorrect message display | -pressure issue -calibration issue -no pressure
FAB-H-REQ-308 | PEEP display range | PEEP shall be displayed within the range 0 to +60 mbar. | -inappropriate display size -incorrect message display -inadequate message display | -device difficult to setup or prepare -use of incorrect control settings -decreased sensitivity
FAB-H-REQ-615 | Volume | Measured values derived from time-varying measurements of airway flow shall share a common specification for range and accuracy. Volume measurements use a proximal hot-wire flow sensor to provide for monitoring of airway flow at the patient connection port. Volume measurements are not available with the single limb breathing circuit configurations, or with the Hot Wire Flow Sensor disabled. | -measurement range out of specification -low measurement accuracy -measurement high | -Pressure Issue -incorrect measurement -Inflation Issue
FAB-H-REQ-59 | Volume display range | Measurements of tidal volume shall cover at least the range 0 to +400 mL. | -fail to deliver tidal volume -inappropriate measurement unit -loss of measurement | -use of device issue -parameter calculation error due to software problem -output below specifications
FAB-H-REQ-60 | Tidal volume measurement accuracy | Tidal Volume values shall have accuracy ±(0.5 mL + 8 % of measurement). | -low measurement accuracy -inaccurate volume -inaccurate measurement | -incorrect measurement -parameter calculation error due to software problem -output below specifications
FAB-H-REQ-40 | Monitor for Vte | The Vte monitor shall be available in following modes: IPPV, SIPPV, SIMV (mand + spon), SIMV+PSV (mand + spon), PSV, CPAP | -unable to program -lack of configuration -unable to connect power | -protective measure issue -inadequate instructions for healthcare professional -particulates
FAB-H-REQ-61 | Vte | The value Vte shall represent the exhaled tidal volume during the period of each mandatory or PSV breath, updated prior to the start of the next mandatory or PSV breath. | -fail to calculate tidal volume -fail to deliver tidal volume -incorrect wave amplitude | -failure to service -volume accuracy issue -parameter calculation error due to software problem
FAB-H-REQ-41 | Monitor for Vte/kg | The Vte/kg monitor shall be available in following modes (defined as Vte, divided by the patient weight): IPPV, SIPPV, SIMV (mand + spon), SIMV+PSV (mand + spon), PSV, CPAP | -inadequate interface specification -inappropriate user -unable to program | -inadequate instructions for healthcare professional -patient data issue -protective measure issue
FAB-H-REQ-42 | Monitor for Vti | The Vti monitor shall be available in following modes: IPPV, SIPPV | -incorrect display format -inadequate user interface -inadequate interface specification | -protective measure issue -device maintenance issue -patient-device incompatibility
FAB-H-REQ-62 | Vti | The value Vti shall represent the delivered tidal volume during the period of each mandatory or PSV breath, updated prior to the start of the next mandatory or PSV breath. Note: Vti is not displayed, but is used in the calculation of airway leak | -fail to deliver tidal volume -fail to update tidal volume -fail to calculate tidal volume | -use of device issue -failure to service -parameter calculation error due to software problem
FAB-H-REQ-43 | Monitor for Leak | The Leak monitor shall be available in following modes: IPPV, SIPPV, SIMV, SIMV+PSV, PSV, CPAP, HFO | -fail to support accessory -missing symbol -inappropriate interface design | -protective measure issue -patient-device incompatibility -particulates
FAB-H-REQ-309 | Leak Specification | The value Leak shall represent the fraction of delivered volume that is lost to airway leak and not returned to the breathing circuit as expired volume, ie (Vti – Vte) / Vti. | -fail to adjust volume -incorrect volume calculation -loss of pressure | -parameter calculation error due to software problem -incorrect measurement -failure to service
FAB-H-REQ-310 | Leak display range | The Leak parameter shall be indicated within range of at least 10 to 50 %, with resolution 1 %. | -incorrect parameter -loss of indication -current out of range | -image resolution poor -protective measure issue -defective item
FAB-H-REQ-311 | Leak measurement accuracy | The Leak parameter shall have accuracy ± 10 %. | -incorrect parameter -current out of range -content out of specification | -incorrect measurement -inadequate instructions for healthcare professional -migration of device or device component
FAB-H-REQ-39 | Monitor for MV-exp | The exhaled minute volume (MV-exp) monitor shall be available in following modes: IPPV, SIPPV, SIMV, SIMV+PSV,PSV, CPAP, HFO | -incorrect volume calculation -fail to adjust volume -incorrect message display | -parameter calculation error due to software problem -patient-device incompatibility -defective component
FAB-H-REQ-312 | Exhaled minute volume (MVexp) specification | The value MVexp shall represent the exhaled minute ventilation averaged over a period of one minute. Note: this parameter may be calculated over a shorter period and extrapolated to one minute. | -incorrect value -incorrect parameter -incorrect measurement | -improper flow or infusion -inaccurate delivery -failure to pump
FAB-H-REQ-313 | MVexp display range | The MVexp parameter shall be indicated within range 0 to 9.9 L. | -incorrect parameter -current out of range -out of range | -decreased sensitivity -device difficult to setup or prepare -defective item
FAB-H-REQ-314 | MVexp measurement accuracy | The MVexp parameter shall have accuracy ± (0.1 L + 8 % of measurement). | -incorrect parameter -accuracy out of specification -low volume | -incorrect measurement -failure to convert to back-up -calibration issue
FAB-H-REQ-50 | Monitor for Mandatory %MV | The “%MV Mand” monitor shall be available in following modes: SIMV, SIMV+PSV | -fail to support interface -incorrect process -incorrect settings | -protective measure issue -device maintenance issue -patient-device incompatibility
FAB-H-REQ-315 | Share %MV mand specification | The value Share %MV mand shall represent the exhaled minute ventilation for mandatory breaths, as a fraction of total exhaled minute ventilation, averaged over a period of one minute. Note: this parameter may be calculated over a shorter period and extrapolated to one minute. | -incorrect offset value -incorrect breath type -incorrect measurement | -inaccurate delivery -fail-safe mechanism issue -failure to pump
FAB-H-REQ-316 | Share %MV Mand display range | The Share %MV Mand parameter shall be indicated within range 0 to 100 %. | -incorrect parameter -loss of indication -current out of range | -decreased sensitivity -incorrect software programming calculations -no display or display failure
FAB-H-REQ-44 | Monitor for O2 | The O2 monitor shall be available in following modes: IPPV, SIPPV, SIMV, SIMV+PSV, PSV, CPAP, nCPAP, DuoPAP, O2 Therapy, HFO | -inadequate user interface -fail to support accessory -incorrect screen display | -patient-device incompatibility -device maintenance issue -no display or display failure
FAB-H-REQ-604 | O2 monitor disable | It shall be possible for the user to disable the O2 Monitor and Alarm functions. Note this is intended to be used when the O2 sensor has reported failed calibration (e.g. due to error or end of life). | -fail to disable alarm -disable failed -fail to disable control | -failure to calibrate -incorrect measurement -use of device issue
FAB-H-REQ-64 | O2 measurement | The ventilator shall provide measurement of volume fraction of oxygen at the outlet. | -incorrect volume calculation -volume out of specification -low volume | -calibration issue -failure to service -incorrect measurement
FAB-H-REQ-65 | O2 display range | The O2 measurement shall cover at least the range 18 to 100 %. | -inappropriate measurement unit -measurement high -out of range | -incorrect measurement -device difficult to setup or prepare -device expiration issue
FAB-H-REQ-66 | O2 measurement accuracy | O2 values shall have accuracy ±(3 % absolute). | -incorrect value -low measurement accuracy -accuracy out of specification | -incorrect measurement -reset issue -dislodged or dislocated
FAB-H-REQ-45 | Monitor for Freq | The Freq monitor shall be available in following modes: IPPV, SIPPV, SIMV, SIMV+PSV, PSV, CPAP, DuoPAP | -loss of state indication -inappropriate interface design -loss of function | -protective measure issue -device maintenance issue -patient-device incompatibility
FAB-H-REQ-67 | Freq monitor | The ventilator shall provide indication of total breathing rate (including mandatory and PSV breaths) provided that the flow sensor is calibrated and connected to the patient circuit. | -loss of indication -invisible indication -incorrect message display | -calibration issue -use of device issue -incorrect measurement
FAB-H-REQ-68 | Freq monitor display range | The Freq measurement shall cover at least the range 0 to 150 bpm. | -inappropriate measurement unit -incorrect measurement -loss of measurement | -use of device issue -device difficult to setup or prepare -deployment issue
FAB-H-REQ-69 | Freq measurement accuracy | Freq values shall have accuracy ±(1 bpm). Note: the breath rate monitor is based on the proximal flow waveform and is independent of mandatory and patient trigger breath delivery. | -accuracy out of specification -incorrect value -incorrect breath type | -incorrect measurement -capturing issue -incorrect display
FAB-H-REQ-51 | Monitor for Flow | The Flow monitor shall be available in following modes: nCPAP, DuoPAP, O2 Therapy | -inadequate flow -inappropriate flow pattern -lack of ambient monitor | -defective component -particulates -protective measure issue
FAB-H-REQ-790 | Flow monitor display range | The Flow monitor shall cover at least the range 0 to 30 L/min. | -inappropriate flow pattern -insufficient flow -inadequate flow | -defective component -leak -insufficient flow or underinfusion
FAB-H-REQ-791 | Flow monitor accuracy | Flow values shall have accuracy ± (0.5 L/min + 8%) | -incorrect value -rotation accuracy out of range -accuracy out of specification | -incorrect measurement -Inflation Issue -imprecision
FAB-H-REQ-605 | Monitor for Vte HFO | The Vte HFO monitor shall be available in HFO mode only. | -unable to program -inadequate user interface -inappropriate user | -defective component -material rigid or stiff -parameter calculation error due to software problem
FAB-H-REQ-792 | Vte HFO monitor display range | The Vte HFO monitor shall cover at least the range 0 to 100 mL. | -measurement range out of specification -insufficient visibility -reduced capacity | -incorrect measurement -device expiration issue -pressure issue
FAB-H-REQ-793 | Vte HFO monitor accuracy | Vte HFO values shall have accuracy ± (0.5 mL + 8%) | -low measurement accuracy -incorrect value -accuracy out of specification | -incorrect measurement -failure to convert to back-up -dose calculation error due to software problem
FAB-H-REQ-606 | Monitor for HFO Amp | The HFO Amp monitor shall be available in HFO mode only. | -unable to program -fail to update pressure -inadequate dhf | -defective component -material rigid or stiff -parameter calculation error due to software problem
FAB-H-REQ-794 | HFO Amp monitor display range | The HFO Amp monitor shall cover at least the range 0 to 120 mbar. | -insufficient brightness -inappropriate fit -insufficient flow | -device expiration issue -device maintenance issue -leak
FAB-H-REQ-795 | HFO Amp monitor accuracy | HFO Amp values shall have accuracy ± (0.5 mbar+ 4%) | -low measurement accuracy -accuracy out of specification -measurement high | -incorrect measurement -parameter calculation error due to software problem -reset issue
FAB-H-REQ-607 | Monitor for HFO Freq | The HFO Freq monitor shall be available in HFO mode only. | -unable to program -lack of use specification -incorrect display format | -defective component -material rigid or stiff -parameter calculation error due to software problem
FAB-H-REQ-796 | HFO Freq monitor display range | The HFO Freq monitor shall cover at least the range 5 to 20 Hz. | -out of range -loss of function -inadequate size | -defective component -leak -device maintenance issue
FAB-H-REQ-797 | HFO Freq monitor accuracy | HFO Amp values shall have accuracy ± (0.5 Hz) | -low measurement accuracy -accuracy out of specification -low accuracy rate | -incorrect measurement -parameter calculation error due to software problem -reset issue
FAB-H-REQ-608 | Monitor for DCO2 | The “DCO2” monitor calculated as (Vte HFO)2 x (HFO Freq) shall be available in HFO mode only. Note: Vte in mL and Freq in Hz results in units for DCO2 of mL2/s | -incorrect display format -inadequate evaluation -inadequate display message | -defective item -parameter calculation error due to software problem -solder joint failure
FAB-H-REQ-798 | DCO2 monitor display range | The DCO2 monitor shall cover at least the range 0 to 10,000 mL2/s Note: Accuracy is not specified for DCO2 as this is derived from Vte HFO and HFO Freq. | -out of range -accuracy out of specification -reduced capacity | -incorrect measurement -use of device issue -deployment issue
FAB-H-REQ-317 | Measurement of dynamic compliance | The ventilator shall provide measurement of dynamic compliance. | -loss of measurement -inaccurate measurement -measurement high | -incorrect measurement -inadequate user interface -use of device issue
FAB-H-REQ-46 | Monitor for Compliance, C | The Compliance monitor (C) shall be available in following modes: IPPV, SIPPV, SIMV, SIMV+PSV, PSV | -fail to follow standard -inadequate verification strategy -loss of performance | -protective measure issue -environmental control or utility issue -device maintenance issue
FAB-H-REQ-318 | Compliance display range | C shall be displayed within range 0 to 50 mL/mbar. | -inappropriate display size -inadequate message display -out of range | -electronic property issue -defective item -no display or display failure
FAB-H-REQ-319 | Compliance measurement accuracy | C shall have accuracy ±(8 % of measurement). | -low measurement accuracy -accuracy out of specification -low accuracy | -incorrect measurement -temperature issue -dislodged or dislocated
FAB-H-REQ-320 | Measurement of airway resistance | The ventilator shall provide measurement of airway resistance. | -measurement high -incorrect measurement -loss of measurement | -incorrect measurement -calibration issue -loss of data
FAB-H-REQ-48 | Monitor for Resistance, R | The Resistance monitor (R) shall be available in following modes: IPPV, SIPPV, SIMV, SIMV+PSV, PSV | -resistance inadequate -inappropriate interface design -loss of state indication | -instruction for use issue -device remains activated -protective measure issue
FAB-H-REQ-321 | Resistance display range | R shall be displayed within range 0 to 200 mbar / L/s | -inappropriate display size -out of range -motion out of range | -defective item -no display or display failure -incorrect display
FAB-H-REQ-322 | Resistance measurement accuracy | R shall have accuracy ±(8 % of measurement). | -low measurement accuracy -low accuracy -inaccurate measurement | -incorrect measurement -failure to convert to back-up -difficult to interrogate
FAB-H-REQ-323 | Measurement of respiratory time constant | The ventilator shall provide measurement of respiratory time constant. | -inadequate time format -long charge time -measurement high | -loss of data -calibration issue -incorrect measurement
FAB-H-REQ-49 | Monitor for Time Constant, RC | The Time Constant monitor (RC) shall be available in following modes: IPPV, SIPPV, PSV | -fail to update time -inadequate time format -inadequate operating mode | -device maintenance issue -patient-device incompatibility -failure to deploy
FAB-H-REQ-324 | RC display range | RC shall be displayed within range 0 to 2.5 s. Note: accuracy for RC is not specified as this parameter is derived from R and C | -inappropriate display size -out of range -accuracy out of specification | -incorrect measurement -calibration issue -Failure to Calibrate
FAB-H-REQ-70 | Rapid Shallow Breathing Index (RSBI) | The ventilator shall provide measurement of Rapid Shallow Breathing Index, defined as the ratio of breath rate and tidal volume, RSBI = Freq / Vte. Note: accuracy for RSBI is not specified as this parameter is derived from Rate and Vte | -low measurement accuracy -inaccurate measurement -inaccurate rate | -incorrect measurement -use of device issue -incorrect display
FAB-H-REQ-47 | Monitor for C20/C | The C20/C monitor shall be available in following modes: IPPV, SIPPV, SIMV, SIMV+PSV, PSV | -inadequate operating mode -incorrect screen display -inappropriate interface design | -protective measure issue -device maintenance issue -patient-device incompatibility
FAB-H-REQ-325 | Over-Distension Index (C20/C) | The ventilator shall provide measurement of C20/C, defined as the compliance calculated over the last 20 % of the inspired volume, divided by the compliance calculated over the whole breath. Note: accuracy for C20/C is not specified as this parameter is derived from other values. | -inaccurate measurement -incorrect measurement -measurement high | -failure to service -incorrect measurement -improper flow or infusion
FAB-H-REQ-326 | C20/C display range | C20/C shall be displayed within range 0.0 to 5.0. | -inadequate message display -motion out of range -incorrect display format | -defective item -device difficult to setup or prepare -use of device issue
FAB-H-REQ-616 | Pulmonary Reactance (Xrs) | Measurement of pulmonary reactance shall be available when the FOT option is enabled and selected. This estimates pulmonary reactance using high-frequency pressure changes imposed onto the ventilation baseline. | -fail to measure -incorrect measurement -fail to update indicator | -programming issue -deployment issue -failure to deliver
FAB-H-REQ-617 | FOT availability | FOT shall be available in all invasive ventilation modes using a dual limb breathing circuit. | -fail to maintain ventilation mode -incorrect ventilation mode -inadequate breathing mode | -improper device output -device packaging compromised -device cleaning issue
FAB-H-REQ-618 | FOT and other functions | FOT shall be available for use together with Volume Guarantee, Volume Limit, SpO2 measuring and PRICO. Note: FOT cannot be used together with Lung recruitment in HFO ventilation mode. | -fail to adjust volume -low volume -inappropriate flow volume | -use of device issue -protective measure issue -parameter calculation error due to software problem
FAB-H-REQ-619 | FOT measurement in modes other than HFO | In conventional modes, the forced oscillation shall be created by the exhalation membrane. During the measurement, ventilation mode is temporarily changed to CPAP. | -fail to enter mode -incorrect mode -incorrect settings | -detachment of device component -mechanical issue -failure to deploy
FAB-H-REQ-620 | FOT measurement in HFO ventilation mode | In HFO mode, the forced oscillation shall be created by the HFO module. During the measurement, HFO is temporarily paused. | -inappropriate pause period -fail to shut off -inappropriate transition point | -no display or display failure -failure to deploy -incorrect software programming calculations
FAB-H-REQ-71 | Measurement of pulmonary reactance | The ventilator shall provide measurement of Xrs, defined as the energy storage capacity of the respiratory system, determined from both elastic and inertive properties, as an option when using a dual-limb breathing circuit configuration only. Note: The algorithm used to calculate Xrs is to be defined in Software Requirements. | -loss of measurement -fail to measure -fail to activate energy | -device misassembled during manufacturing or shipping -delivery system failure -calibration issue
FAB-H-REQ-537 | Control of pulmonary reactance measurement | The pulmonary reactance measurement requires a period of CPAP (to allow lung pressure to stabilize), followed by reactance measurement, prior to resumption of the prior ventilation. The ventilator shall initiate the measurement sequence only in response to a direct user interaction. | -fail to measure -incorrect measurement -inappropriate measurement unit | -pressure issue -inaccurate delivery -calibration issue
FAB-H-REQ-72 | Pulmonary reactance display range | Xrs shall be displayed within range -500 to +100 mbar / L/s. | -inappropriate display size -inadequate message display -out of range | -defective item -incorrect display -positioning issue
FAB-H-REQ-73 | Pulmonary reactance at PEEP / Pmean steps | The ventilator shall provide for measurement of Xrs at a sequence of PEEP or Pmean settings, with PEEP / Pmean incremented with a preset number of steps each with a preset incremental pressure. | -measurement high -low measurement accuracy -fail to measure | -device difficult to setup or prepare -pressure issue -compliance issue
FAB-H-REQ-74 | Xrs / Pressure graph | The ventilator shall provide display of Xrs plotted on the Y axis versus PEEP / Pmean plotted on the X axis. Note: this plot enables the clinician to determine the level of PEEP / Pmean associated with the highest (least negative) pulmonary reactance. | -incorrect plot -fail to provide indicator -unreadable display | -failure to service -compliance issue -device expiration issue
FAB-H-REQ-621 | Xrs / Pressure graph reset | The FOT Graph is cleared if any of the following actions are performed: a new FOT sequence is started, the device is restarted, the patient range is changed, device date / time is changed | -fail to clear message -fail to clear data -incorrect graph | -not audible alarm -reset issue -incorrect display
FAB-H-REQ-622 | Conditions that disable FOT | FOT is turned OFF and cannot be re-enabled while any of the following actions occurs or alarms are active: Any system alarm: Air supply pressure, Check Blender, Exhalation calibration, Oxygen supply pressure, Tube Occlusion; Any flow sensor alarm: Clean flow sensor, Flow sensor defect, Flow sensor not connected Disconnection / Tube blocked alarm: Check ET tube, Patient disconnected; Any system failure: Checksum conductor PIC, Checksum monitor PIC, COM interface, Cooling fan defect, DIO interface, I2C interface, Input pressure blender, Low physical memory, parallel interface, Safety relay defect, SPI interface, Voltage monitoring; Alarm actions: Applying Manual breath, Applying or pre-set another ventilation mode, Changing Date / time, Changing Patient range, Changing the Bias Flow setting, Changing the value of any Ventilation parameters except O2, O2 Flush and trigger sensitivity, Changing Trigger type, Changing VL or VG state, Turning OFF Flow sensor, Turning ON lung recruitment" | -fail to activate alarm -fail to activate air-in-line alarm -fail to turn on device | -device maintenance issue -deployment issue -Pressure Issue
FAB-H-REQ-623 | FOT data export for post data analysis | The system shall export the FOT settings and the measurement results (Xrs) at each pressure level via system log files | -incorrect measurement -inappropriate measurement unit -inadequate measure unit | -device maintenance issue -use of device issue -environmental control or utility issue
FAB-H-REQ-624 | Capnometry monitoring | The ventilator shall provide measurements of capnometry parameters when a capnometer device is in use. | -measurement high -loss of measurement -measure activity failed | -reset issue -inadequate user interface -failure to sense
FAB-H-REQ-75 | Capnometry Units of Measure | The ventilator shall provide for the operator to configure CO2 values to be displayed in units of mmHg, kPa, or Vol%. | -fail to display data -inadequate message display -incorrect display format | -device difficult to setup or prepare -incorrect display -failure to align
FAB-H-REQ-327 | Display of EtCO2 | The ventilator shall provide display of end-tidal CO2, EtCO2. | -inadequate message display -inadequate configuration -loss of indication | -incorrect display -loss of data -poor quality image
FAB-H-REQ-52 | Monitor for EtCO2 | The EtCO2 monitor shall be available in following modes (when capnometry module is connected and enabled): IPPV, SIPPV, SIMV, SIMV+PSV, PSV, CPAP, nCPAP, DuoPAP, O2 Therapy, HFO Note: the requirements for EtCO2 monitoring with nCPAP interface, and in HFO mode, are not clear – we may need to revise this requirement. | -inadequate connection -incompatible connection -loss of connection | -protective measure issue -device remains activated -patient data issue
FAB-H-REQ-328 | EtCO2 display range | EtCO2 shall be displayed within range 0 to 150 mmHg, or 0 to 20 kPa. | -inappropriate display size -motion out of range -out of range | -volume accuracy issue -use of device issue -device alarm system issue
FAB-H-REQ-329 | EtCO2 accuracy (with Oridion Micropod option) | EtCO2 with the Oridion micropod option shall have accuracy: ± 2 mmHg (0 to 38 mmHg); ± (5% of reading + 8% of (reading – 39 mmHg)) (39 to 150 mmHg) | -low accuracy -accuracy out of specification -inaccurate measurement | -defective item -poor quality image -human-device interface issue
FAB-H-REQ-330 | Display of FiCO2 | The ventilator shall provide display of inspired CO2, FiCO2. | -inadequate message display -fail to display data -fail to display message | -incorrect display -loss of data -poor quality image
FAB-H-REQ-331 | FiCO2 display range | FiCO2 shall be displayed within range 0 to 150 mmHg, or 0 to 20 kPa. | -inappropriate display size -motion out of range -out of range | -volume accuracy issue -use of device issue -device alarm system issue
FAB-H-REQ-332 | FiCO2 accuracy (with Oridion Micropod option) | FiCO2 with the Oridion micropod option shall have accuracy: ± 2 mmHg (0 to 38 mmHg) ± (5% of reading + 8% of (reading – 39 mmHg)) (39 to 150 mmHg) | -low accuracy -accuracy out of specification -inadequacy shape | -defective item -poor quality image -volume accuracy issue
FAB-H-REQ-333 | PCO2 waveform display | The ventilator shall provide a waveform graphic display of partial pressure of CO2. | -low pressure -fail to produce pressure -inadequate message display | -incorrect display -incorrect measurement -defective component
FAB-H-REQ-625 | Pulse oximetry monitoring | The ventilator shall provide measurements of pulse oximetry parameters when a pulse oximeter device is in use. | -measurement high -measure activity failed -measurement data corrupted | -electronic property issue -inadequate user interface -reset issue
FAB-H-REQ-54 | Monitor for SpO2 | The SpO2 monitor shall be available in following modes (when pulse oximetry module is connected and enabled): IPPV, SIPPV, SIMV, SIMV+PSV, PSV, CPAP, nCPAP, DuoPAP, O2 Therapy, HFO | -incompatible connection -loss of connection -incorrect display format | -device maintenance issue -protective measure issue -device remains activated
FAB-H-REQ-335 | SpO2 display | The ventilator shall provide display of peripheral oxygen saturation, SpO2 | -fail to update oxygen volume -low oxygen concentration -inadequate message display | -incorrect display -defective component -calibration issue
FAB-H-REQ-336 | SpO2 display range | SpO2 shall be displayed within range 1 to 100 %. | -inappropriate display size -incorrect message display -inadequate message display | -use of incorrect control settings -defective item -incorrect display
FAB-H-REQ-337 | SpO2 accuracy | SpO2 shall have accuracy specified only within the range 70 to 100%, as ± 2% pediatric patients, no motion conditions; ± 3% pediatric patients, motion conditions, and all neonatal patients. | -accuracy out of specification -low measurement accuracy -low accuracy | -incorrect software programming calculations -defective item -unintended system motion
FAB-H-REQ-338 | Pulse rate display | The ventilator shall provide display of pulse rate. | -fail to initiate pulse -inadequate message display -inadequate pulse amplitude | -incorrect display -inadequate user interface -poor quality image
FAB-H-REQ-55 | Monitor for Pulse Rate | The Pulse Rate monitor shall be available in following modes (when pulse oximetry module is connected and enabled): IPPV, SIPPV, SIMV, SIMV+PSV, PSV, CPAP, nCPAP, DuoPAP, O2 Therapy, HFO | -incorrect triggered pulse -fail to deliver pulse -incorrect display format | -protective measure issue -device remains activated -programming issue
FAB-H-REQ-339 | Pulse rate display range | Pulse Rate shall be displayed within range 25 to 240 bpm. | -incorrect message display -inappropriate display size -low rate | -defective item -electronic property issue -incorrect display
FAB-H-REQ-340 | Pulse rate accuracy | Pulse Rate shall have accuracy ± 3 bpm (no motion conditions) or ± 5 bpm (motion conditions). | -low rate -low measurement accuracy -fail to initiate pulse | -electronic property issue -use of device issue -incorrect measurement
FAB-H-REQ-341 | Perfusion index display | The ventilator shall provide display of perfusion index. | -incorrect display -inadequate message display -fail to display data | -incorrect display -loss of data -poor quality image
FAB-H-REQ-56 | Monitor for Perfusion | The Perfusion monitor shall be available in following modes (when pulse oximetry module is connected and enabled): IPPV, SIPPV, SIMV, SIMV+PSV, PSV, CPAP, nCPAP, DuoPAP, O2 Therapy, HFO | -inadequate connection -incorrect display format -inadequate display message | -device remains activated -device maintenance issue -protective measure issue
FAB-H-REQ-342 | Perfusion index display range | Perfusion Index shall be displayed within range 0.02 to 20.0 %. | -inadequate message display -incorrect display format -incorrect message display | -defective item -incorrect display -use of incorrect control settings
FAB-H-REQ-343 | Pleth waveform display | The ventilator shall provide a waveform graphic display of the pulse oximeter plethysmographic signal. | -inadequate message display -incorrect message display -loss of signal | -incorrect display -inadequate user interface -poor quality image
FAB-H-REQ-344 | SIQ graphical display | The ventilator shall provide a graphic display of SIQ. Note: SIQ is displayed as a bar indicator in the PRICO screen. | -incorrect message display -inadequate message display -incorrect display format | -incorrect display -no display or display failure -inadequate user interface
FAB-H-REQ-345 | Pulse oximeter averaging time control | The ventilator shall allow configuration of pulse oximeter averaging time with following parameters: Range of control: as 2-4, 4-6, 8, 10, 12, 14 or 16 seconds. Reset default: 10s | -fail to setup configuration -inadequate configuration -incorrect configuration | -use of incorrect control settings -use of device issue -electronic property issue
FAB-H-REQ-346 | Pulse oximeter sensitivity control | The ventilator shall allow configuration of pulse oximeter sensitivity with following parameters: Range of control: Normal, Max, or APOD (advanced probe off detection) Reset default: Normal. | -out of range -fail to setup configuration -incorrect configuration | -device difficult to setup or prepare -use of device issue -unintended movement
FAB-H-REQ-347 | Pulse oximeter fast sat control | The ventilator shall allow configuration of pulse oximeter FastSat algorithm: Range of control: On or Off. Reset default: Off | -fail to setup configuration -inadequate pulse amplitude -incorrect configuration | -inaccurate delivery -use of incorrect control settings -device difficult to setup or prepare
FAB-H-REQ-773 | Pulse oximeter alarm delay control | The ventilator shall allow configuration of pulse oximeter alarm delay time: Range of control: 0, 5, 10 or 15s. Reset default: 10s | -fail to set or reset alarm automatically -fail to setup configuration -inadequate configuration | -incorrect measurement -device difficult to setup or prepare -use of device issue
FAB-H-REQ-76 | Waveform graphics options | Waveform graphic screens shall be provided that plot pressure, flow and / or volume against time using user-configurable axis scaling | -incorrect plot -fail to adjust pressure -fail to hold pressure | -use of device issue -improper flow or infusion -programming issue
FAB-H-REQ-77 | Waveform graphics indication of patient trigger | In waveform graphic screens that show breathing activity plotted against time, patient-triggered breaths shall be identified by use of a distinct colour (green). | -confused breath type -incorrect breath type -fail to initiate breath | -defective component -improper device output -telemetry discrepancy
FAB-H-REQ-6 | Display of Loop Graphics | Loop graphic screens shall be provided that plot volume against pressure, or Flow against volume, using user-configurable axis scaling. | -incorrect plot -inaccurate volume -lack of information flow | -use of device issue -improper flow or infusion -leak
FAB-H-REQ-626 | P-V loops display | Loop plots of volume against pressure shall be displayed with pressure on the horizontal axis and inspired volume on the vertical axis | -incorrect plot -fail to display data -fail to raise pressure | -pressure issue -device markings issue -use of incorrect control settings
FAB-H-REQ-627 | F-V loops display | Loop plots of flow against volume shall be displayed with flowrate on the vertical axis and inspired volume on the horizontal axis | -incorrect plot -incorrect display format -inadequate message display | -use of device issue -device markings issue -compliance issue
FAB-H-REQ-78 | Display of trend graphs | A Trend screen shall be provided that plots measured values, averaged over a sampling period of 30 seconds. | -incorrect plot -fail to reset plot -fail to trend data or parameter | -device maintenance issue -protective measure issue -failure to disinfect
FAB-H-REQ-79 | Trend screen parameter selection | The trend screen shall provide for the following parameters to be selected for display: | -incorrect screen display -fail to provide indicator -incorrect message display | -inadequate user interface -incorrect display -data issue
FAB-H-REQ-80 | Trend screen timebase selection | The trend screen shall allow the operator to select the timebase (time axis scaling) and to scroll through available trend data. | -fail to verify time set -inaccurate trend result -incorrect screen display | -no device output -difficult to position -patient data issue
FAB-H-REQ-81 | Trend storage duration | The ventilator shall provide for up to 5 days continuous trend data storage. | -fail to manage data -lack of storage document -loss of data | -data issue -loss of data -manufacturing or shipping issue associated with device
FAB-H-REQ-82 | Audible alarm priority | The ventilator shall provide audible indication of alarm activation using the priority encoding specified in IEC 60601-1-8. | -fail to activate occlusion alarm -loss of alarm type indication -loss of indication | -no display or display failure -failure to sense -alarm not visible
FAB-H-REQ-83 | Visual alarm indication | The ventilator shall provide visual indication of alarm activation to comply with the 1m and 4m indication classes specified in IEC 60601-1-8. Note: the 1m indication is a text message indication on the GUI. The 4m indication is the coloured alarm bar on the GUI. A red lamp is provided as a backup alarm indicator for the GUI being off. | -fail to activate occlusion alarm -loss of indication -invisible indication | -no display or display failure -alarm not visible -failure to sense
FAB-H-REQ-84 | Audible alarm volume | A control shall be provided to allow the volume of the audible alarm signal to be set to a minimum of three distinct volume settings, within range 49 to 65 dB(A). | -sensor settings out of range -improper alarm -out of range | -pressure issue -use of device issue -device difficult to setup or prepare
FAB-H-REQ-7 | Alarm Audio Pause | The audio pause function shall cause the alarm audible indication to be silenced for a period not exceeding 120 s. | -fail to silence alarm -low audible alarm -loss of indication | -device alarm system issue -improper alarm -Device Alarm System Issue
FAB-H-REQ-628 | Alarm limits screen | An alarm limits screen shall be is provided to allow the operator to change the alarm limits for patient alarms applicable to the current mode of ventilation. | -incorrect limit setting -fail to change -fail to adjust manually | -device alarm system issue -use of device issue -no display or display failure
FAB-H-REQ-85 | Alarms non-latching | All alarm conditions shall be implemented as non-latching alarms. | -fail to clear alarm -lack of audible alarm -fail to generate alarm | -defective item -not audible alarm -improper alarm
FAB-H-REQ-86 | Alarm de-activation message | When an alarm deactivates, if no other alarm remains active, an informational message shall be displayed showing the last alarm active. | -fail to activate alarm -fail to activate message -incorrect message display | -device alarm system issue -Device Alarm System Issue -Device displays error message
FAB-H-REQ-87 | Alarm de-activation message clearance | The informational message showing the last alarm active shall be cleared by action of the alarm audio pause control. | -fail to clear message -fail to clear alarm -loss of message | -device alarm system issue -not audible alarm -improper alarm
FAB-H-REQ-88 | Alarm auto-set | An alarm autoset function shall be provided to allow alarm limits to be preset based on measured values. | -fail to set limit correctly -fail to set or reset alarm manually -fail to provide default setting | -no audible alarm -device alarm system issue -low audible alarm
FAB-H-REQ-89 | Indication of alarm thresholds | Alarm thresholds that are user configurable and related to a monitor value shall be indicated alongside that monitor value. | -inappropriate threshold setting -fail to identify threshold -fail to generate alarm | -output issue -data back-up problem -device maintenance issue
FAB-H-REQ-90 | Alarm prioritized indication and suppression | All alarms when active shall suppress indication of any alarm with the same, or higher, sequence value as tabulated below, with specified exceptions: “Depleted battery!” Alarm over-rides “Charge Battery <15 mins”; “Charge Battery < 30 min” overrides “Charge Battery < 60 mins” Note: High priority alarms are sequence 1 – 7; all medium priority alarms are sequence 8; all low priority alarms are sequence 9. | -fail to indicate battery depletion -confused indication -fail to activate alert | -device alarm system issue -Device Alarm System Issue -defective item
FAB-H-REQ-629 | Interaction between high and low alarm limits | It shall not be possible to set the “Low” limit for any alarm parameter to a value equal or exceeding the “High” limit for the same parameter. Note: for example, “Low Pmean” cannot be set to a value greater or equal to “High Pmean” | -incorrect limit setting -fail to set limit correctly -inappropriate threshold setting | -decrease in pressure -failure to cycle -Blockage within Device or Device Component
FAB-H-REQ-349 | Cooling fan defect alarm | The Cooling fan defect alarm shall have following parameters: Available: In all operating states; Range: N/A; Autoset: N/A; Priority: High; Sequence: 1 | -inadequate alarm states -loss of alarm -fail to activate alarm | -Device Handling Issue -defective component -break
FAB-H-REQ-350 | COM interface alarm | The COM interface alarm shall have following parameters: Available: In all operating states; Range: N/A; Autoset: N/A; Priority: High; Sequence: 1 | -inadequate alarm states -inappropriate interface design -incorrect parameter | -Device Handling Issue -defective component -device maintenance issue
FAB-H-REQ-351 | DIO interface alarm | The DIO interface alarm shall have following parameters: Available: In all operating states; Range: N/A; Autoset: N/A; Priority: High; Sequence: 1 | -inadequate alarm states -improper alarm -fail to indicate priority | -Device Handling Issue -defective component -device maintenance issue
FAB-H-REQ-352 | I2C interface alarm | The I2C interface alarm shall have following parameters: Available: In all operating states; Range: N/A; Autoset: N/A; Priority: High; Sequence: 1 | -inadequate alarm states -fail to indicate priority -incorrect parameter | -Device Handling Issue -defective component -device maintenance issue
FAB-H-REQ-353 | Parallel interface alarm | The Parallel interface alarm shall have following parameters: Available: In all operating states; Range: N/A; Autoset: N/A; Priority: High; Sequence: 1 | -inadequate alarm states -fail to generate alarm -improper alarm | -defective item -Device Handling Issue -defective component
FAB-H-REQ-354 | SPI interface alarm | The SPI interface alarm shall have following parameters: Available: In all operating states; Range: N/A; Autoset: N/A; Priority: High; Sequence: 1 | -inadequate alarm states -fail to indicate priority -improper alarm | -Device Handling Issue -defective component -device maintenance issue
FAB-H-REQ-355 | Depleted battery alarm | The “Depleted Battery!” alarm shall have following parameters: Available: In all operating states; Range: N/A; Autoset: N/A; Priority: High; Sequence: 2 Note: When the threshold for this alarm is reached, ventilation ceases and a warning symbol screen is displayed. | -fail to indicate battery depletion -fail to run on battery -incorrect parameter | -device alarm system issue -break -Device Alarm System Issue
FAB-H-REQ-356 | Charge Battery < 15 mins alarm | The Charge Battery < 15 mins alarm shall have following parameters: | -fail to indicate battery charging status -incorrect state of charge calculation -fail to run on battery | -device alarm system issue -defective component -device maintenance issue
FAB-H-REQ-357 | Checksum controller PIC alarm | The Checksum Controller PIC alarm shall have following parameters: | -fail to schedule check -incorrect measurement -inadequate checklist | -device alarm system issue -defective component -poor quality image
FAB-H-REQ-358 | Checksum monitor PIC alarm | The Checksum Monitor PIC alarm shall have following parameters | -fail to activate alarm -inadequate checklist -incorrect measurement | -defective component -poor quality image -device maintenance issue
FAB-H-REQ-359 | Low memory alarm | The Low Memory – please reboot! alarm shall have following parameters | -fail to reset alarm automatically -fail to provide low battery indicator -fail to activate alarm | -device alarm system issue -data issue -defective component
FAB-H-REQ-360 | Air supply pressure alarm | The Air supply pressure alarm shall have following parameters | -inadequate alarm states -fail to view alarm settings -fail to set or reset alarm automatically | -defective component -device maintenance issue -device alarm system issue
FAB-H-REQ-361 | Oxygen supply pressure alarm | The Oxygen supply pressure alarm shall have following parameters | -inadequate alarm states -improper alarm -fail to set or reset alarm automatically | -defective component -device alarm system issue -not audible alarm
FAB-H-REQ-362 | Safety relay defect alarm | The Safety relay defect alarm shall have following parameters | -safety interlock inadequate -lack of safety wire -inadequate safety information | -defective component -device alarm system issue -device maintenance issue
FAB-H-REQ-363 | Tube occlusion alarm | The Tube occlusion alarm shall have following parameters | -inadequate alarm states -incorrect parameter -loss of alarm | -defective component -device maintenance issue -false alarm
FAB-H-REQ-364 | Voltage monitoring alarm | The Voltage Monitoring alarm shall have following parameters. Note: the voltage monitoring alarm is activated by out of range voltages on the 5V and 12V DC rails on the powerboard and mainboard. | -fail to activate alarm -inadequate alarm states -improper alarm | -device alarm system issue -use of incorrect control settings -solder joint failure
FAB-H-REQ-365 | Battery defect alarm | The Battery Defect alarm shall have following parameters. Note: the battery defect alarm activates if the battery cannot provide sufficient current. | -fail to activate alarm -fail to provide sufficient pressure -incorrect parameter | -device alarm system issue -battery issue -misfire
FAB-H-REQ-366 | Exhalation valve calibration alarm | The Exhalation Calibration alarm shall have following parameters. Note: the exhalation calibration process is a manual process specified in the service manual. | -fail to set or reset alarm manually -fail to calibrate manually -fail to follow risk management process | -device maintenance issue -defective item -Device Handling Issue
FAB-H-REQ-367 | Check blender alarm | The Check Blender alarm shall have following parameters | -fail to activate alarm -fail to view alarm settings -fail to follow process | -device maintenance issue -device alarm system issue -defective component
FAB-H-REQ-368 | Patient disconnect alarm | The Patient Disconnect alarm shall have following parameters | -incorrect parameter -fail to set or reset alarm automatically -fail to activate alarm | -device alarm system issue -false alarm -defective component
FAB-H-REQ-369 | Check ET tube alarm | The Check ET tube alarm shall have following parameters. Note: This alarm is triggered by a total absence of flow detected by the flow sensor. | -fail to activate alarm -loss of detection -fail to disable alarm | -expulsion -defective item -excess flow or over-infusion
FAB-H-REQ-370 | Clean flow sensor alarm | The Clean flow sensor alarm shall have following parameters. | -false alarm -loss of alarm -lack of audible alarm | -defective component -particulates -device maintenance issue
FAB-H-REQ-371 | Flow sensor defect alarm | The Flow sensor defect alarm shall have following parameters | -loss of alarm -incorrect sensor orientation -fail to detect sensor error | -defective component -incorrect display -device maintenance issue
FAB-H-REQ-372 | Flow sensor not connected alarm | The Flow sensor not connected alarm shall have following parameters | -fail to generate alarm -incorrect sensor orientation -fail to test alarm | -defective component -loss of data -patient data issue
FAB-H-REQ-373 | O2 sensor calibration failed alarm | The O2 sensor calibration failed alarm shall have following parameters | -fail to calibrate manually -fail to request calibration -fail to set or reset alarm automatically | -loss of data -false alarm -inaudible voice prompts
FAB-H-REQ-374 | Oxygen sensor defect alarm | The Oxygen sensor defect alarm shall have following parameters | -fail to update oxygen volume -loss of alarm -fail to activate alarm | -defective component -device alarm system issue -incorrect display
FAB-H-REQ-375 | Oxygen sensor used up alarm | The Oxygen sensor used up alarm shall have following parameters | -loss of alarm -fail to update oxygen volume -incorrect warning light | -defective component -device alarm system issue -incorrect display
FAB-H-REQ-376 | O2 value out of range alarm | The O2 value out of range alarm shall have following parameters | -inadequate alarm states -lack of audible alarm -improper alarm | -defective item -device alarm system issue -incorrect measurement
FAB-H-REQ-630 | O2 value out of range alarm disabled when the O2 monitor is disabled | The O2 value out of range alarm shall be disabled when the O2 Monitor has been disabled by the operator. | -fail to disable alarm -disable failed -unnoticeable alarm | -incorrect measurement -defective item -not audible alarm
FAB-H-REQ-377 | SpO2 module disconnected alarm | The SpO2 module disconnected alarm shall have following parameters | -fail to disable alarm -fail to activate alarm -loss of alarm | -defective component -device alarm system issue -device maintenance issue
FAB-H-REQ-378 | SpO2 sensor failure alarm | The SpO2 sensor failure alarm shall have following parameters | -fail to indicate failure cause -incorrect sensor orientation -fail to activate alarm | -defective component -device maintenance issue -device alarm system issue
FAB-H-REQ-379 | Check SpO2 sensor alarm | The check SpO2 sensor alarm shall have following parameters | -inadequate alarm states -incorrect sensor orientation -loss of alarm | -device maintenance issue -device alarm system issue -loss of data
FAB-H-REQ-387 | Low SIQ alarm | The Low SIQ alarm shall have following parameters | -incorrect parameter -fail to view alarm settings -fail to raise medium priority alarm | -device alarm system issue -device maintenance issue -break
FAB-H-REQ-380 | EtCO2 filter line disconnected alarm | The etCO2 FilterLine disconnected alarm shall have following parameters | -loss of alarm -fail to disable alarm -incorrect parameter | -device alarm system issue -defective component -data issue
FAB-H-REQ-381 | EtCO2 module disconnected alarm | The etCO2 module disconnected alarm shall have following parameters | -incorrect parameter -loss of alarm -fail to activate alarm | -defective component -device alarm system issue -data issue
FAB-H-REQ-382 | EtCO2 sampling line occlusion alarm | The occlusion in etCO2 sampling line alarm shall have following parameters | -incorrect parameter -incorrect sample size -parameter confused | -device alarm system issue -defective component -defective item
FAB-H-REQ-383 | EtCO2 sensor fault alarm | The etCO2 sensor faulty alarm shall have following parameters | -inadequate alarm states -false alarm -incorrect warning light | -device alarm system issue -break -defective component
FAB-H-REQ-384 | Apnea alarm | The Apnea alarm shall have following parameters | -inadequate alarm states -improper alarm -fail to set or reset alarm automatically | -device alarm system issue -defective component -device maintenance issue
FAB-H-REQ-385 | High PIP alarm | The High PIP alarm shall have following parameters. Note: This alarm is the ‘High Pressure’ alarm and applies even in CPAP mode with no ventilator breaths. | -inadequate alarm states -improper alarm -loss of alarm | -device alarm system issue -Device Alarm System Issue -use of device issue
FAB-H-REQ-386 | High PEEP alarm | The High PEEP alarm shall have following parameters | -fail to view alarm settings -fail to activate alarm -fail to sort alarm message | -device alarm system issue -false alarm -device maintenance issue
FAB-H-REQ-388 | Low PIP alarm | The Low PIP alarm shall have following parameters | -inadequate alarm states -fail to view alarm settings -incorrect parameter | -device alarm system issue -device maintenance issue -break
FAB-H-REQ-389 | Charge battery < 30 mins alarm | The “Charge Battery < 30 mins” alarm shall have following parameters | -fail to indicate battery charging status -fail to run on battery -incorrect state of charge calculation | -defective component -device alarm system issue -data issue
FAB-H-REQ-390 | Charge battery < 60 mins alarm | The “Charge Battery < 60 mins” alarm shall have following parameters | -fail to indicate battery charging status -incorrect state of charge calculation -fail to activate alarm | -data issue -defective component -device alarm system issue
FAB-H-REQ-391 | High breath rate alarm | The “High Breath Rate” alarm shall have following parameters. Note: the “High Breath Rate” alarm is activated when the monitored “Freq” (patient breath rate) exceeds the limit. | -fail to activate alarm -incorrect breath type setting -inappropriate cycle breath time | -device alarm system issue -defective item -incorrect display
FAB-H-REQ-392 | High EtCO2 alarm | The High EtCO2 alarm shall have following parameters | -inadequate alarm states -incorrect parameter -fail to activate alarm | -device alarm system issue -false alarm -break
FAB-H-REQ-393 | Low EtCO2 alarm | The Low EtCO2 alarm shall have following parameters | -inadequate alarm states -improper alarm -fail to activate alarm | -device alarm system issue -incorrect display -device maintenance issue
FAB-H-REQ-394 | High FiCO2 alarm | The High FiCO2 alarm shall have following parameters | -fail to raise high priority alarm -inadequate alarm states -fail to follow process | -device alarm system issue -false alarm -break
FAB-H-REQ-395 | High tube leak alarm | The “High Tube Leak” alarm (available in all modes with dual limb breathing circuit) shall have following parameters | -incorrect parameter -parameter confused -fail to activate alarm | -defective component -device maintenance issue -defective item
FAB-H-REQ-396 | High Minute Volume alarm | The High Minute Volume alarm shall have following parameters | -fail to raise high priority alarm -inappropriate rated burst volume -inadequate alarm states | -device alarm system issue -false alarm -not audible alarm
FAB-H-REQ-397 | Low minute volume alarm | The Low Minute Volume alarm shall have following parameters | -inadequate alarm states -fail to sort alarm message -inappropriate rated burst volume | -device alarm system issue -not audible alarm -defective component
FAB-H-REQ-398 | Low PEEP alarm | The Low PEEP alarm shall have following parameters | -fail to view alarm settings -fail to sort alarm message -fail to activate alarm | -device alarm system issue -device maintenance issue -break
FAB-H-REQ-399 | High Pulse Rate alarm | The High Pulse Rate alarm (available when SpO2 monitoring is active) shall have following parameters | -fail to activate alarm -fail to view alarm settings -fail to disable alarm | -device alarm system issue -not audible alarm -device remains activated
FAB-H-REQ-400 | Low Pulse Rate alarm | The Low Pulse Rate alarm (available when SpO2 monitoring is active) shall have following parameters | -fail to activate alarm -fail to view alarm settings -fail to disable alarm | -device alarm system issue -not audible alarm -device remains activated
FAB-H-REQ-401 | High SpO2 alarm | The High SpO2 alarm (available when SpO2 monitoring is active) shall have following parameters | -fail to activate alarm -inadequate alarm states -improper alarm | -device alarm system issue -not audible alarm -device maintenance issue
FAB-H-REQ-402 | Low SpO2 alarm | The “Low SpO2” alarm (available when SpO2 monitoring is active) shall have following parameters | -fail to activate alarm -loss of alarm -improper alarm | -device alarm system issue -not audible alarm -device maintenance issue
FAB-H-REQ-403 | Low Perfusion Index alarm | The “Low Perfusion Index” alarm (available when SpO2 monitoring is active) shall have following parameters | -fail to activate alarm -incorrect parameter -insufficient alert type | -device alarm system issue -not audible alarm -device maintenance issue
FAB-H-REQ-404 | Volume not reached alarm | The “Volume Not Reached” alarm (available when Volume Guarantee is active) shall have following parameters | -fail to activate alarm -incorrect volume calculation -inaccurate volume | -use of device issue -not audible alarm -device alarm system issue
FAB-H-REQ-406 | Power Failure alarm | The Power Failure alarm shall have following parameters | -fail to activate alarm -fail to set or reset alarm automatically -fail to indicate failure cause | -defective component -power conditioning issue -device alarm system issue
FAB-H-REQ-407 | Tidal volume limit alarm | The Tidal Volume Limit alarm (available when Volume Limit is set) shall have following parameters. Note: the mandatory breath is terminated when the volume limit is reached. | -fail to set limit correctly -fail to reset alarm automatically -fail to disable alarm | -use of device issue -not audible alarm -device alarm system issue
FAB-H-REQ-631 | Acoustic audio alarm test failed alarm | The Acoustic audio alarm test failed alarm (with label “Please perform the acoustic audio alarm test!”) shall have following parameters. Note: this alarm is activated by operator confirmation that audio alarm indication is at fault. | -fail to activate alarm -audio mute failed -fail to generate alarm | -device alarm system issue -not audible alarm -use of device issue
FAB-H-REQ-632 | High DCO2 alarm | The “High DCO2” alarm (available in HFO mode only) shall have following parameters | -incorrect parameter -parameter confused -fail to activate alarm | -defective component -device alarm system issue -use of device issue
FAB-H-REQ-633 | Low DCO2 alarm | The “Low DCO2” alarm (available in HFO mode only) shall have following parameters | -incorrect parameter -parameter confused -fail to activate alarm | -defective component -device alarm system issue -use of device issue
FAB-H-REQ-634 | High Pmean alarm | The “High Pmean” alarm (available in HFO mode only) shall have following parameters | -incorrect parameter -fail to activate alarm -parameter confused | -defective component -device alarm system issue -use of device issue
FAB-H-REQ-635 | Low Pmean alarm | The “Low Pmean” alarm (available in HFO mode only) shall have following parameters | -incorrect parameter -fail to activate alarm -parameter confused | -defective component -device alarm system issue -use of device issue
FAB-H-REQ-408 | Alarm site indication for pneumatic faults | In the event of a pneumatic fault (oxygen supply pressure, air supply pressure, blender fault, breathing circuit tube occlusion, prox line occlusion, or exhalation valve fault) the site of the fault shall be indicated graphically on the user interface annotated onto a pneumatic diagram. | -incorrect fault status -loss of pressure -valve stuck | -unintended movement -defective item -leak
FAB-H-REQ-409 | Low Inlet Pressure Alarm | The ventilator shall switch to using the alternative gas in the event of one supply falling below the low inlet pressure alarm limit. Note: if the oxygen supply fails, the ventilator will operate on air inlet alone, and vice versa. | -fail to update oxygen volume -fail to drive piston -fail to initiate infusion | -use of device issue -no display or display failure -failure to deploy
FAB-H-REQ-503 | Alarm for tube occlusion | The ventilator shall provide an alarm indication for occlusion of the inspiratory, expiratory or proximal pressure sense lines. (Applies to IPPV, SIPPV, SIMV, SIMV+PSV, PSV, CPAP and HFO modes with dual limb breathing circuit only). | -loss of alarm type indication -loss of indication -fail to raise pressure | -Device Alarm System Issue -pressure issue -Incorrect Measurement
FAB-H-REQ-504 | Outlet flow disabled during occlusion | In the event of activation of the occlusion alarm, the ventilator shall terminate outlet flow for a period of at least 1s, prior to attempting to resume ventilation. | -fail to terminate process completely -fail to prevent flow -fail to disable ventilation | -Fracture -protective measure issue -missing test results
FAB-H-REQ-410 | Check Blender alarm fault code | In the event of a blender fault, the “Check Blender” alarm message shall be accompanied by a code indicating the nature of the fault. | -fail to indicate failure cause -loss of message -fail to detect error | -Device displays error message -no display or display failure -device maintenance issue
FAB-H-REQ-505 | Dual oxygen sensor | The ventilator shall use a dual oxygen sensor to implement an “O2 Cell Defect” alarm in the event of failure of the oxygen sensor. Note: this allows detection of sensor fault based on comparison of the voltages from two independent measurement cells. | -fail to detect sensor error -fail to detect error -loss of detection | -Pressure Issue -calibration issue -defective component
FAB-H-REQ-411 | Flow Sensor Faults | On activation of the flow sensor fault alarm, the operator shall be provided means to either re-calibrate, or to disable, the proximal flow sensor. | -fail to activate alarm -fail to disable alarm -fail to recover sensor | -difficult to open or close -Pressure Issue -imprecision
FAB-H-REQ-412 | Depleted Battery Alarm before shutdown | The ventilator shall shut down no sooner than one minute following activation of the 'Depleted battery' alarm. | -fail to activate alarm -fail to reset alarm automatically -fail to disable alarm | -improper device output -device remains activated -patient-device incompatibility
FAB-H-REQ-521 | Switch to internal battery | The ventilator shall switch to using the internal battery without interruption to ventilation on activation of the Power Failure alarm. | -fail to activate alarm -fail to activate air-in-line alarm -fail to run on battery | -Connection Issue -unintended movement -temperature issue
FAB-H-REQ-91 | Capnometry Faults | The ventilator shall provide High Priority alarm indications for the following capnometer fault conditions: etCO2 module disconnected etCO2 filter line disconnected Occlusion in etCO2 sampling line etCO2 sensor faulty | -fail to indicate priority -loss of alarm type indication -fail to generate alarm | -device alarm system issue -incorrect display -break
FAB-H-REQ-536 | Capnometry disabled | All capnometry measurements (numerics and graphical) shall be disabled during any EtCO2 module fault or disconnect alarm state. | -fail to disable alarm -disable failed -loss of measurement | -device alarm system issue -defective item -alarm not visible
FAB-H-REQ-92 | Oximeter Faults | The ventilator shall provide alarm indications for the following oximeter fault conditions (sub-bullets are additional informational messages, not discrete alarm states) | -fail to generate indication -loss of indication -fail to indicate failure cause | -device alarm system issue -Device Alarm System Issue -not audible alarm
FAB-H-REQ-93 | Alarm System Fault | The ventilator shall provide a technical alarm indication (using a red alarm lamp) in the event of any system fault that prevents the display of IEC 60601-1-8 compliant visual alarm indication (for example, failure of the NetDCU processor, alarm PIC, or their power supply rails). | -fail to generate indication -loss of indication -invisible indication | -no display or display failure -alarm not visible -Device Alarm System Issue
FAB-H-REQ-94 | Logging of indicated alarms | All indicated alarms (excluding the Alarm System Fault alarm) are recorded into an alarm log file, together with the alarm priority and a timestamp. Note: alarms that are suppressed by a higher priority alarm, and hence not indicated, are also not logged. | -log file overwritten -insufficient log space -inadequate log events | -device alarm system issue -Device Alarm System Issue -Device displays error message
FAB-H-REQ-95 | Alarm log review | The system shall provide capability to review the content of the alarm log. | -inadequate log events -unable to review loop -fail to sort alarm message | -device alarm system issue -power conditioning issue -failure to align
FAB-H-REQ-96 | Alarm log storage size | The ventilator shall provide storage for up to 1,000 alarm messages. | -fail to sort alarm message -insufficient storage space -lack of storage document | -loss of data -manufacturing or shipping issue associated with device -self-activation or keying
FAB-H-REQ-97 | Alarm log preserved on power failure | The alarm log shall be preserved during power failure. Note: alarm log data prior to the most recent power cycle is stored and can be exported, but cannot be reviewed on screen. | -fail to manage data -configuration data corrupted -insufficient log space | -data issue -power conditioning issue -loss of data
FAB-H-REQ-98 | Patient data entry fields | The ventilator user interface shall incorporate a patient data screen with provision for entry of the following patient data | -inadequate user interface -fail to configure patient information -incorrect user manual | -data issue -application security issue -image resolution poor
FAB-H-REQ-99 | Patient data entry reset | The patient data screen shall incorporate means to reset all patient data to blank. | -blank screen -fail to retrieve patient treatment data -incorrect data format | -patient data issue -communication or transmission issue -data issue
FAB-H-REQ-100 | Patient data entry | The patient data screen shall incorporate means to accept data entered by the operator. | -incorrect data entered -incorrect data format -fail to transfer data | -application security issue -patient data issue -communication or transmission issue
FAB-H-REQ-101 | Deletion of stored trend data | On accepting changes to the patient data, the stored trend data shall be deleted. | -fail to discard data -data back-up failed -incorrect data format | -patient data issue -application security issue -data issue
FAB-H-REQ-103 | Patient data entry cancel | The patient data screen shall incorporate means to cancel changes to the patient data. | -undesired patient data changes -incorrect data -incorrect data format | -patient data issue -application security issue -communication or transmission issue
FAB-H-REQ-104 | Patient data entry changes | The patient data screen shall incorporate means for the operator to change the data displayed. | -lack of change control -loss of data -fail to incorporate | -patient data issue -data issue -application security issue
FAB-H-REQ-106 | Data reset on power on | The patient data reset function shall be actioned on powering on the ventilator. | -fail to set or reset alarm automatically -fail to reset -fail to auto-reset message | -application security issue -data issue -device alarm system issue
FAB-H-REQ-122 | User interface controls | The ventilator shall incorporate a user interface with a color graphical display with touch screen control, a rotary control, and dedicated keys. | -inadequate user interface -inappropriate color scheme -lack of display | -compliance issue -inadequate user interface -failure to convert to back-up
FAB-H-REQ-123 | User interface lamps | The ventilator shall have dedicated lamps to indicate the following functions: System Fault (technical alarm for Alarm System failure or power off); Battery Power (operating from internal battery); Battery Charging Status (not charging [Off]; charging [blinking]; fully charged [On]); Mains Power (operating from external AC supply) | -fail to indicate ac or battery power -fail to indicate battery charging status -loss of indication | -power conditioning issue -device-device incompatibility -alarm not visible
FAB-H-REQ-636 | System Fault alarm lamp operation | The System Fault alarm lamp shall blink red after power is turned off (until the alarm silence is actioned) and during “Alarm System Fault” alarm state. | -fail to activate alarm -fail to disable alarm -fail to turn on device | -Device Alarm System Issue -power conditioning issue -device alarm system issue
FAB-H-REQ-637 | Battery power lamp operation | The Battery Power lamp shall illuminate (yellow) when the ventilator is operating from battery power. | -loss of illumination -fail to run on battery -inadequate lighting | -battery issue -device-device incompatibility -Low Battery
FAB-H-REQ-638 | Battery charging lamp operation | The Battery charging lamp shall blink when charging the battery (whether or not the ventilator is turned on) and shall illuminate solid when the battery is fully charged. | -fail to indicate battery charging status -loss of illumination -fail to measure battery voltage | -battery issue -device-device incompatibility -misfire
FAB-H-REQ-639 | Mains power lamp operation | The Mains Power lamp shall illuminate (yellow) when the ventilator is connected to wall AC power. | -illumination transmit failed -loss of illumination -inadequate lighting | -aborted charge -Power Source Issue -device-device incompatibility
FAB-H-REQ-124 | User interface membrane keys | The ventilator shall have a set of ten dedicated membrane keys for dedicated functions. | -invalid key -loss of function -missing information | -device unsafe to use in environment -failure to convert to back-up -defective component
FAB-H-REQ-125 | User interface dedicated indication | The ventilator user interface shall incorporate dedicated indication (icons or text) for the following | -fail to incorporate -loss of indication -confused indication | -no display or display failure -inadequate user interface -incorrect display
FAB-H-REQ-507 | Primary Outlet (“Insp”) | The ventilator outlet shall be a 22 mm male / 15 mm female coaxial conical connector per ISO 5356-1:2015. | -connector mismatched -suction out of range -diameter out of specification | -defective component -compliance issue -device unsafe to use in environment
FAB-H-REQ-508 | Primary outlet obstruction detection | The ventilator shall have means to detect an obstruction at the primary outlet. | -loss of detection -fail to detect errors -invisible indication | -device sensing issue -misfocusing -invalid sensing
FAB-H-REQ-509 | Fresh Gas outlet for Single Limb breathing circuits | A secondary outlet (the 'FG' or 'Fresh Gas' outlet) shall be provided for connection to single limb circuits. | -inadequate connection -single connector mechanism failed -loss of connection | -application program issue -device alarm system issue -degraded
FAB-H-REQ-510 | Fresh Gas outlet connector | The secondary outlet shall be a hose barb connector, with nominal OD 5.3 mm, hose barb OD 6.3 mm. Note: the existing breathing circuits use 4.85 mm ID / 7.95 mm OD tubing for this port. | -tubing blocked -connector mismatched -tubing malformed | -defective component -split -material rigid or stiff
FAB-H-REQ-511 | Removable Exhalation Valve Diaphragm Holder | The exhalation valve shall incorporate a removable polycarbonate diaphragm holder that includes all parts wetted by the exhaled patient gases.Note: cleaning and disinfection requirements, materials and methods are specified in section 7.4 later. | -fail to incorporate -fail to incorporate step sequence -difficult to remove | -improper device output -device handling issue -device cleaning issue
FAB-H-REQ-512 | Exhalation Valve connection for breathing circuit | The exhalation valve diaphragm holder shall incorporate a 22 mm male / 15 mm female coaxial conical connector per ISO 5356-1:2015 for connection from the breathing circuit expiratory limb. | -fail to incorporate -wrong connector -loss of connection | -loose or intermittent connection -defective component -deployment issue
FAB-H-REQ-513 | Proximal pressure measurement | The ventilator shall incorporate a pressure sensor to monitor the pressure at the patient airway interface using a proximal pressure line. | -fail to monitor feedback -loss of sensor data -incorrect sensor orientation | -pressure issue -fluid leak -decrease in pressure
FAB-H-REQ-514 | Proximal pressure line connector | The connector for the proximal pressure line shall be a cylindrical hose connector, with nominal OD 4.0 mm. Note: the existing breathing circuits use 4.0mm ID / 6.0mm OD tubing for this port. | -connector mismatched -pin length out of specification -angle out of specification | -component missing -disconnection -Detachment of Device Component
FAB-H-REQ-515 | Proximal pressure line purge flow | The ventilator shall provide purge flow to the proximal pressure line. Note: the purge flowrate is a function of blender outlet pressure, which varies depending on ventilation mode and breathing circuit configuration, typically within the range 150 – 350 mL/min for outlet flow of 8 L/min. | -inaccurate flow rate -pressure out of range -improper flow | -device operates differently than expected -failure to deploy -electronic property issue
FAB-H-REQ-413 | Flow sensor connector | The port for connection of the proximal flow sensor shall be a uniquely keyed, proprietary, dedicated latching connector. Note: this connector is on the rear panel. | -key stuck on -connector mismatched -port mismatched | -break -compliance issue -loose or intermittent connection
FAB-H-REQ-414 | Capnometer module connector | The port for connection of a capnometer module shall be a uniquely keyed, proprietary, dedicated latching connector with shrouded contacts. Note: this connector is on the rear panel. | -key stuck on -port mismatched -leakage | -fitting problem -loose or intermittent connection -component incompatible
FAB-H-REQ-415 | Pulse oximeter module connector | The port for connection of a pulse oximeter module shall be a uniquely keyed, proprietary, dedicated latching connector. Note: this connector is on the rear panel. | -key stuck on -port mismatched -unique number violated | -fitting problem -loose or intermittent connection -Device Handling Issue
FAB-H-REQ-416 | Nebulizer / USB port connector | A USB-A socket shall be provided for connection of the nebulizer, or a FLASH memory device. Note: no other devices are qualified for connection to this port. | -loss of connection -insufficient memory size -incompatible connection | -improper or incorrect procedure or method -device alarm system issue -computer software issue
FAB-H-REQ-126 | Ventilator mass (HFO model) | The fabian ventilator (model Fabian HFO) shall have a mass not exceeding 21 kg (excluding cart and accessories). | -mass out of specification -weight out of range -out of range input | -defective component -material integrity issue -patient-device incompatibility
FAB-H-REQ-640 | Ventilator mass (HFO Light model) | The fabian ventilator (model Fabian HFO Light) shall have a mass not exceeding 15 kg (excluding cart and accessories). | -mass out of specification -weight out of specification -weight out of range | -defective component -material integrity issue -patient-device incompatibility
FAB-H-REQ-127 | Ventilator dimensions | The fabian ventilator shall have dimensions not exceeding 30 cm (w) x 37 cm (h) x 42 cm (d). | -out of dimension -diameter out of specification -density out of specification | -defective component -inadequate storage -device-device incompatibility
FAB-H-REQ-136 | RS232 port for connection to a PDMS | The ventilator shall incorporate an RS232 port for connection to a PDMS. | -loss of connection -fail to connect -port mismatched | -device markings issue -loose or intermittent connection -improper or incorrect procedure or method
FAB-H-REQ-137 | RS-232 port DC out | The RS232 port shall incorporate a 5V DC output capable of providing a minimum of 500 mA. | -incorrect output -fail to incorporate -port mismatched | -battery issue -device unsafe to use in environment -device misassembled during manufacturing or shipping
FAB-H-REQ-138 | Ethernet port | The ventilator shall incorporate an Ethernet port, with a tamper-proof cover. Note: this port is for future expansion and is not qualified for use at this release. | -port mismatched -port broken or damaged -loss of function | -user used incorrect product for intended use -use of device issue -market issue
FAB-H-REQ-139 | Intellibridge / VueLink protocol | The ventilator shall provide means to communicate with a PDMS using the RS232 port using the Philips Intellibridge/VueLink protocol. | -loss of communication -loss of state indication -loss of message | -device difficult to setup or prepare -use of device issue -electromagnetic interference (emi)
FAB-H-REQ-140 | Aculink (Proprietary protocol) | The ventilator shall provide means to communicate with a PDMS using the RS232 port using a proprietary protocol (Aculink). | -loss of communication -incorrect communication information -loss of state indication | -use of device issue -inaccurate delivery -device difficult to setup or prepare
FAB-H-REQ-142 | Nurse call port | The ventilator shall incorporate a dedicated port for connection to a Nurse Call system, with normally open and normally closed contacts, enabled during activation of any alarm. | -loss of call -fail to open circuit -fail to connect | -device unsafe to use in environment -Failure to adhere or bond -Difficult to Position
FAB-H-REQ-143 | Nurse call switching capability | The Nurse Call connection shall be compatible with switching up to 1A DC or AC, at signalling voltages up to 30 VDC. | -ups switch failed -power on or off switch failed -fail to manage switch | -loose or intermittent connection -fitting problem -blocked connection
FAB-H-REQ-14 | USB memory device interface | The ventilator shall support the use of a FLASH memory device connected to the USB port, for export of trend data, alarm logs, and system information. Note: This is only indicated for use when the ventilator is not connected to a patient, but this is not restricted by device software. | -fail to store data -fail to test memory address -data corrupted | -device operates differently than expected -use of device issue -failure to sense
FAB-H-REQ-15 | Video output | The ventilator shall provide an HDMI video output, labelled as not for clinical use. | -inadequate video output channel -loss of video output -lack of audio output | -inadequate user interface -use of device issue -failure to conduct
FAB-H-REQ-149 | Patient size configuration | The ventilator shall provide means for the user to specify patient size as either neonatal or pediatric. Note: this configuration is implemented on the calibration screen – the flow sensor used must correspond to the patient size setting. | -fail to configure patient information -incorrect sample size -fail to adjust size | -user used incorrect product for intended use -application security issue -incorrect display
FAB-H-REQ-641 | Alarm delay setting | An “Alarm Delay” control shall be provided with following parameters. Note: The Alarm Delay specifies the duration of an inactivation period, applicable to all alarms, following reset of an active alarm and during which that alarm function is inhibited. This is distinct from the “SpO2 Alarm Delay” setting. | -fail to set or reset alarm automatically -fail to set or reset alarm manually -fail to activate alarm | -fitting problem -device alarm system issue -use of device issue
FAB-H-REQ-642 | Pressure units configuration | The ventilator shall provide means for the user to specify units of measure for airway pressure parameters. Note: This configuration changes the label for units, however the device does not correct for the ~2% difference between values expressed in mbar or cmH2O. To maintain accuracy specifications, the calibration procedures for inspiratory and proximal pressure transducers specified in the Service Manual need to be executed using the units of measure intended to be used. | -inadequate measure unit -incorrect measurement -lack of configuration item | -incorrect measurement -use of device issue -incorrect display
FAB-H-REQ-643 | Ppsv configuration | The ventilator shall provide means for the user to specify Ppsv as either an absolute value (default), or as an increment above set PEEP. | -incorrect settings -inadequate message display -inadequate evaluation | -device markings issue -electromagnetic interference (emi) -use of device issue
FAB-H-REQ-150 | Capnometry units configuration | The ventilator shall provide means for the user to specify units of measure for capnometry values | -inadequate measure unit -incorrect measurement -loss of measurement | -incorrect measurement -loss of data -capturing issue
FAB-H-REQ-151 | Local time and date configuration | The ventilator shall provide means for the user to specify the local time and date. | -inappropriate time format -lack of start time -lack of delivery date | -use of device issue -programming issue -inadequate user interface
FAB-H-REQ-152 | Altitude at the site of use | The ventilator shall provide means for the user to specify the altitude at the site of use, to allow for barometric compensation for controlled and measured values | -fail to expel air -fail to adjust volume -fail to monitor acceleration | -incorrect measurement -device difficult to setup or prepare -excess flow or over-infusion
FAB-H-REQ-153 | Manual breath maximum duration | The ventilator shall provide means for the user to specify the maximum duration for a manual breath, in the range 2 to 30 s, with accuracy ± 1s. Reset default: 2 s | -fail to provide default setting -fail to display default setting -incorrect breath type | -failure to align -device difficult to setup or prepare -device alarm system issue
FAB-H-REQ-154 | O2 Flush maximum duration | The “Flush Time” control shall have the following parameters | -inappropriate control flow -incorrect time and date -inappropriate time format | -defective component -inadequate storage -inadequate user interface
FAB-H-REQ-644 | Automatic O2 calibration configuration | The ventilator shall provide means for the automatic calibration of the oxygen sensor to be either single point calibration (21 %) or two point calibration (21 % and 100 %) using a control with the following parameters. Note: Two point calibration requires connection of a medical oxygen supply, with %O2 > 99% (not oxygen93). | -fail to calibrate manually -fail to request calibration -fail to reset sensor | -medical gas supply problem -inaccurate delivery -calibration issue
FAB-H-REQ-645 | Timing paradigm | The ventilator shall provide means for the timing of mandatory breaths to be configured as Set as Frequency and I-Time; Set as I-Time and E-Time. using a control with the following parameters: | -incorrect breath type setting -inappropriate cycle breath time -incorrect breath period | -use of device issue -electronic property issue -device difficult to setup or prepare
FAB-H-REQ-646 | Use BTB for Vt | The ventilator shall provide means for the tidal volume monitors to be calculated as BTB (breath to breath) enabled or disabled using a control with the following parameters | -fail to calculate tidal volume -fail to update tidal volume -fail to deliver tidal volume | -use of device issue -device difficult to setup or prepare -electronic property issue
FAB-H-REQ-647 | Manual breath in HFO configuration | The ventilator shall provide means for the manual breath function in HFO mode to be configured as enabled, or disabled, using a control with the following parameters | -fail to configure infusion parameters -fail to maintain ventilation mode -fail to setup configuration | -device difficult to setup or prepare -electromagnetic interference (emi) -data issue
FAB-H-REQ-648 | Bias flow configuration | The ventilator shall provide means for the breathing gas to be provided via the main outlet (“Internal”) or via the FG outlet (“External”) using a control with the following parameters. Note: in nCPAP, DuoPAP and O2 Therapy modes the breathing gas is always provided via the FG outlet. This configuration is intended for use with inhaled nitric oxide systems, which require unidirectional flow through the metering accessory. | -inadequate ventilation -fail to control flow -fail to deliver energy | -failure to deploy -use of device issue -unintended movement
FAB-H-REQ-649 | Language configuration | The ventilator shall provide means to configure the software user interface language as one of | -inadequate software configuration -lack of user interface evaluation plan -fail to upgrade software | -device difficult to setup or prepare -inadequate user interface -electronic property issue
FAB-H-REQ-16 | Hospital Settings
FAB-H-REQ-506 | Alarm preset compliance to standards | Restoring controls from Hospital Settings shall not result in the following restrictions being exceeded: Low SpO2 < 80 %; Low FiO2 < 18 % | -fail to retain settings -inappropriate interval settings -inappropriate environment | -device remains activated -patient data issue -device maintenance issue
FAB-H-REQ-8 | Specification for the environment of use | The ventilator shall conform to specifications when used within the following environmental conditions: Temperature +10 to +40 °C, Relative Humidity 10 to 90%, non-condensing, Ambient pressure 70 to 106 kPa (10.2 to 15.0 PSIA), Note: the ambient pressure range is intended to equate to approximately -1350 to +10,000 feet altitude. | -inappropriate environment condition -temperature out of range -lack of use environment | -environmental control or utility issue -electronic property issue -loss of or failure to bond
FAB-H-REQ-107 | Specification for shipping and transport conditions | The ventilator within its shipping package shall conform to specifications following shipping and transport for up to 72h within the following environmental conditions: Temperature -20 to +60 °C, Relative Humidity 10 to 95%, non-condensing, Ambient pressure 50 to 106 kPa (7.2 to 15.0 PSIA) | -inappropriate shipping environment or condition -inappropriate storage environment or condition -inappropriate environment condition | -communication or transmission issue -no device output -date-related software issue
FAB-H-REQ-650 | Specification for storage conditions | The ventilator within its shipping package shall conform to specifications following storage within the following environmental conditions: Temperature 0 to +40 °C, Relative Humidity 10 to 90%, non-condensing, Ambient pressure 70 to 106 kPa (10.2 to 15.0 PSIA) | -inappropriate storage environment or condition -inappropriate shipping environment or condition -out of specification | -environmental control or utility issue -communication or transmission issue -temperature issue
FAB-H-REQ-108 | Shipping carton specification | Each ventilator shall be provided with a shipping carton that meets the requirements for shipping and transport testing (FAB-H-REQ-110) and package labelling (see section 7.3.1). | -shipping test failed -missing test result -missing warning label | -compliance issue -use of device issue -fail-safe design failure
FAB-H-REQ-110 | Shipping and transport testing | The ventilator in its intended shipping carton or packaging shall conform to the required tested by ASTM D4169 (Distribution Cycle DC13, Assurance Level 1) or to ISTA 2A. | -shipping test failed -package broken or damaged -shipping damage | -packaging issue -failure to charge -failure to cycle
FAB-H-REQ-111 | Gas supply rated pressure range | The ventilator shall conform to specifications when used with gas supplies within the following range of inlet pressure: Inlet pressure 200 to 600 kPa (29 to 87 PSIG) | -pressure out of range -out of specification -out of range | -pressure issue -device alarm system issue -calibration issue
FAB-H-REQ-112 | Gas inlet particulate filters | The ventilator shall be provided with a particulate filter at each gas inlet. Note: the inlet particulate filters provide protection for the valves and blender flow sensor from contamination. | -inadequate filter -missing filter -fail to provide indicator | -fluid leak -use of device issue -failure to reset
FAB-H-REQ-113 | Gas inlet pressure regulators | Pressure regulators at each gas inlet shall be rated to at least 10 bar inlet pressure. | -fail to update pressure -over pressure -inappropriate classification | -leak -Leak -detachment of device or device component
FAB-H-REQ-114 | Gas inlet fittings | The gas inlets for oxygen and air shall use gas-specific NIST connectors. | -fail to update oxygen volume -fail to compatible with oxygen supply -wrong connector | -use of device issue -defective component -data back-up problem
FAB-H-REQ-115 | Gas supply maximum rated flow | The ventilator shall be designed to draw no more than 40 lpm from either inlet connection. | -less flexible -port less durable -reading out of specification | -misconnection -no device output -detachment of device or device component
FAB-H-REQ-116 | Air inlet water trap | A water trap shall be available for use at the air inlet. | -unable to use -water resistant failed -water mark | -delivered as unsterile product -split -inadequate instructions for healthcare professional
FAB-H-REQ-117 | Electrical power rating | The ventilator shall conform to specifications when used with a wall AC supply within the following rated range. | -fail to follow standard -inadequate specification -out of specification | -device maintenance issue -device alarm system issue -no display or display failure
FAB-H-REQ-118 | Power inlet connector | The wall AC inlet connector shall be a socket for a power cord with an IEC 60320 connector and a retention clip. | -connector mismatched -hole in connector -fail to retain settings | -fitting problem -loose or intermittent connection -compliance issue
FAB-H-REQ-119 | AC power status indication | The ventilator shall incorporate visual indication of the status of AC power connection (see FAB-H-REQ-639). | -fail to indicate ac power -loss of indication -invisible indication | -no display or display failure -optical issue -defective item
FAB-H-REQ-120 | Battery run time | The ventilator shall incorporate a rechargeable internal battery capable of run time of at least 2.5 hours (non-HFO) / 1 hour (HFO mode). | -fail to run on battery -internal battery short -non-rechargeable battery | -battery issue -deployment issue -defective component
FAB-H-REQ-121 | Battery recharging time | The battery shall be capable of recharging from fully discharged in no more than 5 hours. | -fail to run on battery -fail to charge battery to full -unable to access battery | -battery issue -defective component -device handling issue
FAB-H-REQ-128 | Ingress protection class | The fabian ventilator shall have ingress protection classification IP22 in accordance with IEC 60529:1989+A1:1999+A2:2013. | -inadequate ventilation -iec regulation violated -inadequate protection | -device unsafe to use in environment -defective component -electrical issue
FAB-H-REQ-129 | Electrical safety class | The fabian ventilator shall be electrical safety Class I / Internally Powered (protectively earthed when connected to wall power, internally powered when not connected to external power). | -inappropriate software safety class -inadequate safety information -loss of power | -Device-Device Incompatibility -failure to deliver energy -aborted charge
FAB-H-REQ-130 | Applied parts electrical isolation | The patient applied parts (excluding the breathing circuit) shall have isolation Type BF Note: this applies to the flow sensor, pulse oximeter and nebulizer connections. | -incorrect circuit type -improper isolation -inadequate profile | -use of device issue -electronic property issue -patient data issue
FAB-H-REQ-651 | Applied parts electrical isolation – breathing tube interface | The breathing circuit shall have isolation Type B. | -incorrect circuit type -incorrect breath type setting -inadequate breathing mode | -electronic property issue -poor quality image -parameter calculation error due to software problem
FAB-H-REQ-9 | Acoustic Emissions | Excluding audible alarm indication, the fabian ventilator shall produce sound pressure level not exceeding 52 dB(A) in normal operation. | -sound pressure level out of range -loss of audible alarm -low audible alarm | -use of device issue -delivery system failure -deployment issue
FAB-H-REQ-10 | Mechanical Pressure Relief at main outlet | The ventilator shall incorporate a mechanical pressure relief to prevent the pressure at the ventilator INSP outlet from exceeding 125 mbar. | -low pressure -fail to reduce pressure -fail to prevent flow | -calibration issue -failure to deploy -Failure of device to Self-Test
FAB-H-REQ-652 | Mechanical Pressure Relief at FG outlet | The ventilator shall incorporate a mechanical pressure relief to prevent the pressure at the ventilator FG outlet from exceeding 600 mbar. | -low pressure -fail to reduce pressure -pressure control failed | -calibration issue -Failure of device to Self-Test -failure to deploy
FAB-H-REQ-11 | Electronic Pressure Relief | Electronic pressure relief shall be provided at the ventilator INSP outlet during periods of active Tube Occlusion alarm, see section 7.1.9.2 | -fail to update pressure -fail to adjust pressure -inaccurate pressure reading | -defective component -Fracture -environmental control or utility issue
FAB-H-REQ-12 | Loss of Inlet Gas | In the event of loss of pressure at one gas inlet, such that the required flow for that gas cannot be provided, the ventilator shall switch to ventilation with the remaining inlet gas. | -fail to update pressure -fail to manage switch -switch invoked inappropriately | -use of device issue -pressure issue -no display or display failure
FAB-H-REQ-13 | System Fault state | In the event of loss of software control, the ventilator shall disable the blender flow control valves (such that flow is stopped). | -fail to disable control -loss of control -valve stuck | -protective measure issue -compliance issue -leak
FAB-H-REQ-653 | System Resistance | Total system resistance as measured at the patient connection port shall comply with following specifications: Note: 'inspiratory resistance' and 'expiratory resistance' refer to the pressure change at the patient connection port induced by an inspiratory flow or an expiratory flow respectively, taking the breathing system as a whole – there are no separate requirements for the resistance of the inspiratory limb or the expiratory limb of the breathing system. | -fail to follow standard -insufficient pressure -inner out of specification | -pressure issue -device inoperable -environmental control or utility issue
FAB-H-REQ-131 | Inspiratory resistance, pediatric | The inspiratory resistance when using a pediatric dual limb breathing circuit, at 15 lpm shall not exceed 6 mbar (i.e. 24 mbar/L/s) when measured at the patient connection port. | -undesired contact -inappropriate rate calculation -over dose | -Disconnection -failure to cycle -device stops intermittently
FAB-H-REQ-132 | Expiratory resistance, pediatric | The expiratory resistance when using a pediatric dual limb breathing circuit, at 15 lpm shall not exceed 6 mbar (i.e. 24 mbar/L/s) when measured at the patient connection port. | -high resistance -low resistancy -reduced capacity | -device stops intermittently -inappropriate prompts -Defective Alarm
FAB-H-REQ-133 | Inspiratory resistance, infant | The inspiratory resistance when using an neonatal dual limb breathing circuit, at 2.5 lpm shall not exceed 6 mbar (i.e. 144 mbar/L/s) when measured at the patient connection port. | -fail to hold inspiratory -low transmittance -fail to reduce pressure | -Decreased Sensitivity -inadequate instructions for healthcare professional -Use of Incorrect Control Settings
FAB-H-REQ-134 | Expiratory resistance, infant | The expiratory resistance when using an neonatal dual limb breathing circuit, at 2.5 lpm shall not exceed 6 mbar (i.e. 144 mbar/L/s) when measured at the patient connection port. | -high resistance -low resistancy -excessive dose | -Decreased Sensitivity -inappropriate prompts -user used incorrect product for intended use
FAB-H-REQ-523 | Self test sensors and calibration | The ventilator shall verify all sensors and their calibration as part of the start-up self test, and prevent ventilation if fault detected. | -fail to detect sensor error -fail to establish test criteria -fail to calibrate manually | -no pressure -protective measure issue -Pressure Issue
FAB-H-REQ-135 | Alarm indicators self test | The ventilator shall implement a self test at startup that allows the operator to confirm correct operation of the following. The nurse call indicator (if connected) | -fail to initiate test -fail to confirm -fail to indicate failure cause | -device expiration issue -use of device issue -no display or display failure
FAB-H-REQ-167 | Expected service life | The ventilator shall be designed for an operating life of 40,000 hours, or eight years (whichever is lesser) assuming all planned maintenance is carried out per the published schedule. | -material life short -reduced life cycle -low life cycle | -use of device issue -capturing issue -failure to reset
FAB-H-REQ-168 | Continuous use duration | The ventilator shall be designed for continuous operation for a period of at least 30 days. | -fail to retain settings -reduced service life -redundant process | -failure to service -image resolution poor -defective component
FAB-H-REQ-169 | Blender reverse flow | The blender construction shall prevent reverse gas flow from the outlet manifold into the blender. | -air flow blocked -manifold detached -infuser leakage | -Leak -leak -inadequate storage
FAB-H-REQ-170 | Electrical connection mechanical retention | All internal cable connections shall incorporate means for mechanical retention. | -loss of connection -fail to retain settings -missing symbol | -device handling issue -difficult to open or close -fitting problem
FAB-H-REQ-171 | Internal fasteners | All PCBAs and other internal components shall be mounted with vibration-proof fasteners (for example, screws with spring washers or thread-lock). | -component damaged -component incompatible -loss of connection | -Torn Material -fitting problem -compliance issue
FAB-H-REQ-172 | Cooling / extraction fan | There shall be a cooling fan mounted between the power electronics and the blender and designed to push air out of the enclosure. Note: the air flow pathways prevent oxygen leaking from the blender from reaching the power electronics. | -fail to prevent flow -fail to update oxygen volume -no fluid flow | -fluid leak -leak -Leak
FAB-H-REQ-173 | Cooling fan grille | The cooling fan outlet shall be protected with a grille that prevents access to the test finger specified in IEC 60601-1 clause 5.9.2.1. | -fail to activate test -unauthorized access -fail to support accessory | -human-device interface issue -therapeutic or diagnostic output failure -pressure issue
FAB-H-REQ-785 | Cooling fan safety marking | The cooling fan shall be a safety marked component. | -inappropriate safety factor -insufficient safety clearance -missing warning label | -component missing -material rigid or stiff -device markings issue
FAB-H-REQ-786 | Mains inlet connector | The AC inlet shall be a safety marked and fused power entry module conforming to IEC 60320. | -inadequate safety information -inappropriate safety factor -inadequate markings | -compliance issue -component missing -defective item
FAB-H-REQ-787 | Mains supply cord | The power cord shall be a safety marked component. | -inadequate markings -missing warning label -incorrect display value | -compliance issue -human-device interface issue -Difficult to Position
FAB-H-REQ-174 | Mains part insulation | All mains parts shall have dual insulation to the enclosure and ventilator electronic control circuits. | -loss of electrical insulation -inadequate insulation -lack of colored insulation | -electrical shorting -electrical issue -device unsafe to use in environment
FAB-H-REQ-175 | Power supply safety marking | The internal AC/DC power supply shall be a safety-marked component. | -power source incompatibility -loss of power -fail to identify safety | -unintended movement -failure to conduct -device handling issue
FAB-H-REQ-527 | Power board classification | The powerboard shall be classified as a high integrity component subject to purchasing controls. | -low sensitivity -incorrect power state -loss of state indication | -device maintenance issue -improper alarm -defective item
FAB-H-REQ-526 | Battery classification | The rechargeable battery shall be classified as a high integrity component subject to purchasing controls. | -fail to run on battery -battery broken or damaged -low battery | -improper alarm -device maintenance issue -data issue
FAB-H-REQ-525 | Nurse call relay classification | The nurse call relay shall be classified as a high integrity component subject to purchasing controls. | -loss of call -inappropriate user -unsecured data | -use of device issue -device maintenance issue -defective item
FAB-H-REQ-528 | Pressure regulator classification | The pressure regulators shall be classified as high integrity components subject to purchasing controls. | -lack of components with high-integrity characteristics -measurement high -inappropriate classification | -improper alarm -defective item -device cleaning issue
FAB-H-REQ-529 | Proportional valve classification | The proportional valves shall have proof pressure at least 10 bar, and be classified as high integrity components subject to purchasing controls. | -low pressure -inappropriate pressure level -inaccurate pressure reading | -defective item -detachment of device or device component -defective component
FAB-H-REQ-531 | Pressure sensor classification | The pressure sensors shall be classified as high integrity components subject to purchasing controls. | -lack of components with high-integrity characteristics -measurement high -loss of sensor data | -improper alarm -device maintenance issue -insufficient flow or underinfusion
FAB-H-REQ-532 | Membrane switch assembly classification | The membrane switch assembly shall be classified as a high integrity component subject to purchasing controls. | -fail to manage switch -incorrect labeling -lack of characteristics | -device maintenance issue -break -detachment of device or device component
FAB-H-REQ-540 | Membrane switch activation | All membrane switches shall have edge-triggered activation. | -fail to activate -tool activation failed -fail to activate control | -misconnection -biocompatibility issue -Delivery System Failure
FAB-H-REQ-176 | Over-voltage protection | The ventilator shall provide over-voltage protection to protect from power regulator failure. | -power control failed -overcurrent protection failed -loss of power | -defective component -failure to power-up -loss of power
FAB-H-REQ-519 | Isolation for SIP/SOPs | The ventilator shall provide isolation to the external SIP/SOPs (RS232, Ethernet, and Nurse Call) to meet requirements of IEC 60601-1. | -improper isolation -inadequate shielding -resistance inadequate | -installation-related problem -compliance issue -lack of maintenance documentation or guidelines
FAB-H-REQ-520 | Flow sensor technology | The blender flow sensors shall be hot-wire type sensors. Note: hot wire flow sensors inherently measure mass flow, and are insensitive to sample gas temperature effects. | -inaccurate sensor measurement data -inaccurate flow rate -inaccurate measurement | -insufficient flow or underinfusion -device markings issue -use of device issue
FAB-H-REQ-192 | Biocompatibility evaluation | All parts of the fabian ventilator that contact the patient, patient cells or fluids, the user, or the breathing gas shall be evaluated to ISO 10993-1 and ISO 18562-1 for biocompatibility. Note: this includes the internal dry gas pathways, the flow sensor, the exhalation valve membrane holder and exhalation valve membrane, and the ventilator enclosure. | -fail to retrieve patient diagnostics data -iso regulation violated -fail to retrieve patient treatment data | -compliance issue -inaccurate delivery -electrical shorting
FAB-H-REQ-193 | FDA biocompatibility guidance | The ventilator shall comply with the requirements of FDA guidance "Final Guidance on "Use of International Standard ISO 10993-1, Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process". Note: A Biocompatibility Evaluation Plan (BIO-RD-0017.0-EP) will be created and executed to verify compliance. | -iso regulation violated -fail to follow standard -device malfunction | -compliance issue -fail-safe design failure -failure to conduct
FAB-H-REQ-654 | Ventilator unit packaging | Unit packaging (for shipping) of the ventilator shall include the permissible environmental conditions along with the following symbols: ISO 7000-0621 symbol for 'Fragile', ISO 7000-0632 symbol for 'Temperature limit' , ISO 7000-2620 symbol for 'Humidity limit' , ISO 7000-2621 symbol for 'Atmospheric pressure limit'. | -inappropriate environment condition -missing symbol -inadequate conditions | -device markings issue -compliance issue -packaging issue
FAB-H-REQ-109 | Shipping carton labelling | The shipping carton shall be labelled to show limit on stacking height, environmental conditions for shipping and storage. | -inadequate labeling -missing warning label -adjustment height out of specification | -packaging issue -Difficult to Insert -no device output
FAB-H-REQ-655 | Legibility | All markings and labels shall be clearly legible per IEC 60601-1 clause 7.1.2. | -missing brand marking -inadequate labeling -missing warning label | -improper or incorrect procedure or method -battery issue -medication error
FAB-H-REQ-656 | Durability | All markings and labels shall be sufficiently durable to remain clearly legible during the expected service life of the product per IEC 60601-1 clause 7.1.3. | -missing brand marking -inadequate labeling -lack of labeling | -device markings issue -material integrity issue -improper or incorrect procedure or method
FAB-H-REQ-518 | Healthcare Facility Use | The fabian ventilator shall be labelled as intended for hospital and institutional healthcare facility use, including during intra-hospital transport. | -inadequate labeling -incorrect labeling -confused indication | -device markings issue -particulates -incorrect display
FAB-H-REQ-522 | Use with external clinical monitoring | The ventilator shall be labelled as intended for use in settings in which vital signs monitoring is mandatory standard of care. | -inadequate labeling -lack of labeling -incorrect labeling | -use of incorrect control settings -defective component -use of device issue
FAB-H-REQ-177 | Printed operators’ manual | Each ventilator saleable unit packaging shall include an IFU in print form, as appropriate for the Intended Market. | -missing ifu in package -inadequate ifu -missing packaging | -inadequate instructions for healthcare professional -device markings issue -instruction for use issue
FAB-H-REQ-657 | Operators Manual translations | The Operators Manual shall be available in the following languages. | -fail to support multiple languages -lack of local language -lack of application specification | -instruction for use issue -failure to deploy -inadequate instructions for healthcare professional
FAB-H-REQ-658 | Consult Accompanying Documents | The ventilator shall be marked with the ISO 7010-M002 symbol, “Refer to instruction manual / booklet” | -inappropriate symbol -inadequate markings -missing symbol | -device markings issue -compliance issue -instruction for use issue
FAB-H-REQ-179 | Marking on video output | Marking on the device shall include a safety sign that the video output is not for use on patients. | -invalid signal -missing warning label -loss of indication | -inadequate user interface -device handling issue -protective measure issue
FAB-H-REQ-180 | Marking for wall power connection | Marking on the device shall include a safety sign to connect to wall power when not in transit use. | -inadequate use safety design -loss of connection -inappropriate safety factor | -Failure to Power-Up -no flow -fitting problem
FAB-H-REQ-181 | WEEE marking | The ventilator shall be marked with the "wheelie bin" symbol denoting "Disposal as WEEE". | -inappropriate marking symbol -missing symbol -weee violated | -device markings issue -incorrect display -inadequate user interface
FAB-H-REQ-659 | Particular storage and handling | The ventilator labels shall include any particular storage and/or handling instructions. Specific instructions: labelling for ingress protection, labelling of mass of product | -inadequate labeling -lack of storage document -incorrect labeling | -use of device issue -fail-safe design failure -inadequate instructions for healthcare professional
FAB-H-REQ-660 | Particular warnings and precautions relevant to immediate operation | The ventilator labels shall include any relevant warnings and cautions. Specific warnings and cautions: do not obstruct chassis fan exhaust, read instructions (symbol) | -missing warning label -inadequate warning -fail to generate warning | -use of device issue -inadequate user interface -incorrect display
FAB-H-REQ-661 | Prescription Use Only | The ventilator shall have a label specifying prescription use only (e.g. “RX Only”). | -incorrect prescription information -inadequate labeling -lack of sickness info | -inadequate user interface -defective component -improper or incorrect procedure or method
FAB-H-REQ-662 | UDI label | The ventilator shall have a UDI label | -inadequate labeling -incorrect labeling -missing brand | -defective component -component missing -device markings issue
FAB-H-REQ-663 | Rated supply voltage range & frequency | The ventilator shall be labelled with the rated range of supply voltage, frequency, current and power. | -incorrect labeling -lack of labeling -inadequate labeling | -protective measure issue -use of device issue -compliance issue
FAB-H-REQ-664 | Ingress protection | The ventilator shall be labelled with the ingress protection classification according to IEC 60529, as IP 22 | -inadequate labeling -lack of labeling -incorrect labeling | -compliance issue -defective item -defective component
FAB-H-REQ-665 | Applied parts | Ports for connection of the following applied parts shall be labelled with the symbols for Type BF Applied Part, as defined in IEC 60417-5333: Flow Sensor, EtCO2 sensor, Pulse oximetry sensor, Nebulizer Note: this label is not required on ports intended for non-conductive pneumatic connections to the breathing circuit. | -inadequate labeling -incorrect labeling -loss of connection | -size incorrect for patient -defective item -Material Integrity Issue
FAB-H-REQ-666 | High-pressure gas inlet ports - gas name | Each of the high-pressure gas inlet ports shall be labelled with the gas name or chemical symbol “AIR” for medical nominal 4 bar, and “O2” for oxygen). | -inadequate labeling -incorrect labeling -lack of labeling | -use of device issue -pumping issue -no display or display failure
FAB-H-REQ-667 | High-pressure gas inlet ports - rated pressure range | The ventilator shall be labelled with the following rated supply pressure ranges: O2 : 200 - 600 kPa (29 - 87 psig), Air: 200 - 600 kPa (29 - 87 psig) | -inadequate labeling -incorrect labeling -fail to update pressure | -environmental control or utility issue -no device output -failure to deploy
FAB-H-REQ-668 | High-pressure gas inlet ports - rated flow rate range | The ventilator shall be labelled with the following rated flow requirement for both high-pressure gas inlets: Flow: 0 to 40 L/min. | -fail to flow check -inadequate labeling -incorrect labeling | -use of device issue -compliance issue -defective component
FAB-H-REQ-669 | Manifold - inspiratory port | The ventilator inspiratory gas outlet shall be labelled with the text, “Insp” | -inadequate labeling -incorrect labeling -lack of labeling | -failure to deploy -defective component -device-device incompatibility
FAB-H-REQ-670 | Manifold - proximal port | The ventilator proximal pressure port shall be labelled with the text, “Prox”. | -inadequate labeling -incorrect labeling -lack of labeling | -failure to deploy -pressure issue -device difficult to setup or prepare
FAB-H-REQ-671 | Manifold - expiratory port | The exhalation valve shall be marked with the text, “Exp”. | -incorrect message display -inadequate message display -invisible indication | -material rigid or stiff -component or accessory incompatibility -fitting problem
FAB-H-REQ-672 | Cooling fan exhaust | The ventilator cooling fan exhaust shall be labelled with the ISO 7000-0434 symbol, 'Caution' and the IEC 60417-5035 symbol, 'Output'. | -inadequate labeling -incorrect labeling -lack of labeling | -compliance issue -component incompatible -incompatibility problem
FAB-H-REQ-673 | Manufacturing - manufacturer info | The ventilator shall be labelled with the name or trademark and contact information for the manufacturer, together with ISO 7000-3082 symbol for 'Manufacturer'. | -lack of labeling -loss of label -incorrect labeling | -device markings issue -compliance issue -application security issue
FAB-H-REQ-674 | Manufacturing | The ventilator shall be labelled with a type reference (part number), serial number and the date of manufacture. Note: type reference may be labelled with the ISO 7000-2493 symbol; Serial number may be labelled with the ISO 7000-2498 symbol; Date of manufacture may be labelled with the ISO 7000-2497 symbol. | -missing part number -incorrect labeling -inadequate labeling | -device markings issue -use of device issue -improper or incorrect procedure or method
FAB-H-REQ-675 | NRTL certification | The ventilator label(s) shall allow space for the addition of NRTL scheme certification (i.e. MET mark) labels. | -inadequate labeling -lack of labeling -inappropriate marking symbol | -use of incorrect control settings -reset issue -no display or display failure
FAB-H-REQ-676 | Disposal - WEEE directive | The ventilator shall be labelled with reference to information for proper disposal to comply with European Directive 2002/96/EC, Waste Electrical and Electronic Equipment. | -incorrect labeling -inadequate labeling -lack of reference | -compliance issue -incompatibility problem -no device output
FAB-H-REQ-677 | Fuse holder | The ventilator shall be labelled with the type and full rating for the fuse-holders: 2 x T 1,25A. | -incorrect labeling -inadequate labeling -lack of fuse | -component incompatible -failure to deploy -use of device issue
FAB-H-REQ-678 | Rear panel - Ethernet and RS-232 ports | The Ethernet & RS-232 ports shall be labelled with the IEC 60417-5850 symbol for 'Serial Interface'. | -missing symbol -incorrect labeling -inadequate labeling | -compliance issue -component or accessory incompatibility -component incompatible
FAB-H-REQ-679 | Rear panel - nurse call connector port | The nurse call connector port shall be labelled with the IEC 60417-5307 symbol for “alarm, general”. | -missing symbol -inadequate labeling -loss of call | -compliance issue -fitting problem -component incompatible
FAB-H-REQ-680 | Rear panel - nebulizer port (text) | The nebulizer output port shall be labelled with the text, 'NEB'. | -inadequate labeling -incorrect labeling -lack of labeling | -component incompatible -failure to deploy -incorrect software programming calculations
FAB-H-REQ-681 | Rear panel - HDMI port (Video Out) | The HDMI port shall be labelled with the text, 'Video Out'. | -inadequate labeling -incorrect labeling -inadequate message display | -component incompatible -material rigid or stiff -device difficult to setup or prepare
FAB-H-REQ-682 | Rear panel - HDMI port (Service Only) | The HDMI port shall be labelled with the text, 'Service Only'. | -inadequate labeling -incorrect labeling -mislabeled | -component incompatible -material rigid or stiff -device difficult to setup or prepare
FAB-H-REQ-683 | Rear panel - equipotential connection | The equipotential connection shall be labelled with the IEC 60417-5021 symbol for 'Equipotentiality'. | -inadequate labeling -incorrect labeling -missing symbol | -device markings issue -compliance issue -incompatibility problem
FAB-H-REQ-684 | Rear panel - flow sensor port | The HWFS connector port shall be labelled with the text, 'Flow sensor'. | -inadequate labeling -incorrect labeling -lack of labeling | -component incompatible -fitting problem -material rigid or stiff
FAB-H-REQ-685 | Rear panel - pulse oximeter connector port | The connector port for pulse oximeter connection shall be labelled with the text, 'SpO2'. | -incorrect labeling -inadequate labeling -lack of labeling | -loose or intermittent connection -fitting problem -misconnection
FAB-H-REQ-686 | Rear panel - capnometer connector port | The connector port for capnometer connection shall be labelled with the text, 'EtCO2 '. | -lack of labeling -inadequate labeling -incorrect labeling | -loose or intermittent connection -fitting problem -component incompatible
FAB-H-REQ-687 | LED - Warning | The Warning LED shall be labelled with the ISO 7010-W001 symbol for 'General warning sign'. | -fail to generate warning -inadequate labeling -incorrect labeling | -device markings issue -incompatibility problem -Component Incompatible
FAB-H-REQ-688 | LED - Battery | The Battery LEDs shall be labelled with the ISO 7000-0247 symbol for “Battery charging” | -fail to indicate battery charging status -incorrect labeling -inadequate labeling | -compliance issue -device markings issue -incompatibility problem
FAB-H-REQ-689 | LED - Mains | The Mains LED shall be labelled with the ISO 60417-5534 symbol for 'Power plug'. | -inadequate labeling -incorrect labeling -missing symbol | -device markings issue -compliance issue -incompatibility problem
FAB-H-REQ-690 | Masimo branding | The ventilator shall be marked with the proprietary symbol for Masimo Signal Extraction Technology (SET®). | -loss of signal -inappropriate symbol -incorrect message display | -device markings issue -use of device issue -compliance issue
FAB-H-REQ-691 | Detachable components | The following operator-detachable components of the ventilator shall be marked with the name or trademark of the manufacturer and a model or type reference: Exhalation valve diaphragm holder, Power cord | -inappropriate symbol -incorrect display -missing logo | -data issue -compliance issue -protective measure issue
FAB-H-REQ-692 | Membrane button - Nebulizer | The Nebulizer membrane button shall be marked with the following icon: | -missing symbol -incorrect message display -unclear symbol | -sticking -user used incorrect product for intended use -inadequate instructions for healthcare professional
FAB-H-REQ-693 | Rotary encoder - consult accompanying documents | The rotary encoder shall be marked with the ISO 7010-M002 symbol for 'Refer to instruction manual/booklet'. | -inappropriate marking symbol -inadequate markings -missing symbol | -device markings issue -compliance issue -incompatibility problem
FAB-H-REQ-694 | Rotary encoder - key icon | The rotary encoder shall be labelled with a “KEY” icon | -lack of labeling -inadequate labeling -confused indication | -material rigid or stiff -incorrect software programming calculations -incompatibility problem
FAB-H-REQ-695 | Battery - consult accompanying documents | The ventilator internal battery shall be marked with the ISO 7010-M002 symbol for 'Refer to instruction manual/booklet'. | -inappropriate marking symbol -inadequate markings -incorrect logo | -compliance issue -device markings issue -device-device incompatibility
FAB-H-REQ-696 | Grounding terminal | The ground post internally to the ventilator at the AC inlet shall be labelled as a protective grounding terminal using the IEC 60417-5019 symbol. | -incorrect labeling -lack of labeling -inadequate labeling | -component missing -device difficult to setup or prepare -failure to service
FAB-H-REQ-697 | Manufacturing - manufacturer info | Accessories shall be labelled with the name or trademark and contact information for the manufacturer. | -inadequate labeling -incorrect labeling -lack of labeling | -device markings issue -particulates -medication error
FAB-H-REQ-698 | Manufacturing | Accessories shall be labelled with the ISO 7000-2492 symbol for 'Batch code' or the ISO 7000-2498 symbol for 'Serial number' with a batch or serial number, respectively, and the ISO 7000-2493 symbol for 'Catalogue number' with a type reference (part number or model designator). | -missing symbol -inadequate labeling -incorrect labeling | -device markings issue -compliance issue -therapeutic or diagnostic output failure
FAB-H-REQ-699 | Manufacturing - date of manufacture | Accessories shall be labelled with the ISO 7000-2497 symbol for 'date of manufacture', or the ISO 7000-2607 symbol for 'Use by date'. | -lack of labeling -inadequate labeling -missing symbol | -device markings issue -compliance issue -packaging issue
FAB-H-REQ-700 | UDI | The following accessory components shall each have a UDI label: Ventilator Breathing System filters, Machine filter kit, Flow sensors | -inadequate labeling -incorrect labeling -lack of labeling | -defective component -application network issue -Leak
FAB-H-REQ-701 | HWFS - flow direction | The flow sensor shall be marked with the direction of flow. | -incorrect sensor orientation -loss of sensor data -incorrect message display | -compliance issue -imprecision -inaccurate flowrate
FAB-H-REQ-702 | Accessories packaging | Packaging for all accessories shall be marked with a description of the contents. | -missing packaging -missing brand marking -inadequate markings | -device markings issue -inadequate instructions for healthcare professional -fitting problem
FAB-H-REQ-703 | Accessories packaging - UDI | All shipping packages for the ventilator and accessories shall have UDI labelling appropriate to the package and content. | -incorrect package labeling -inadequate labeling -incorrect labeling | -no device output -device markings issue -material integrity issue
FAB-H-REQ-704 | Accessories packaging - single use | Packaging for accessories intended for single-use shall be marked with the at least one of the following words: 'SINGLE USE' , 'DO NOT REUSE' , ISO 7000-1051 symbol for 'Do not reuse'. | -missing symbol -missing brand marking -missing brand | -inadequate instructions for healthcare professional -market issue -device packaging compromised
FAB-H-REQ-705 | Accessories packaging - manufacturing info | Packaging for accessories shall include the following symbols: ISO 7000-2492 symbol for 'Batch code' or ISO 7000-2498 symbol for 'Serial number' with a batch or serial number, respectively; ISO 7000-2493 symbol for 'Catalogue number’ with a type reference (part number or model designator). | -missing symbol -inappropriate marking symbol -inadequate markings | -device markings issue -computer software issue -compliance issue
FAB-H-REQ-706 | Service part packaging - battery | Packaging for a replacement battery shall comply with IATA Special Provision A199. Notes: A199 includes requirements “Any electrical battery or battery powered device having the potential of dangerous evolution of heat that is not prepared so as to prevent a short-circuit (e.g. in the case of batteries, by the effective insulation of exposed terminals; or, in the case of equipment, by disconnection of the battery and protection of exposed terminals) is forbidden from transport.” Any waybill accompanying a consignment of these batteries must contain the words “Non-restricted” and “Special Provision A199”. NiMH batteries are not within scope of UN 3481 which applies to Lithium-based secondary cells and batteries. | -insufficient package -inappropriate coating -inappropriate classification | -battery issue -device-device incompatibility -compliance issue
FAB-H-REQ-182 | IFU specification of servicing intervals | The instructions for use shall include specification of planned maintenance service intervals. | -lack of maintenance process -inadequate instruction -inadequate service | -inadequate instructions for healthcare professional -instruction for use issue -no fail-safe mechanism
FAB-H-REQ-707 | Symbols explained | The meaning of all symbols and safety signs shall be explained in the instructions for use. | -inappropriate symbol -unclear symbol -missing symbol | -device markings issue -protective measure issue -medication error
FAB-H-REQ-708 | IFU Disclosure of performance limitations | The IFU shall quantify and state any degradation in performance or accuracy specification at altitude. | -degradation of essential performance -degradation of parts -loss of function due to altitude change | -use of device issue -incorrect measurement -improper device output
FAB-H-REQ-709 | IFU disclosure of power input rating | The IFU shall disclose the rated range of supply voltage, frequency and power for the input from the supply mains. | -fail to detect power status -inadequate ifu -fail to check voltage periodically | -use of device issue -market issue -optical issue
FAB-H-REQ-710 | IFU disclosure of pressurized gas input rating | The IFU shall disclose the rated range of pressure at each high-pressure gas inlet as 200 to 600 kPa | -fail to produce pressure -fail to raise pressure -inadequate ifu | -pressure issue -improper alarm -failure to service
FAB-H-REQ-711 | IFU disclosure of pressurized gas input - compatibility | The IFU shall disclose the maximum flow rate for each gas input port as no greater than 40 L/min. Note: This addresses the requirements of ISO 80601-2-12 for both transient and time averaged flow rating, and the maximum flowrate is less than 60 L/min. | -fail to flow check -inadequate ifu -low rate | -use of device issue -device markings issue -failure to service
FAB-H-REQ-712 | Ventilator testing errors | The IFU shall include disclosure of the measurement uncertainty for each disclosed tolerance. | -inadequate ifu -lack of uid -fail to identify risk | -inadequate instructions for healthcare professional -use of device issue -protective measure issue
FAB-H-REQ-713 | Additional marking - accessories | If marking any accessory (with indication of any limitations or adverse effects on the ventilator performance) is not practicable, this information shall be placed in the IFU. | -inappropriate marking symbol -loss of indication -lack of indication | -no display or display failure -use of device issue -tidal volume fluctuations
FAB-H-REQ-714 | Additional marking - physiological effects | The IFU shall include a statement that latex is not incorporated into any materials that contact the patient gas pathway or the accessible surfaces of the device or proprietary accessories. | -incorrect incorporated -inadequate ifu -fail to configure drug information | -device handling issue -protective measure issue -electrical shorting
FAB-H-REQ-715 | Additional marking - units of measure | All disclosures of performance specification shall use SI units, except that the units listed below may be used as alternative. | -lack of use specification -incorrect specification -lack of specification | -use of device issue -improper device output -device cleaning issue
FAB-H-REQ-716 | Additional marking - units of measure | All performance specifications in the IFU relating to gas flow or volume within the ventilator breathing system shall be expressed under BTPS conditions. | -inadequate ifu -lack of application specification -fail to conduct task analysis | -use of device issue -improper device output -incorrect display
FAB-H-REQ-717 | Additional marking - units of measure | All performance specifications in the IFU relating to gas flow or volume not within the ventilator breathing system (at the gas inlets) shall be expressed under STPD20 conditions. Note: ISO 80601-2-12 requires gas condition STPD for flow and volume not related to the breathing system, but does not specify which standard temperature is to be used. STPD is STPD based on standard temperature of 20 °C. | -inadequate ifu -fail to update oxygen volume -lack of specification | -use of device issue -improper device output -failure to service
FAB-H-REQ-718 | Identification of manufacturer contact details | The IFU shall identify the equipment by including, as applicable, the following: Name or trade name of the manufacturer, Contact information to which the responsible organization can refer (importer / distributor in each registration domain), Model or type reference. | -lack of manufacturer -inadequate ifu -confused ifu | -inadequate instructions for healthcare professional -device markings issue -no display or display failure
FAB-H-REQ-719 | Identification of user and use environment | The accompanying documents shall specify any special skills, training and knowledge required of the intended operator or the responsible organization and any restrictions on locations or environments in which the ventilator equipment can be used. | -lack of training material -inadequate document -lack of training | -use of device issue -application security issue -degraded
FAB-H-REQ-720 | IFU inclusion of intended use and indications | The IFU shall document: The use of the ventilator equipment as intended by the manufacturer, The frequently used functions, Any known contraindication(s) to the use of the equipment, Intended patient population, Intended application body parts, Those parts of the equipment that shall not be serviced or maintained while in-patient use. | -inadequate ifu -fail to document method -loss of function | -inadequate instructions for healthcare professional -use of device issue -inadequate user interface
FAB-H-REQ-721 | IFU disclosure of risks of reuse of single-use parts | For any single-use component of the ventilator, its parts or accessories, the IFU shall disclose the information on known characteristics and technical factors that could pose a risk if the ventilator, its parts or accessories would be reused. | -inadequate ifu -loss of data -inappropriate indication | -no display or display failure -inadequate user interface -protective measure issue
FAB-H-REQ-722 | IFU revision control | The IFU shall disclose the date of issue or the revision of the IFU. | -inadequate ifu -lack of delivery date -fail to inform regulator on problem | -inadequate instructions for healthcare professional -use of device issue -device expiration issue
FAB-H-REQ-723 | IFU warning – protective earth | The IFU shall include a warning statement to the effect: “WARNING: To avoid the risk of electric shock, this equipment must only be connected to a supply main with protective earth.” | -lack of warning -inadequate ifu -loss of warning | -inadequate instructions for healthcare professional -fitting problem -detachment of device component
FAB-H-REQ-724 | IFU warning - interference | The IFU shall provide the operator or responsible organization with warnings regarding any significant risks of reciprocal interference posed by the presence of the ventilator during specific investigations or treatments. | -fail to generate warning -inadequate ifu -loss of warning | -invalid sensing -no display or display failure -failure to align
FAB-H-REQ-725 | IFU warning for EMC | The IFU shall include information regarding potential electromagnetic or other interference between the ventilator and other devices together with advice on ways to avoid or minimize such interference. | -interference with monitoring device -inadequate ifu -loss of data | -application security issue -use of device issue -data issue
FAB-H-REQ-726 | IFU warning for positioning | The IFU shall include a warning statement to the effect that the ventilator shall not be covered or positioned in such a way that the operation or performance of the ventilator is adversely affected, including applicable examples. | -lack of warning -inadequate ifu -loss of warning | -failure to service -use of device issue -no display or display failure
FAB-H-REQ-727 | IFU warning to have alternative means of ventilation | The IFU shall include a warning statement to the effect that, in case of ventilator failure, the lack of immediate access to appropriate alternative means of ventilation can result in patient death. | -inadequate ifu -lack of deliverable -loss of warning | -protective measure issue -unintended movement -no pacing
FAB-H-REQ-728 | IFU warning regarding accessories | The IFU shall include a warning statement to the effect that adding attachments or other components or sub-assemblies to the ventilator breathing system can change the pressure gradient across the ventilator breathing system and that such changes to the ventilator breathing system can adversely affect the ventilator performance. | -lack of warning -inadequate ifu -fail to identify change | -pressure issue -failure to service -device expiration issue
FAB-H-REQ-729 | IFU warning regarding breathing system filters | The IFU shall include a warning statement to the effect that nebulization or humidification can increase the resistance of breathing system filters and that the operator needs to monitor the breathing system filter frequently for increased resistance and blockage. | -inadequate ifu -lack of warning -loss of warning | -pressure issue -improper device output -no display or display failure
FAB-H-REQ-730 | IFU warning regarding hyperbaric chamber | The IFU shall include a warning statement to the effect that the ventilator shall not be used in a hyperbaric chamber. | -fail to generate warning -inadequate ifu -lack of warning | -no display or display failure -inadequate instructions for healthcare professional -parameter calculation error due to software problem
FAB-H-REQ-731 | IFU warning regarding gas nebulizer | The IFU shall include recommendation against the use of a gas-powered nebulizer and a warning statement to the effect that the ventilator accuracy can be affected by the gas added by use of a nebulizer. | -lack of warning -inadequate ifu -loss of warning | -Human-Device Interface Issue -detachment of device component -No Known Device Problem
FAB-H-REQ-732 | IFU warning regarding HFO limitation on battery | The IFU shall include a warning statement in the HFO section to the effect that when the device is using the internal battery, HFO amplitude is limited to 60mbar | -lack of warning -inadequate ifu -loss of warning | -no display or display failure -instruction for use issue -failure to read input signal
FAB-H-REQ-733 | IFU - warnings and safety notices | The IFU shall include a warning statement in the HFO section to the effect that during HFO ventilation the distinct recognition of disconnection and measurement of the minute volume is limited | -lack of warning -loss of warning -inadequate ifu | -no display or display failure -protective measure issue -disconnection
FAB-H-REQ-734 | IFU - warnings and safety notices | The IFU shall include a warning statement in the HFO section to the effect that the use of a heated inspiratory and expiratory tubing system is strongly recommended in HFO mode | -lack of warning -inadequate ifu -fail to generate warning | -disconnection -no display or display failure -device unsafe to use in environment
FAB-H-REQ-735 | IFU - warnings and safety notices | The IFU shall include a warning statement in the FOT section to the effect that the “Manual Breath” can be activated during FOT, but this will deactivate FOT. To re-enable FOT after a Manual Breath /Sustained Inflation, you must restart FOT from the beginning. | -lack of warning -fail to activate message -loss of warning | -no display or display failure -inadequate instructions for healthcare professional -instruction for use issue
FAB-H-REQ-736 | IFU - warnings and safety notices | The IFU shall include a warning statement in the HFO section to the effect that the Start / Stop button can be used to interrupt mechanical ventilation for two minutes, but this will deactivate FOT. | -inadequate ifu -fail to disable ventilation -fail to deactivate | -no display or display failure -break -inadequate instructions for healthcare professional
FAB-H-REQ-737 | IFU - warnings and safety notices | The IFU shall include a warning statement in the FOT section to the effect that any FOT measurements taken during Standby are NOT valid. | -lack of warning -inadequate ifu -loss of warning | -parameter calculation error due to software problem -inadequate instructions for healthcare professional -device alarm system issue
FAB-H-REQ-738 | IFU - warnings and safety notices | The IFU shall include a warning statement in the FOT section to the effect that inaccurate measurements may be caused by: Flow sensor not calibrated or defective; Incorrect pressure measurement line application or use; Excessive patient spontaneous breathing (strong active respiratory drive); | -inaccurate measurement -inaccurate delivery -loss of warning | -incorrect software programming calculations -fluid leak -improper flow or infusion
FAB-H-REQ-739 | IFU - equipment description | The IFU shall include the location of all natural-rubber-latex-based components. | -inadequate ifu -lack of position indicator -lack of configuration item | -inadequate instructions for healthcare professional -protective measure issue -instruction for use issue
FAB-H-REQ-740 | IFU - installation | The IFU shall contain: a reference to where the installation instructions are to be found; or contact information for persons designated by the manufacturer as qualified to perform the installation. | -lack of reference -inadequate ifu -lack of installation document | -device markings issue -inadequate instructions for healthcare professional -use of device issue
FAB-H-REQ-741 | IFU - isolation | The IFU shall contain an instruction not to position the ventilator so that it is difficult to disconnect the power cord from the wall outlet. | -inadequate ifu -lack of uid -version incorrect | -Power Source Issue -fail-safe design failure -electromagnetic interference (emi)
FAB-H-REQ-742 | IFU - start-up procedure | The IFU shall contain the necessary information for the operator to bring the ventilator into operation including such items as any initial control settings, connection to or positioning of the patient, etc. | -inadequate ifu -lack of help information -lack of uid | -incorrect display -data issue -inadequate instructions for healthcare professional
FAB-H-REQ-743 | IFU - start-up procedure | The IFU shall detail any treatment or handling needed before the ventilator, its parts, or accessories can be used. | -inadequate ifu -lack of procedure -unclear information | -inadequate user interface -use of device issue -degraded
FAB-H-REQ-744 | IFU - start-up procedure | The IFU shall disclose a method by which all of the alarm signals can be functionally tested to determine if they are operating correctly. Note: Portions of this test method may be automatically performed by the ventilator or may require operator action | -fail to establish test criteria -fail to bidirectional communication -fail to indicate valid position | -use of device issue -fail-safe design failure -pressure issue
FAB-H-REQ-745 | IFU - start-up procedure | The IFU shall include the following: a method of verifying all operator-adjustable alarm system functions, the time duration from start-up to providing essential performance. | -inadequate ifu -version incorrect -unable to place | -pressure issue -no display or display failure -failure to align
FAB-H-REQ-746 | IFU - operating instructions | The IFU shall include the information on characteristics and technical factors known to the manufacturer that could pose a risk if the RGM, its parts or accessories were re-used. | -inadequate ifu -lack of risk assessment -lack of risk control measure | -no display or display failure -device markings issue -protective measure issue
FAB-H-REQ-747 | IFU - operating instructions | The IFU shall include the date of issue or revision of the IFU. | -lack of delivery date -inadequate ifu -lack of expiration date | -inadequate instructions for healthcare professional -no display or display failure -use of device issue
FAB-H-REQ-748 | IFU - operating instructions | The IFU shall include a step by step description for the usage of FOT | -inadequate ifu -lack of labeling -lack of evaluation method | -inadequate instructions for healthcare professional -use of device issue -protective measure issue
FAB-H-REQ-749 | Messages | The IFU shall list all system messages, error messages and fault messages that are generated, unless these messages are self-explanatory. | -fail to display message -fail to generate indication -inadequate ifu | -device markings issue -device alarm system issue -solder joint failure
FAB-H-REQ-750 | Messages | The IFU shall include an explanation of messages including important causes, and possible action(s) by the operator, if any, that are necessary to resolve the situation indicated by the message. | -fail to display message -fail to provide indicator -inadequate ifu | -inadequate instructions for healthcare professional -device markings issue -protective measure issue
FAB-H-REQ-751 | Shutdown procedure | The IFU shall contain the necessary information for the operator to safely terminate the operation of the ventilator. | -inadequate ifu -lack of uid -missing brand | -inadequate instructions for healthcare professional -no display or display failure -use of incorrect control settings
FAB-H-REQ-752 | Cleaning, disinfection and sterilization | The IFU shall contain, for ventilator parts or accessories that can become contaminated through contact with the patient or with body fluids or expired gases during normal use and single fault condition, the following: details about cleaning and disinfection or sterilization methods that may be used, and list the applicable parameters such as temperature, pressure, humidity, time limits and number of cycles that such ventilator parts or accessories can tolerate. Note: This requirement does not apply to any material, component, accessory or ventilator parts that is marked as intended for a single-use unless the manufacturer specifies that the material, component, accessory or the ventilator is to be cleaned, disinfected, or sterilized before use. " | -inadequate ifu -missing part number -accessory missing | -inaccurate delivery -application security issue -use of device issue
FAB-H-REQ-753 | Cleaning, disinfection and sterilization | The IFU shall identify which portions of the gas pathways through the ventilator can become contaminated with body fluids or expired gases during both normal condition and single fault condition. | -fail to identify bubble -fail to identify threshold -inadequate ifu | -use of device issue -fail-safe design failure -device-device incompatibility
FAB-H-REQ-754 | Maintenance | The IFU shall instruct the operator or responsible organization in sufficient detail concerning preventive inspection, maintenance and calibration to be performed by them, including the frequency of such maintenance. | -inadequate ifu -fail to follow standard -fail to monitor feedback | -inadequate instructions for healthcare professional -use of incorrect control settings -device markings issue
FAB-H-REQ-755 | Maintenance | The IFU shall provide information for the safe performance of such routine maintenance necessary to ensure the continued safe use of the ventilator. | -lack of information for safety -fail to follow safe practices -inadequate ifu | -inadequate instructions for healthcare professional -no fail-safe mechanism -defective component
FAB-H-REQ-756 | Maintenance | The IFU shall identify the parts on which preventative inspection and maintenance shall be performed by service personnel, including the periods to be applied, but not necessarily including details about the actual performance of such maintenance. | -lack of maintenance process -fail to identify threshold -inadequate ifu | -device markings issue -inadequate instructions for healthcare professional -protective measure issue
FAB-H-REQ-757 | Maintenance | The IFU shall include the procedures for calibration before or during use. | -lack of tracking procedure -inadequate ifu -lack of test specification | -inadequate instructions for healthcare professional -device markings issue -incorrect display
FAB-H-REQ-758 | Maintenance | The IFU shall include the methods and frequency of routine inspection and testing. | -inadequate ifu -inappropriate method -inappropriate test design | -use of device issue -protective measure issue -parameter calculation error due to software problem
FAB-H-REQ-759 | Environmental protection | The IFU shall include advice on the proper disposal of calibration gases. Note: This applies to calibration gas for the capnometry functions | -inadequate ifu -fail to follow plan -lack of help information | -inadequate instructions for healthcare professional -no display or display failure -fail-safe design failure
FAB-H-REQ-760 | Reference to technical description | The IFU shall contain the information specified below. The permissible environmental conditions of use including conditions for transport and storage; All characteristics of the ventilator, including range(s), accuracy, and precision of the displayed values or an indication where they can be found; Any special installation requirements such as the maximum permissible apparent impedance of supply mains; A description of the means of isolating the ventilator from the supply mains, if such means is not incorporated in the ventilator; A warning statement that addresses the hazards that can result from unauthorized modification of the ventilator; Example: “WARNING: No modification of this equipment is allowed”; Example: “WARNING: Do not modify this equipment without authorization of the manufacturer” ; Example: “WARNING: If this equipment is modified, appropriate inspection and testing must be conducted to ensure continued safe use of the equipment.” | -inadequate ifu -lack of use environment -missing information | -inadequate instructions for healthcare professional -instruction for use issue -environmental control or utility issue
FAB-H-REQ-801 | Instructions for inspection following cleaning | The Operators Manual shall include instructions for inspection of each accessory for damage following cleaning and disinfection. | -inadequate instruction -accessory missing -lack of filtration process | -inadequate instructions for healthcare professional -instruction for use issue -protective measure issue
FAB-H-REQ-417 | Exhalation valve diaphragm processing | The exhalation valve diaphragm shall be capable of cleaning and disinfection per the instructions in the Operators Manual for no less than 30 cycles. | -failed to be disinfected -fail to clean adequately -fail to be washed | -material integrity issue -particulates -device maintenance issue
FAB-H-REQ-761 | Exhalation valve diaphragm cleaning process | The exhalation valve diaphragm shall be compatible with the following cleaning and disinfection process: Cleaning with pH neutral detergent (ValSure® Cleaner, Steris) at a temperature of 80 to 95 °C for 5 minutes; followed by Steam sterilization at 134 °C for 2 minutes (half cycle). | -incompatible solvent -component incompatible -part mismatched | -device cleaning issue -component or accessory incompatibility -material rigid or stiff
FAB-H-REQ-418 | Exhalation valve diaphragm holder cleaning cycles | The exhalation valve diaphragm holder shall be capable of cleaning and disinfection per the instructions in the Operators Manual for no less than 30 complete cleaning cycles, or for a maximum of one year (if sooner). | -failed to be disinfected -fail to clean adequately -expired cleaning agent | -mechanical issue -material integrity issue -defective item
FAB-H-REQ-419 | Exhalation valve emptying during use | The instructions for use shall include instructions for emptying the exhalation valve and diaphragm assembly of water and secretions. Note: this is carried out during use on a single patient, in order to maintain pressure performance, and is not intended as an infection control procedure. | -inadequate instruction -confused instruction -lack of procedure | -use of device issue -inadequate instructions for healthcare professional -disconnection
FAB-H-REQ-420 | Exhalation valve diaphragm holder cleaning process | The exhalation valve membrane holder shall be compatible with the following cleaning and disinfection process: Cleaning with pH neutral detergent (ValSure® Cleaner, Steris) at a temperature of 80 to 95 °C for 5 minutes; followed by Steam sterilization at 134 °C for 2 minutes (half cycle). | -incompatible solvent -component incompatible -inappropriate usage environment | -device cleaning issue -component or accessory incompatibility -material rigid or stiff
FAB-H-REQ-421 | Ventilator enclosure cleaning cycles | The ventilator enclosure shall be designed for a minimum of 400 cleaning cycles. | -fail to clean adequately -reduced washable -insufficient wetting | -defective component -device maintenance issue -accessory incompatible
FAB-H-REQ-422 | Ventilator enclosure cleaning materials | The ventilator enclosure shall be compatible with cleaning with each of the following materials, when used in accordance with manufacturers instructions: CaviWipes™ (Metrex); 10% solution of 5.25% sodium hypochlorite bleach (e.g. Clorox), used with lint-free textile cleaning cloths; Hydrogen peroxide 5% (e.g. 35% food grade H2O2, diluted 1:6 with distilled water), used with lint-free textile cleaning cloths. | -material incompatibility -incompatible solvent -component incompatible | -protective measure issue -defective item -device cleaning issue
FAB-H-REQ-423 | Flow sensor cable cleaning cycles | The flow sensor cable shall be designed for a minimum of 400 cleaning cycles. | -cleaning cycle failed -reduced washable -less air input | -defective component -material rigid or stiff -Break
FAB-H-REQ-424 | Flow sensor cable cleaning materials | The flow sensor cable shall be compatible with cleaning with each of the following materials, when used in accordance with manufacturers instructions: CaviWipes™ (Metrex); 10% solution of 5.25% sodium hypochlorite bleach (e.g. Clorox), used with lint-free textile cleaning cloths; Hydrogen peroxide 5% (e.g. 35% food grade H2O2, diluted 1:6 with distilled water), used with lint-free textile cleaning cloths. | -material incompatibility -component incompatible -inappropriate cleanser | -device cleaning issue -protective measure issue -defective item
FAB-H-REQ-799 | Flow sensor cleaning cycles | The flow sensor shall be designed for a minimum of 30 cleaning cycles. | -cleaning cycle failed -duty cycle low -expired cleaning agent | -defective component -Pressure Issue -optical issue
FAB-H-REQ-802 | Flow sensor cleaning materials | The flow sensor shall be compatible with cleaning with the following cleaning and disinfection process: Cleaning with pH neutral detergent (ValSure® Cleaner, Steris) at a temperature of 80 to 95 °C for 5 minutes; followed by Steam sterilization at 134 °C for 2 minutes (half cycle). | -sensor damaged -device damaged by inappropriate cleaning -component incompatible | -device cleaning issue -improper chemical reaction -component or accessory incompatibility
FAB-H-REQ-762 | CE marking to GSPR of EU MDR | The ventilator shall be CE marked to demonstrate conformity to the GSPR of the EU Medical Devices Regulation | -lack of ce mark -fail to comply with standards or regulation -inappropriate marking symbol | -accessory incompatible -defective component -incorrect display
FAB-H-REQ-763 | MET marking for Canada | The ventilator shall be marked with MET mark for distribution in Canada. | -inappropriate marking symbol -device marked incorrectly -missing symbol | -device markings issue -compliance issue -device packaging compromised
FAB-H-REQ-764 | Critical component marking | Off the shelf components identified as High Integrity components (Critical Components list) shall be marked with an NTRL mark. Note: This applies to nurse call relay, pressure regulators, proportional valves, pressure sensors, rechargeable battery, AC/DC convertor. It does not apply to custom parts identified as High Integrity components: powerboard PCBA, membrane switch assembly. | -inadequate markings -inappropriate marking symbol -missing parts | -use of device issue -compliance issue -device markings issue
FAB-H-REQ-183 | Compliance with EN ISO 80601-2-12:2020 | The fabian ventilator shall comply with ISO 80601-2-12:2020 ("Medical electrical equipment - Part 2-12: Particular requirements for basic safety and essential performance of critical care ventilators"). | -inadequate safety information -inappropriate safety factor -fail to comply with standards or regulation | -compliance issue -protective measure issue -fitting problem
FAB-H-REQ-184 | Compliance with EN 60601-1 ed 3.2 | The fabian ventilator shall comply with EN 60601-1:2005 + A1:2012 + A2:2020 Medical electrical equipment – Part 1: General requirements for basic safety and essential performance | -fail to validate basic safety -en regulation violated -inadequate safety information | -compliance issue -fitting problem -protective measure issue
FAB-H-REQ-185 | Compliance with EN 60601-1-2 ed 4.1 | The fabian ventilator shall comply with EN 60601-1-2:2014 + A1:2020 Medical electrical equipment –Part 1-2: General requirements for basic safety and essential performance – Collateral Standard: Electromagnetic disturbances – Requirements and tests. | -en regulation violated -inadequate safety information -incorrect electrical pathway | -compliance issue -fitting problem -use of device issue
FAB-H-REQ-186 | Compliance with EN 60601-1-6 ed 3.2 | The fabian ventilator shall comply with EN 60601-1-6:2010+A1:2013 + A2:2020 Medical electrical equipment –Part 1-6: General requirements for basic safety and essential performance – Collateral Standard: Usability. | -inadequate safety information -en regulation violated -inappropriate safety factor | -compliance issue -fitting problem -use of device issue
FAB-H-REQ-187 | Compliance with EN 60601-1-8 ed 2.2 | The fabian ventilator shall comply with EN 60601-1-8:2006+A1:2012+A2:2020 Medical electrical equipment –Part 1-8: General requirements for basic safety and essential performance – Collateral Standard: General requirements, tests and guidance for alarm systems in medical electrical equipment and medical electrical systems. | -inadequate safety information -en regulation violated -part misoriented | -compliance issue -use of device issue -fitting problem
FAB-H-REQ-765 | Compliance with EN 60601-1-9 ed 1.2 | The fabian ventilator shall comply with EN 60601-1-9:2007+A1:2013+A2:2020 Medical electrical equipment – Part 1-9: General requirements for basic safety and essential performance – Collateral Standard: Requirements for environmentally conscious design | -inadequate safety information -en regulation violated -part misoriented | -compliance issue -fitting problem -use of device issue
FAB-H_REQ-766 | Compliance with EN 60601-1-10 ed 1.2 | The Fabian ventilator when operating in PRICO function shall comply with EN 60601-1-10:2007+A1:2013+A2:2020 Medical electrical equipment – Part 1-10: General requirements for basic safety and essential performance – Collateral Standard: Requirements for the development of physiologic closed-loop controllers | -en regulation violated -inadequate safety information -inappropriate safety factor | -compliance issue -fitting problem -device unsafe to use in environment
FAB-H-REQ-188 | Compliance with EN ISO 80601-2-55 | The fabian ventilator oxygen monitoring and etCO2 monitoring functions shall comply with EN ISO 80601-2-55:2018 Medical electrical equipment -- Part 2-55: Particular requirements for the basic safety and essential performance of respiratory gas monitors. | -inadequate safety information -iso regulation violated -inappropriate safety factor | -compliance issue -computer software issue -Difficult to Position
FAB-H-REQ-189 | Compliance with EN ISO 80601-2-61 | The fabian ventilator pulse oximetry functions shall comply with EN ISO 80601-2-61:2017 Medical electrical equipment -- Part 2-61: Particular requirements for basic safety and essential performance of pulse oximeter equipment. | -iso regulation violated -incorrect triggered pulse -inadequate safety information | -compliance issue -protective measure issue -fail-safe design failure
FAB-H-REQ-190 | Compliance with EN 62304 | The fabian ventilator shall comply with EN 62304:2006+A1:2016 Medical device software – Software life cycle processes. | -reduced device life cycle -en regulation violated -fail to comply with standards or regulation | -compliance issue -application program version or upgrade problem -material disintegration
FAB-H-REQ-191 | Compliance with EN 62366-1 | The fabian ventilator shall comply with EN 62366-1:2015 Medical devices – Application of usability engineering to medical devices. | -device alarmed -device damaged -device malfunction | -use of device issue -application network issue -device cleaning issue
FAB-H-REQ-767 | Compliance with ISO 80601-2-87:2021 | The fabian ventilator when operating in HFO mode shall comply with ISO 80601-2-87:2021 Medical electrical equipment - Part 2-87: Particular requirements for basic safety and essential performance of high frequency ventilators. | -inappropriate safety factor -inadequate safety information -iso regulation violated | -compliance issue -use of device issue -device unsafe to use in environment
FAB-H-REQ-768 | Compliance with ISO FDIS 80601-2-90:2021 | The fabian ventilator when operating in O2 Therapy mode shall comply with ISO FDIS 80601-2-90:2021 Medical electrical equipment - Part 2-87: Particular requirements for basic safety and essential performance of high flow oxygen therapy equipment. | -iso regulation violated -inappropriate safety factor -inadequate safety information | -compliance issue -device unsafe to use in environment -computer software issue
FAB-H-REQ-769 | ISTA 2A Shipping Test | The fabian ventilator when packaged in the shipping packaging shall comply with ISTA 2A | -incorrect package labeling -package compromised -package broken or damaged | -device packaging compromised -packaging issue -sticking
FAB-H-REQ-17 | Factory configuration | The ventilator shall provide means to enter license codes to enable and / or disable the following options | -invalid access code -unsecured data -fail to insert | -installation-related problem -use of device issue -electronic property issue
FAB-H-REQ-781 | Non-slip feet for shelf use | The ventilator shall incorporate non-slip feet and a flat base to allow for use placed on a shelf. | -slip coating -lack of keyboard -incorrect configuration | -Device Inoperable -human factors issue -inaccurate delivery
FAB-H-REQ-789 | Mounting mechanism | The ventilator shall incorporate a mounting mechanism to allow for use mounted to a cart, pedestal or pendant | -fail to incorporate -inappropriate mount -fail to connect | -inaccurate delivery -manufacturing or shipping issue associated with device -device maintenance issue
FAB-H-REQ-144 | User calibration for the prox flow sensor | The ventilator shall provide means for the user to calibrate the proximal flow sensor (drift correction). | -calibration failed -inaccurate sensor measurement data -inaccurate measurement | -failure to calibrate -Failure to Calibrate -incorrect measurement
FAB-H-REQ-145 | Prox flow sensor calibration prompt | The ventilator shall provide a prompt to the operator to calibrate the proximal flow sensor at startup (if it is connected) and on each occasion that a proximal flow sensor is connected to the ventilator. | -fail to request calibration -fail to detect sensor error -prompt failed | -Failure to Calibrate -excess flow or over-infusion -failure to calibrate
FAB-H-REQ-146 | Prox flow sensor disconnect notification | The ventilator shall provide a notification on each occasion that a proximal flow sensor is disconnected from the ventilator. | -loss of notification -tubing disconnect -sensor failure | -unable to obtain readings -failure to deploy -failure to service
FAB-H-REQ-147 | Oxygen sensor calibration | The ventilator shall provide automated means for calibration of the oxygen sensor, at start-up and at intervals not exceeding 24 hours, without interruption to the patient inspired oxygen concentration. | -fail to calibrate manually -out of calibration -fail to request calibration | -excess flow or over-infusion -medical gas supply problem -failure to service
FAB-H-REQ-530 | Oxygen sensor calibration errors | The oxygen sensor calibration process in use shall not cause the oxygen calibration to become corrupt due to a slow response from the oxygen cell. Note: the pneumatic design results in typical response time of the cell to a switch to calibration gas less than 3 s, However this is a function of supply gas pipeline pressure, ventilation flow rate, and circuit back-pressure, and calibration may fail if the response is very slow. If this occurs in clinical use (not at startup or PM) the oxygen sensor calibration should be retained. | -software corruption -fail to detect errors -loss of measurement | -medical gas supply problem -no flow -imprecision
FAB-H-REQ-148 | Capnometry zero | When a capnometry device is in use, the ventilator shall provide means for the user to perform zero calibration for the capnometry sensor. | -fail to provide indicator -fail to request calibration -fail to calibrate manually | -failure to calibrate -no display or display failure -incorrect display
FAB-H-REQ-155 | Specification of planned maintenance | Planned maintenance shall be specified in a Service Manual, and / or service procedure documents. | -lack of maintenance documentation -lack of documenting procedure -lack of maintenance process | -device markings issue -Device Handling Issue -Material Integrity Issue
FAB-H-REQ-156 | Blender flow sensor calibration | Planned maintenance documentation shall include a procedure to calibrate the blender flow sensors. | -lack of tracking procedure -incorrect preventative maintenance -lack of test specification | -device markings issue -Deployment Issue -accessory incompatible
FAB-H-REQ-157 | Exhalation valve characterization | Planned maintenance documentation shall include a procedure to characterize the exhalation valve. Note: exhalation valve characterization stores the pressure curve vs drive current. | -inadequate document -lack of procedure -inadequate test procedure | -pressure issue -defective component -component or accessory incompatibility
FAB-H-REQ-158 | Pressure sensor calibration | Planned maintenance documentation shall include a procedure to calibrate the inspiratory and proximal pressure sensors. | -lack of tracking procedure -lack of procedure -lack of evaluating procedure | -no device output -pressure issue -failure to deliver energy
FAB-H-REQ-524 | RTC battery replacement | Planned maintenance documentation shall include replacement interval and instructions for the RTC battery. | -missing maintenance -fail to plan deliverable -inadequate task analysis | -device markings issue -parameter calculation error due to software problem -use of incorrect control settings
FAB-H-REQ-159 | Replacement part installation instructions | All replacement parts shall be provided with instructions for installation, and which confirm what software version, software re-installation, calibration and processes are required following installation. | -software version incorrect -inadequate software configuration -lack of implementation process | -compliance issue -device markings issue -incompatibility problem
FAB-H-REQ-161 | Operating hours counter | The ventilator shall provide a counter for operating hours with a resolution of one hour and maximum no less than 100,000 Note: The operating hours counter is stored in the mainboard. The exchange procedure should provide for the counter to be copied from the old mainboard to the new mainboard. | -inadequate service -long response time -lack of document | -incorrect software programming calculations -fail-safe mechanism issue -electrical power problem
FAB-H-REQ-162 | Service due date indication | The ventilator shall incorporate means to set, and display, the next due service date. | -lack of expiration date -incorrect time and date -lack of delivery date | -electronic property issue -use of incorrect control settings -instruction for use issue
FAB-H-REQ-163 | Planned maintenance frequency | Planned maintenance activities shall not be required at more frequent than 12-month intervals. | -misassembly during maintenance -tool activation failed -delivery delayed | -Delivery System Failure -insufficient flow or underinfusion -Failure to adhere or bond
FAB-H-REQ-164 | Alarm disable during servicing | The ventilator shall provide means to temporarily disable alarms, only accessible through a password-controlled service screen. | -fail to disable alarm -disable pacing -loss of state (activate or deactivate or alarm or warning) | -unauthorized access to computer system -instruction for use issue -loss of power
FAB-H-REQ-165 | Alarm system re-enable | The alarm system shall be re-enabled on power-cycle. Note: the alarm disable function is solely intended to facilitate maintenance and calibration without causing technician annoyance due to nuisance alarms while not in clinical use. | -fail to set or reset alarm automatically -fail to reset alarm automatically -fail to activate alarm | -Device displays error message -no audible alarm -device alarm system issue
FAB-H-REQ-166 | Software installation procedure | A software installation procedure shall be documented that includes verification of the complete installation hardware and software configuration and that generates an installation report to be provided to the manufacturer. | -lack of documenting procedure -lack of installation document -lack of document | -installation-related problem -use of device issue -leak

## IFU
category | Inputs | Failure_Modes | Hazards
Maintenance | The fabian HFO device is a ventilator classified as Class IIb (for prolonged use more than 24 hours and less than 30 days) according to the European Medical Devices Directive | -device malfunction -inappropriate classification -inappropriate user | -installation-related problem -defective component -electronic property issue
Maintenance | Inspection according to manufacturer specifications is required every 12 months | -maintenance does not comply to manufacturers recommendations -fail to perform inspection -inadequate task analysis | -use of device issue -electronic property issue -device contamination with biological material
Maintenance | Maintenance must be performed by ACUTRONIC Medical Systems trained personnel with access to appropriate test and measuring equipment | -lack of maintenance process -fail to perform inspection -inadequate training | -no device output -Failure to adhere or bond -human-device interface issue
Maintenance | Only use original ACUTRONIC Medical Systems parts for repairs | -missing parts -component damaged -device damaged by inappropriate cleaning | -therapeutic or diagnostic output failure -unauthorized access to computer system -Inflation Issue
Liability | In the event of improper equipment maintenance or repairs by any persons not associated with ACUTRONIC Medical Systems AG, Service or improper use, all liability for the functionality is transferred to the owner or operator. | -improper device position -loss of function -loss of transmission | -therapy delivered to incorrect body area -Insufficient Flow or Underinfusion -degraded
Liability | ACUTRONIC Medical Systems AG, assumes no liability for damages caused by the nonobservance of preceding notices. The preceding notices DO NOT extend the warranty and liability terms of the ACUTRONIC Medical Systems AG, sales terms and delivery conditions | -void warranty -inappropriate usage -inappropriate user | -use of device issue -device handling issue -dose calculation error due to software problem
Intended use | The fabian HFO is intended for premature infants, new-borns as well as children weighing up to 30 kg. | -premature ventricular contraction -weight out of range -low weight capacity | -instruction for use issue -incorrect display -inadequate user interface
Intended use | The fabian HFO is intended for “in-patient use” in hospitals, medically used rooms and intrahospital patient transport | -patient data breached -confused patient notification -clinical setting environment violated | -user used incorrect product for intended use -use of incorrect control settings -reset issue
Intended use | The fabian HFO is an electronically microprocessor-controlled ventilator | -fail to output current -fail to suit -software can not recover from reset | -inadequate user interface -failure to reset -reset issue
Intended use | The fabian HFO ventilates with excess pressure based on the continuous-flow principle | -inappropriate pressure -excessive vibration -inaccurate pressure reading | -pressure issue -use of incorrect control settings -device difficult to setup or prepare
Intended use | Oxygen is metered by the integrated Air / O2 blender | -fail to vent air out -fail to pump -fail to hold pressure | -air leak -excess flow or over-infusion -calibration issue
Intended use | The oxygen concentration is measured internally with a galvanic oxygen sensor | -over oxygen concentration -fail to update oxygen volume -incorrect measurement | -Pressure Issue -incorrect software programming calculations -failure to recalibrate
Intended use | The ventilator is intended for the following ventilation methods: Continuous Positive Airway Pressure (CPAP), High and Low flow oxygen therapy (HFNC: O2 Therapy), High Frequency Oscillation (HFO), (membrane principle), Intermittent Positive Pressure Ventilation (IPPV), Pressure Support Ventilation (PSV), Synchronized Intermittent Mandatory Ventilation (SIMV), Synchronized Intermittent Mandatory Ventilation combined with PSV (SIMV + PSV), Synchronized Intermittent Positive Pressure Ventilation (SIPPV), Ventilation nCPAP / DUOPAP with variable flow generators (NIV), (Infant Flow™, Infant Flow™ LP, Inspire™, Medijet®) | -inappropriate ventilation support -inadequate ventilation -incorrect ventilation mode | -Pressure Issue -Occlusion within Device -decrease in pressure
Intended use | The equipment is operated by a physician or by a physician’s orders by a professional with technical training in this task, any operator must be trained on this equipment, be familiar with the operator’s manual and have knowledgeable use of the equipment. | -inappropriate training -inadequate training -inadequate training material | -Torn Material -decreased sensitivity -radio signal problem
Intended use | fabian HFO is NOT approved for use in a homecare environment | -inappropriate usage environment -fail to approve -incorrect product classification | -no flow -device packaging compromised -market issue
Scope of delivery | The fabian HFO product includes the following items: One fabian HFO Ventilator, One Accessory kit( One Flow Sensor reusable, One Flow Sensor Cable, One Test Lung), Two Infant Flow Ventilator Tubes, One Power Cable (# country specific), One Operating Manual (# country specific) | -accessory missing -inappropriate indication -component incompatible | -imprecision -Pressure Issue -fluid leak
Scope of delivery | The duration of indirect patient contact is defined as prolonged (i.e. more than 24 hours and less than 30 days) | -long response time -short circuit -inappropriate user | -infusion or flow issue -defective item -Battery Issue
Contraindications | Severe airflow obstruction and intracranial-hypertension would contraindicate the use of the fabian HFO neonatal and infant ventilator | -airway obstructed -inadequate humidification -inappropriate flow volume | -Failure to Cycle -use of device issue -Device Inoperable
Contraindications | In the event of ventilation for several hours or more, care must be taken for optimal conditioning of the respiratory gases (warmth, humidification) to optimize secretion mobilization and prevent damage to mucous membranes | -fail to maintain ventilation mode -inadequate service -fail to retain settings | -Biocompatibility Issue -Pressure Issue -incorrect measurement
fabian Front connections | Devices with serial number prefix AI / AL: External Bias Flow (FG - Fresh Gas) port and port for nCPAP system based on flow generators (single limb systems); Expiratory limb port; Proximal Pressure port; Inspiratory limb port/ HFO port | -incorrect pin number -port mismatched -inappropriate interface design | -compliance issue -Material Integrity Issue -device markings issue
fabian Front connections | Devices with serial number prefix 20 /AK/AH: Inspiratory Limb port/ center port for connecting nCPAP system based on flow generators (single limb systems); Expiratory limb port; Proximal Pressure port; HFO Port | -incorrect pin number -port mismatched -loss of connection | -fitting problem -pressure issue -device difficult to setup or prepare
Rear panel | Hardware with HDMI output (SN AI-01500 and AL-00400 or higher): Video Out HDMI connection; USB connection for connection of a USB powered nebulizer; Network jack for data management (DISABLED); RS232 interface, service, PDMS; Flow Sensor Connector; Nurse Call Connector, max switching voltage 30V DC; Loudspeaker (Audio); Fan; Connector for etCO2 module (optional); Connector for SpO2 module; Connection for Oxygen “O2” supply 2.0 to 6.0 bar / max. 40 L / min; Connection for pressurized Air supply 2.0 to 6.0 bar / max. 40 L / min; Main Power Connector with fuse holder; Terminal stud for potential equalization | -loss of output -loss of electrical connection -fail to connect | -application program issue -therapeutic or diagnostic output failure -fitting problem
Rear panel | Hardware with Video input: Connector for 24V DC external power supply (No charging); Network jack for data management, PDMS; USB port for data output, Software update and connection for Masimo SpO2 module; DB9 RS-232 port for PDMS; Flow Sensor 7-pin Connector; Nurse Call 3-pin Connector; Video In, VGA (NOT USED); Loudspeaker (Audio); Fan; CO2 sensor (Optional); Nebulizer (Not used); O2 supply connector 2.0 to 6.5 bar/ 40 L/min; Pressurised air connector 2.0 to 6.5 bar /40 L/min; Power Connector (fuse 1.25 AT); Equipotential connection | -loss of video input -loss of power -loss of electrical connection | -compliance issue -data issue -problem with software installation
Rear panel | Initial hardware model: Connector for 24VDC external power supply; Network jack for data management, PDMS; USB port for data output and Software update, Connection for Masimo SpO2 module; DB9 RS-232 port for Service, CO2 option or PDMS; Flow Sensor 7-pin Connector; Nurse Call 3-pin Connector; Loudspeaker (Audio); Fan; Optional Ports; Power Connector (Fuse 1 AT); Equipotential connection; O2 Supply Connector 2.0 to 6.5 bar /max. 40 L/min; Pressurized Air Connector 2.0 to 6.5 bar /max. 40 L/min | -loss of power -insufficient power capacity -loss of electrical connection | -data issue -problem with software installation -Low Battery
Rear panel | General hardware warning characteristics: DO NOT connect Ethernet, Nurse call, USB, RS-232 (CO2), Flow Sensor port to anything other than specified devices | -loss of warning -fail to connect -incorrect warning | -defective component -insufficient flow or underinfusion -device operates differently than expected
Rear panel | General hardware warning characteristics: DO NOT connect anything to Ethernet, Nurse call, and RS232, while operating on battery power. | -loss of warning -fail to connect -incorrect warning | -use of device issue -solder joint failure -defective component
Rear panel | General hardware warning characteristics: ONLY connect IEC 60601-1 or IEC 60950-1 compliant devices to Ethernet, Nurse call, and RS232 ports. | -inadequate warning -lack of warning -loss of warning | -electrical issue -device-device incompatibility -loose or intermittent connection
Rear panel | The Potential Equalization Pin is for additional safety and can be connected to an equipotential zone. Adhere to local guidelines when using this PIN. The guidelines may vary between countries, localities and power companies. Always keep the PIN for Potential equalization accessible. | -loss of connection -pin length out of specification -connector mismatched | -compliance issue -use of device issue -user used incorrect product for intended use
Rear panel | Maximum Connected Loads: Nurse call : Isolated relay output, Max contact load: 30 VDC @ 1A; USB : 5V @ 150 mA max; RS-232 : Signal Levels: EIA/TIA-232 Standard, Pin 9 Power: 5V @ 500 mA max (models without DC input), Isolated; Ethernet : N/A (Disabled); Flow Sensor : Maximum load is one Flow sensor, (Only ACUTRONIC Medical Systems Flow sensors can connect), Max. Voltage: 5 V, Max operating current: 300 mA per hotwire -> 600 mA total; SpO2 : Max. Voltage: 3.3 V, Max. operating current: 200 mA; etCO2 : Max Voltage: 5 V, max operating Current: 700 mA | -loss of connection -power out of range -over load | -defective component -use of incorrect control settings -misconnection
Control panel options | The Touch screen (1) allows the direct control of the ventilator parameters by pressing defined buttons on the Graphic User Interface (GUI). | -insensitive touch screen -unresponsive touch screen -incorrect screen display | -inadequate user interface -human-device interface issue -failure to conduct
Control panel options | Access to Quick Launch settings is on the left side of the touch screen. | -incorrect settings -confused settings -sticky button | -human-device interface issue -application security issue -unauthorized access to computer system
Control panel options | The Key / Control Panel (2) with Rotary Pulse encoder (combines a key and a rotary encoder). | -key stuck on -incorrect encoder count -loss of control | -defective item -therapeutic or diagnostic output failure -component incompatible
Function buttons | Home: Displays the Main Screen for selecting the Respiration mode. | -incorrect screen display -inadequate display message -incorrect display format | -failure to sense -incorrect display -No Display or Display Failure
Function buttons | Graphics: Switches to Curve / loop display. | -incorrect display format -inappropriate interface design -loss of visual | -Material Integrity Issue -Difficult to Position -Malposition of Device
Function buttons | Manual Breath: Used to apply a manual breath. Available in all the Ventilation modes. In HFO mode, it can be disabled in the Ventilation menu. | -fail to initiate breath -incorrect breath type -inadequate breath type | -no display or display failure -fluid leak -user used incorrect product for intended use
Function buttons | Alarm limits: Switches to the Alarm Limits Configuration. | -fail to raise low priority alarm -fail to set limit correctly -fail to disable alarm | -device alarm system issue -Device Alarm System Issue -defective item
Function buttons | Alarm Silence: For acknowledging and audibly silencing alarms for a maximum duration of 120 seconds. | -fail to silence alarm -low alarm volume -audio mute failed | -device alarm system issue -Device Handling Issue -low audible alarm
Function buttons | Alarm Silence: Subsequent alarms with higher priority are visually displayed during periods of alarm silence. | -fail to generate alarm -incorrect message display -fail to activate alarm | -device alarm system issue -Device Alarm System Issue -alarm not visible
Function buttons | menu / Calibration: Used to access the Configuration and Calibration menus: Press once to open the Configuration menu; Press again to open the Calibration menu. | -fail to setup configuration -fail to access -fail to configure infusion parameters | -human-device interface issue -unexpected therapeutic results -Suction Issue
Function buttons | Nebulizer: Obsolete function | -device malfunction -device stops intermittently -blank screen | -device operates differently than expected -Sticking -failure to deploy
Function buttons | O2 Flush: Used to start O2 Flush. Flush concentration and time can be preset. | -incorrect concentration -low filter effect -inappropriate infusion mode selected | -no display or display failure -medication error -medical gas supply problem
Function buttons | Start / Stop: Used to temporarily stop or commence respiration. During mechanical ventilation this provides a two-minute pause in ventilation. In nCPAP, DUOPAP and O2 Therapy modes this acts as a Standby mode (therapy is disabled until re-enabled by the operator). | -stop failed -fail to disable ventilation -inappropriate pause period | -improper or incorrect procedure or method -delivery system failure -Device Handling Issue
Function buttons | ON / OFF: Used to switch the device ON or OFF. | -use device incorrectly -fail to turn on device -device failsafe failed | -material integrity issue -inadequate training -Connection Issue
Function buttons | The Rotary Pulse encoder combines a push button with a rotary encoder for executing various settings, selections and confirmation options | -fail to indicate button status -fail to initiate pulse -incorrect settings | -therapeutic or diagnostic output failure -inadequate user interface -failure to prime
Display areas | The information system features two key display areas: The Touch screen, LED indicators | -incorrect display format -incorrect message display -inadequate message display | -no display or display failure -incorrect display -device markings issue
Display areas | The display screen shows various information, setting and display areas depending on the display settings or menu. | -incorrect screen display -incorrect display format -incorrect message display | -incorrect display -no display or display failure -use of incorrect control settings
Display areas | Information bar indicating: Battery status, Time / Date, Status information, System and Alarm information. | -fail to indicate battery level -loss of state indication -loss of indication | -battery issue -charging issue -data back-up problem
Display areas | Graphic area (single graph or multigraph with loops or curves). | -area out of specification -inappropriate visual interface -incorrect display format | -poor quality image -defective item -degraded
Display areas | Parameter settings and controls area. | -use incorrect control settings -confused settings -inaccurate settings | -use of incorrect control settings -output issue -reset issue
Display areas | Operating mode selection | -inappropriate program selected -selection failed -fail to select signal | -use of device issue -improper or incorrect procedure or method -defective item
Display areas | Numeric area / alarm limits. | -incorrect limit setting -inadequate traceability -inappropriate variable value | -defective item -device alarm system issue -alarm not visible
Display areas | Depending on the display mode selected, individual areas can be shown / hidden. | -incorrect message display -inadequate message display -invisible settings | -no display or display failure -incorrect display -use of incorrect control settings
Information bar | The Information bar displays from general information to displaying alarms. It is divided into three sections. | -fail to generate alarm -incorrect message display -fail to display message | -no display or display failure -deployment issue -leak
Information bar | The Information bar indicates the following among other things: Neonatal or Pediatric mode, Patient Data available, Manual Breath in HFO mode ON / OFF, Nurse Call Active / Inactive, Battery Charge Status | -loss of indication -incorrect logo -inappropriate user | -data issue -device handling issue -improper device output
Information bar | Ventilation mode, information and Alarm Display. | -incorrect message display -incorrect display -incorrect display format | -use of incorrect control settings -failure to sense -Device Alarm System Issue
Information bar | Alarm Silence (time remaining until alarm reactivates) Time / Date. | -fail to silence alarm -sound mute failed -device paused | -device alarm system issue -Device Alarm System Issue -improper alarm
Information bar | Pediatric mode. For patients above 10 kilograms body weight, use of Pediatric mode is recommended. Extended range for flow and volume is available. | -inappropriate flow volume -weight out of range -low volume | -use of device issue -protective measure issue -capturing issue
Information bar | Neonatal mode. Used for patients up to 10 kilograms body weight. | -weight out of range -heavy weight -inappropriate user | -Material Integrity Issue -user used incorrect product for intended use -application program issue
Information bar | The Patient Symbol indicates patient data associated with the current ventilator operation is saved to the device. | -loss of indication -fail to transfer data -data back-up failed | -user used incorrect product for intended use -application security issue -inadequate user interface
Information bar | Ventilation mode / Information / and Alarm display. Displays Ventilation mode and, if applicable, additional current information and notices. | -incorrect ventilation mode -inadequate information -inadequate alarm states | -incorrect display -protective measure issue -use of device issue
Information bar | A RED blinking information bar also provides alerts to active alarms. | -fail to activate alarm -loss of alarm type indication -red regulation violated | -defective item -use of incorrect control settings -use of device issue
Information bar | If Screen Lock is in use, the symbol for Locked Screen is displayed. | -incorrect screen display -fail to lock -incorrect message display | -incorrect display -No Display or Display Failure -no display or display failure
Numeric field / alarm limits | In the Numeric area, all measured values are displayed together with the set limits relevant in the selected Ventilation mode. | -incorrect display value -incorrect message display -inadequate message display | -incorrect display -no display or display failure -output below specifications
Numeric field / alarm limits | There can be multiple pages of Numeric areas. | -lack of configuration item -inadequate interface specification -inadequate process | -optical issue -Item Contaminated during Manufacturing or Shipping -defective component
Numeric field / alarm limits | To go to the next page, press the button (1) below the Numeric area. | -fail to indicate button status -indistinct button -sticky button | -human-device interface issue -use of device issue -no display or display failure
Numeric field / alarm limits | Depending on the Ventilation mode the data is updated as average / minute or by breath. | -incorrect data -fail to update configuration -lack of update | -failure to service -data issue -application security issue
Numeric field / alarm limits | The upper and lower limits can be manually or automatically adjusted in the alarm limits menu. | -fail to adjust manually -incorrect limit setting -unable to automatic off | -break -device difficult to setup or prepare -positioning issue
Numeric field / alarm limits | To adjust the alarm limit manually: Press the Alarm Limits button; Touch the Parameter (2) you wish to change; The button turns from dark Blue to Green, With the Rotary Pulse encoder, change the value; To confirm the setting, either touch the Parameter button again or push Rotary Pulse encoder again. | -fail to adjust manually -incorrect limit setting -adjustment failed | -device difficult to setup or prepare -failure to align -sticking
Numeric field / alarm limits | To automatically set the alarm limits, select parameter and touch the AUTOSET button. | -incorrect limit setting -fail to set limit correctly -fail to reset alarm automatically | -use of incorrect control settings -low audible alarm -use of device issue
Numeric field / alarm limits | The alarms can be set between the following ranges: Apnea [2 to 20, OFF], CPAP [mbar], DCO2 [mL2 / sec], Frequency [bpm], Leak [%], Minute Volume [liter], PEEP [mbar], PIP [mbar], Pmean [mbar] | -fail to test alarm -improper alarm -fail to reset alarm automatically | -use of incorrect control settings -pressure issue -parameter calculation error due to software problem
Numeric field / alarm limits | WARNING: The respiratory rate and minute volume monitors are calculated as averages using an update period of 6 seconds. If the monitored breath rate falls below 11 bpm, these values are not reported. | -incorrect rate calculation -fail to update oxygen volume -inaccurate rate | -use of device issue -failure to service -failure to reset
Graphics display | Displays the current Pressure, Volume or Flow measurements as a diagram. Use the Graphics key to access the Graphics menu; Three waveforms are simultaneously displayed. | -loss of pressure -fail to display message -fail to display point data | -use of device issue -incorrect display -device markings issue
Graphics display | You can also switch to Loop Display view. In this view two loops and one of three waves are displayed. LOOP Display: Pressure / Volume, Volume / Flow | -incorrect display format -fail to manage switch -fail to raise pressure | -device markings issue -incorrect display -use of device issue
LED indicators | Warning LED: This LED illuminates or blinks Red when a system alarm is triggered. | -led stuck on -illumination transmit failed -fail to generate alarm | -device alarm system issue -defective item -Device Alarm System Issue
LED indicators | Battery LED: This LED illuminates Yellow in Battery operation and Green when the battery is fully charged. The LED blinks Green when the Battery is charging. | -fail to indicate battery charging status -reduced capacity -device burned | -battery issue -device-device incompatibility -misfire
LED indicators | Mains LED: This LED illuminates Green in Mains operation and switches OFF in battery operation | -illumination transmit failed -fail to energize -out of range | -solder joint failure -device handling issue -electrical issue
LED indicators | When powered by the external 24 VDC connector, the Mains and Battery LED will NOT illuminate. | -loss of illumination -illumination transmit failed -fail to infuse | -Battery Issue -degraded -Loss of or Failure to Bond
LED indicators | If the device shuts down with warning LED blinking simultaneously with battery LED, this represents a device fault. Then follow these steps: Immediately take the device out of service and provide an alternative means of ventilation to the patient; Report the incident to your local distributor | -loss of warning -system fault displayed -fail to activate alert | -improper or incorrect procedure or method -therapeutic or diagnostic output failure -fitting problem
LED indicators | DO NOT use the device until it has been checked by suitably trained and qualified personnel. | -inappropriate training -unqualified use -lack of training indication | -use of device issue -failure to deliver -dose calculation error due to software problem
Ventilation menu | Push the buttons and their statuses are indicated by various colors: Light Blue: Push button; Dark Blue: focused push button; Yellow: (in Functions menu): parameters are monitored by the system; Green: selected push button if the push button features a parameter for setting a value, it can now be changed; Orange: parameter is automatically adjusted by the device; The value that you set indicates the maximum value and will NOT be exceeded. | -fail to indicate button status -confused button -loss of indication | -inadequate user interface -protective measure issue -date-related software issue
Ventilation menu | Depending on the Ventilation mode, the The Information bar (1) indicates current information; Using the push button, you can also switch between the Pressure-, Volume and Flow curve (2); Respiration menu indicates the Configuration Parameters (3) and Measurements (4) relevant in the mode; The Ventilation mode can be changed in the menu (5). | -fail to indicate priority -loss of indication -incorrect message display | -use of device issue -protective measure issue -incorrect display
Ventilation menu | The Ventilation parameters can be set before starting /activating Ventilation mode: For the desired Ventilation mode, tap the key once: the key will turn Yellow; The Configuration parameters for the preselected Ventilation mode can now be adjusted; Tap Ventilation mode again: the key turns Green; ventilation starts with the parameter settings. | -incorrect ventilation mode -inappropriate ventilation support -incorrect parameter | -use of incorrect control settings -failure to deploy -device difficult to setup or prepare
Ventilation menu | The selected parameter can also be confirmed by pressing the Rotary Pulse encoder. | -fail to confirm -incorrect parameter -parameter confused | -human-device interface issue -inadequate user interface -protective measure issue
Ventilation menu | If the setting is active (Yellow button) and no action is taken within 15 seconds, or the preselected mode is NOT confirmed by tapping again, the device continues in the previous mode, any settings are deleted. | -incorrect settings -inappropriate timeout setting -fail to activate alarm | -use of incorrect control settings -use of device issue -incorrect display
Ventilation menu | After confirming a parameter, the audible and visual alarm is automatically suppressed for 15 seconds. | -loss of audible alarm -loss of visual alarm -fail to disable alarm | -device alarm system issue -failure to sense -use of incorrect control settings
Ventilation menu | The audible and visual alarm can be immediately activated by pressing the Alarm Silence button. | -fail to activate alarm -loss of audible alarm -loss of visual alarm | -Device Alarm System Issue -not audible alarm -low audible alarm
Ventilation menu | Press any blank area or another parameter to NOT accept the parameter and keep the previous value. | -blank screen -incorrect parameter -unreadable display | -defective item -use of device issue -insufficient flow or underinfusion
Ventilation menu | Alternating alarms active at the time the mode is switched will automatically be reset. | -fail to reset alarm automatically -fail to set or reset alarm automatically -fail to reset error message automatically | -device operates differently than expected -break -device alarm system issue
Ventilation menu | If parameters are mutually regulating, an arrow will appear on the parameter requiring modification indicating the required direction. | -regulation failed -incorrect parameter -aami regulation violated | -device markings issue -inadequate user interface -user used incorrect product for intended use
Ventilation menu | Some Ventilator parameters are “Locked” to prevent unusually high values. When this value is reached, a key symbol appears and a “Notice Signal” will sound. The high value must be confirmed by pressing the Rotary Pulse encoder again. | -parameter confused -incorrect parameter -inappropriate marking symbol | -decrease in pressure -false alarm -Occlusion within Device
Graphics menu | The Graphics menu can be accessed from the Graphics button on the Keypad | -confused button -fail to access -unable to access | -device difficult to setup or prepare -use of device issue -inadequate user interface
Graphics menu | The following graphics are available in the Graphics menu: curves, loops, trends | -lack of trend analysis -inappropriate visual interface -loss of image | -user used incorrect product for intended use -instruction for use issue -reset issue
Graphics menu | The Graphics menu shows the following curves: Pressure, Flow, Volume | -loss of pressure -inadequate flow -incorrect graph | -use of device issue -incorrect display -inadequate instructions for healthcare professional
Graphics menu | When auto-scaling of graphics is switched OFF, the graph can be adjusted manually: Select desired graph; Press the graph; The selected graph is marked with the “Cursor” symbol and the Scaling cursor is displayed; Use the cursor to scale the graph in the X and Y direction. | -fail to adjust settings -incorrect graph -adjustment failed | -no display or display failure -compliance issue -incorrect display
Graphics menu | The cursor disappears if not used within five seconds, or by tapping the selected graph again | -fail to move cursor -fail to cancel suspend mode -invisible indication | -use of device issue -user used incorrect product for intended use -inadequate instructions for healthcare professional
Graphics menu | The Freeze key stops the curve being updated. | -key stuck on -lack of update -fail to update configuration | -failure to deploy -break -defective item
Graphics menu | The Configuration parameters and measurements in the Display area continue being updated. | -fail to update configuration -incorrect parameter -incorrect configuration | -failure to deploy -incorrect display -use of incorrect control settings
Graphics menu | Press the Freeze key again to continue updating the Curve data. | -key stuck on -invalid key -fail to update display | -unauthorized access to computer system -device maintenance issue -patient data issue
Graphics menu | Depending on the Ventilation mode selected, the Graphics menu displays the Configuration parameters and measurements relevant in this mode. | -incorrect ventilation mode -incorrect message display -inadequate message display | -use of incorrect control settings -use of device issue -device operates differently than expected
Graphics menu | The Loops menu features the following loops: Pressure / Volume (P/ V Loop), Volume / Flow (V / F Loop) | -fail to adjust volume -loss of pressure -incorrect configuration | -incorrect display -incorrect measurement -failure to service
Graphics menu | The following parameters for the respective Curve can be selected in the upper area of the display: Pressure, Volume, Flow | -incorrect parameter -parameter confused -fail to update pressure | -protective measure issue -incorrect measurement -material rigid or stiff
Graphics menu | The Freeze key is used to pause Loop updating. The Settings parameters and measurements in the display area will only be shown after a Loop has been saved. The key will turn Green. Press the Freeze key again to continue updating the Loop Data. | -fail to update display -fail to update configuration -key stuck on | -failure to deploy -incorrect display -use of incorrect control settings
Graphics menu | After freezing the loop, it can be saved to later compare it against the Reference Loop. Save Loop: Freeze Loop, Push Freeze Key, The Loop will be saved until a New Loop is saved. | -fail to attach within time frame -fail to deliver -unable to access | -use of device issue -device maintenance issue -no display or display failure
Graphics menu | Comparing loops: Switch to Loops Display, Freeze current Loop, Current Loop is shown in Blue. the Saved Loop is shown in Black. | -inadequate markings -loss of indication -loss of grounding | -inadequate user interface -electrical issue -protective measure issue
Graphics menu | The device trending function automatically saves an average of measurements every 30 seconds. | -fail to save configuration -loss of measurement -inadequate time format | -device operates differently than expected -loss of power -Power Source Issue
Graphics menu | Measurements of up to five days can be recorded. Trend Data is automatically deleted in the following cases: Date and / or Time is modified, New Patient data is entered, Software Update | -fail to record data -inappropriate data flow -fail to maintain record | -deployment issue -patient data issue -data issue
Graphics menu | A confirmation message will always be displayed before Trend Data is deleted: Parameter / Selection, Graphics, Parameter values, Timeline, Adjust Timeline (minimum 30 minutes to a maximum of five days), . Shift Timeline, Update Data (refresh) | -incorrect message display -inadequate message display -fail to activate message | -positioning issue -image resolution poor -device difficult to setup or prepare
Graphics menu | The Trend menu features three simultaneous parameters as curves. | -inaccurate trend result -incorrect parameter -incorrect program parameter | -incorrect display -use of incorrect control settings -inadequate user interface
Graphics menu | The Parameters displayed can be modified by activating the Select List on the left side. Available parameters are: % Mvmand, Compliance, DCO2, EtCO2, FiO2, Freq, HF Amplitude, Leak, MV, Pinsp, Pmean, Resistance, RSBI, SpO2, SpO2 PI, SpO2 PR, VTe | -incorrect parameter -parameter confused -configuration changed | -no display or display failure -inadequate user interface -incorrect display
Graphics menu | The Timeline can be shifted for all three parameters simultaneously: Change timeline via Rotary Pulse encoder. | -incorrect parameter -inefficient position -inappropriate transition point | -use of incorrect control settings -device difficult to setup or prepare -component incompatible
Preparing for operation | DO NOT use any accessory if individual package is damaged. If the packaging of the ventilator is damaged use the device only if the start-up device check is performed successfully. | -device damaged -package broken or damaged -component damaged | -protective measure issue -user used incorrect product for intended use -reset issue
Preparing for operation | For initial installation and when lifting the fabian HFO, disconnect all cables and circuits from the ventilator, grab the ventilator with both hands from each lateral side, see arrows indicating on the left side in the figure below, and lift the ventilator carefully. | -circuit broken or damaged -missing route -part mismatched | -device-device incompatibility -use of device issue -fitting problem
Preparing for operation | Always secure the fabian HFO using the dedicated screw at the bottom of the device | -deposit in device or device ingredient -device damaged -unsecured cap | -No Known Device Problem -detachment of device component -device expiration issue
Preparing for operation | Connect the fabian with power cable to a suitable power outlet. | -connection broken or damaged -loss of connection -defective power supply unit | -loss of power -fitting problem -detachment of device component
Preparing for operation | DO NOT connect the device to a power outlet strip. Power fluctuations from the system may trip the circuit breaker in the power outlet strip causing a power shut OFF to the device. | -circuit broken or damaged -fail to indicate ac or battery power -loss of connection | -loss of power -Failure to Power-Up -Power Source Issue
Preparing for operation | An IEC60601-1 approved power outlet strip with valid amperes rating by the manufacturer for connection to a ventilator. | -fail to approve -switch to ac power failed -fail to supply power | -fitting problem -failure of device to self-test -application program issue
Preparing for operation | The device can be operated with 100 to 240 VAC and automatically adjusts to voltage without manual switch-over being necessary. | -fail to adjust settings -inappropriate operation environment -fail to adjust pressure | -detachment of device or device component -pressure issue -electrical issue
Preparing for operation | warning: The use of fuses must correspond with the value and type approved for the fabian: The fuse type and value is printed on the back of the fabian; Replacement of the fuses can only be carried out by trained staff. Replace fuses with identical values. Failing to do so, can cause fire hazards. | -lack of fuse -defective switch -incorrect electrode placement | -use of device issue -use of incorrect control settings -medication error
Preparing for operation | warning: To avoid risk of electric shock, this equipment must only be connected to a Supply Main with protective earth ground. | -protective earth conductor failed -electrical shock -fail to supply power | -device cleaning issue -fitting problem -compliance issue
Preparing for operation | Connect the compressed air and oxygen supply tube to the back of the device and to the Central Gas Supply. | -fail to air dry -fail to update oxygen volume -fail to rotate | -positioning issue -detachment of device component -leak
Preparing for operation | Use water trap (see Accessory list) in case humidity can be present in the compressed air pipeline. | -humidity out of range -water resistant failed -inadequate humidification | -improper or incorrect procedure or method -Torn Material -split
Preparing for operation | If there is no central gas supply, gas supply from cylinders is also possible. Use cylinder with approved regulator per CGA/ISO standards, including fill indicators. | -fail to generate indication -fail to provide indicator -fail to supply power | -not audible alarm -failure to sense -particulates
Preparing for operation | The inputs are coded to prevent wrong connection. | -wrong connection -loss of connection -invalid input | -user used incorrect product for intended use -accessory incompatible -market issue
Preparing for operation | Medical grade oxygen and medical grade air (dust-free, oil-free and dry, 2.0 bar to 6.0 bar) is required | -oxygen out of specification -fail to update oxygen volume -inadequate breathing mode | -Failure to adhere or bond -Failure to Cycle -output above specifications
Preparing for operation | WARNING: ELECTRICAL HAZARD NEVER use anti-static or electrically conductive tubes. | -fail to activate system -device burned -fail to identify safety | -device handling issue -crack -inadequate user interface
Preparing for operation | WARNING: The pressure gradient on the fabian system of the ventilator measured on the patient connection port can increase when accessories or other components are attached to the system. Change in pressure gradient can adversely affect the performance of the fabian system. | -pressure increased -incorrect measurement -fail to adjust pressure | -Connection Issue -Power Source Issue -inadequate filtration process
Preparing for operation | Connect the tubing set: Connect Inspiratory tube to Insp. Port; Connect Expiratory tube to Exp. Port; Connect Proximal Pressure measuring tube to Prox. Port; Use FG Connection when connecting to an INOvent Flow sensor. | -tubing malformed -fail to connect -tubing blocked | -loose or intermittent connection -Disconnection -Connection Issue
Preparing for operation | Always hold tubes by the collar when connecting and disconnecting to prevent damage. | -loss of connection -tubing detached -unable onehand to detach | -Use of Device Issue -Device Disinfection or Sterilization Issue -Difficult to Insert
Preparing for operation | The Expiration membrane can be easily removed by turning the membrane holder counterclockwise. | -fail to remove -unable onehand to remove -difficult to remove | -incorrect software programming calculations -component incompatible -unable to obtain readings
Preparing for operation | Pay attention to the correct installation position of the Expiration Membrane. The labeling must be legible upon installation. | -incorrect installation -missing installation info -inadequate labeling | -positioning issue -detachment of device or device component -detachment of device component
Preparing for operation | WARNING: The HFO outlet connector must be connected with the Inspiratory limb. Connection to Expiratory limb will result in incorrect ventilation! | -incorrect ventilation mode -connector mismatched -loss of connection | -misconnection -Delivery System Failure -failure to deploy
Preparing for operation | connect nitric oxide (NO) tubing sets: Only INOmax and INOvent NO systems are approved by ACUTRONIC Medical Systems. If other manufacturers’ NO systems are used, they must be validated first before use with the fabian. | -set leakage -fail to approve -tubing leakage | -Failure to adhere or bond -solder joint failure -no audible alarm
Preparing for operation | Bias flow selection for NO system usage in HFO mode: Select the Bias Flow extern, the button will turn Green when selected; A WARNING message appears. Select the Confirm Box to switch to External Bias Flow. | -incorrect message display -inadequate warning -inadequate message display | -missing test results -insufficient flow or underinfusion -device maintenance issue
Preparing for operation | Bias Flow intern: Use when there is no External NO Flow sensor; Bias Flow extern: Usage of a NO device needs to measure the Bias Flow with an External Flow sensor. | -inaccurate measurement -incorrect measurement -inappropriate measurement unit | -Incorrect Measurement -Pressure Issue -fluid leak
Preparing for operation | NOTE: When switching between HFO to a Conventional Ventilation mode, the External NO Flow sensor can be left in this position. | -incorrect sensor orientation -sensor settings out of range -sensor failure | -imprecision -mechanical jam -no flow
Preparing for operation | NOTE: This bias flow setting has no impact on conventional modes but on HFO mode only. Therefore, in conventional mode, the bias flow setting can be ignored. | -incorrect mode -incorrect settings -incorrect limit setting | -imprecision -use of incorrect control settings -medical gas supply problem
Preparing for operation | WARNING: When the flow sensor is deactivated, it must be removed from the patient circuit. | -energy deactivation failed -fail to remove -fail to reset sensor | -positioning issue -unstable -pacing inadequately
Preparing for operation | WARNING: With the external bias flow setting, flow will come from the FG port, and no flow from the Insp port. | -inappropriate flow pattern -inaccurate flow rate -incorrect settings | -fluid leak -Blockage within Device or Device Component -compliance issue
Patient circuit assembly | We recommend the use of “Single Use Patient Circuits ” on the device. Best performance is achieved with dual limb heated systems. | -single use violated -incorrect circuit type -fail to activate system | -use of device issue -device operates differently than expected -inadequate user interface
Patient circuit assembly | Connect the distal Inspiratory limb to Inspiration port (Insp) on the ventilator and Connect proximal end Inspiratory limb to humidifier chamber inlet. | -port blocked -loss of connection -fail to connect | -Disconnection -loose or intermittent connection -Detachment of Device Component
Patient circuit assembly | Connect distal end heated Inspiratory limb to humidifier chamber outlet and connect proximal end heated Inspiratory limb to patient connection. | -connection broken or damaged -fail to connect -loss of connection | -Disconnection -loose or intermittent connection -failure to cycle
Patient circuit assembly | Connect distal end Expiratory limb to air filter and then connect them to Expiration port (Exp) on the ventilator and Connect proximal end Expiratory limb to patient connection. | -fail to connect -air leak -loss of connection | -loose or intermittent connection -leak -difficult to open or close
Patient circuit assembly | Connect proximal end Proximal Pressure tubing to patient connection and Connect distal end Proximal Pressure tubing to Proximal Pressure port (Prox) on ventilator. | -tubing blocked -fail to connect -tubing detached | -blockage within device or device component -Detachment of Device Component -loose or intermittent connection
Patient circuit assembly | Connect the Inspiratory and Expiratory limb heater connectors with corresponding connectors on humidifier chamber heater base. | -connector mismatched -incompatible connection -loss of connection | -fitting problem -Detachment of Device Component -Device displays error message
Patient circuit assembly | Recommended positioning of temperature probe for humidifier | -inappropriate position -position out of range -fail to monitor battery temperature | -incorrect measurement -Biocompatibility Issue -output issue
Patient circuit assembly | One of the major sources of problems in neonatal ventilation is, that excessive rainout is affecting accuracy of flow measurement and therefore compromising its use in various modes of ventilation. We recommend carefully following the manufacturer’s instructions for use of the Humidifier to minimize the risk of rainout in the Patient Circuit. | -incorrect measurement -fail to measure -inadequate measure unit | -incorrect measurement -Pressure Issue -unintended movement
Patient circuit assembly | If patient is positioned in an incubator, the following setup is recommended: Place temperature probe outside the incubator and use Incubator extension line. | -inappropriate position -position out of range -unable to place | -user used incorrect product for intended use -inadequate user interface -instruction for use issue
Patient circuit assembly | If the patient is in the open warmer unit, the following setup must be used: | -inappropriate patient type -inappropriate user -inappropriate shock | -user used incorrect product for intended use -patient data issue -break
Patient circuit assembly | If Reusable patient circuits are used, they should be used with inspiratory and expiratory limb heater wire or alternatively, a water trap in the expiratory limb: Assemble the tubing as required: Connect Inspiratory / Expiratory line and Pressure limb to Y-piece; Connect Flow-sensor to Ypiece; Connect Flow-sensor cable to Flow-sensor. | -fail to calibrate circuit compensation -lack of functional requirement -fail to configure patient information | -fitting problem -device markings issue -fluid leak
Patient circuit assembly | Prepare Humidifier according to manufacturer specifications and connect to tubing: If the Humidifier does NOT feature an Inspiratory line heater, Water Traps must be installed in the Inspiratory and Expiratory Limbs; When using an Inspiratory heater, a Water Trap must be installed at the Expiratory end. | -fail to install -fail to connect -loss of connection | -use of device issue -loose or intermittent connection -insufficient flow or underinfusion
Patient circuit assembly | NOTE: We strongly suggest using an Inspiratory and Expiratory Heated Circuit System in HFO mode. | -incorrect circuit type -circuit broken or damaged -inappropriate program selected | -failure to sense -fitting problem -Failure to Cycle
Patient circuit assembly | WARNING: Ensure that the Flow-sensor IS NOT exposed to excessive Humidification. | -over noise exposure -inappropriate indication -over flow | -device handling issue -device cleaning issue -charging problem
Patient circuit assembly | Connect nCPAP tubing set: Connect nCPAP adapter to the fabian FG / nCPAP port; Connect 22 mm adapter connector to Humidifier Chamber; Connect nCPAP system Pressure measuring limb to fabian Prox port; Connect nCPAP system to Humidifier Chamber; Apply nCPAP system to Patient according to manufacturer specifications. | -tubing malformed -tubing broken or damaged -tubing mismatched | -loose or intermittent connection -Detachment of Device Component -Disconnection
Patient circuit assembly | Connect flow sensor in NIV trigger: Preparing the nCPAP Infant Flow / Infant Flow LP / Inspire Tubing System and Connect Flow sensor to exhalation tube | -sensor damaged -fail to prepare -inadequate flow | -Pressure Issue -air leak -Leak
System start-up | Switch ON the Ventilator: Use the ON / OFF button (1) on the device to power ON the equipment; When the device is powered ON, the Start-up screen will appear, and the equipment will perform a “Self-test”; The SOFTWARE REVISION, the CHECK SUMS and the status of the interfaces will be checked; After a successful power ON, the device will prompt the user to perform the acoustic audio alarm test; In order to proceed, you must acknowledge that the alarm is audible, continue using the ventilator without audio alarm or shut down the ventilator; After the system test Is complete, the Calibration menu will appear. | -confused button -indistinct button -fail to indicate button status | -no display or display failure -device operates differently than expected -power conditioning issue
System start-up | Perform Flow sensor calibration: Select the appropriate Flow sensor (Neonatal or Pediatric) and set the patient’s body weight to see the VTe BW [mL/kg] measurement; The O2 will automatically be calibrated after leaving the Calibration menu; EtCO2 and SpO2 modules can be switched ON here as well. | -fail to calibrate manually -inaccurate measurement -incorrect measurement | -Pressure Issue -fluid leak -imprecision
Device check | Always perform a Device Check before each ventilator use or after changing the patient circuit. | -incorrect circuit type -fail to check voltage periodically -safety check failed | -Connection Issue -device operates differently than expected -output issue
Device check | The fabian ventilator does not automatically compensate gas fraction measurements for changes in barometric pressure during use. The oxygen sensor is calibrated in the hospital prior to use and this calibration remains in effect. This calibration process is part of the normal pre-use checks for the ventilator. The pre-use calibration should not be skipped if the ventilator has been moved to a location with significant change in altitude. | -fail to calibrate manually -fail to compensate -incorrect measurement | -failure to recalibrate -programming issue -use of device issue
Device check | Gas Supply check: Air and Oxygen supply hoses are correctly connected | -fail to compatible with oxygen supply -fail to update oxygen volume -oxygen delivery compromised | -medical gas supply problem -human factors issue -failure to disinfect
Device check | Breathing System (Dual limb) check: Expiration membrane holder and membrane correctly installed. Patient circuit assembled correctly according to manufacturer’s instructions | -incorrect installation -incorrect manufacturer name -fail to schedule check | -improper device output -failure to adhere or bond -audible noise
Device check | Breathing System (Single limb) check: Patient circuit assembled correctly according to manufacturer’s instructions | -incorrect circuit type -incorrect manufacturer name -fail to setup valve correctly | -failure to adhere or bond -improper device output -Ambient Temperature Issue
Device check | Switch ON test: Alarm tone is audible and alarm lamp flashes Red during switch-ON test. Self-test successful. | -loss of audible alarm -fail to generate alarm -alarm audio stuck on | -device alarm system issue -not audible alarm -Device Alarm System Issue
Device check | Calibration: Calibrate Flow sensor, Calibrate O2 sensor (occurs automatically after leaving Calibration menu). | -inaccurate sensor measurement data -fail to recover sensor -fail to calibrate manually | -Failure to Calibrate -failure to calibrate -Pressure Issue
Device check | Leakage Test: Pressure of 80 ±4 mbar is achieved on the pressure graph. Ensure that Pmanual and other pressures are reduced to safe levels prior to connecting to patient. | -inappropriate pressure level -fail to reduce pressure -pressure out of range | -pressure issue -mechanical issue -use of device issue
Device check | Infant Flow LP Leakage Test: Pressure of 5 mbar is achieved on the pressure graph | -fail to pass leak test -fail to produce pressure -low pressure | -fluid leak -Pressure Issue -Fluid Leak
Device check | Function Test: Monitor ventilation parameters to ensure the values are within the following ranges. Pinsp: 20 ±2 mbar, PEEP: 5 ±1 mbar, O2: 30 ±2 Vol % | -incorrect ventilation mode -out of range -current out of range | -device operates differently than expected -incorrect display -user used incorrect product for intended use
Device check | Alarms check: Disconnect the Inspiration tube and block the inspiratory port on the ventilator. Alarm: Tube Occlusion | -tubing blocked -fail to hold inspiratory -fail to check voltage periodically | -Failure to adhere or bond -application network issue -use of incorrect control settings
Device check | CAUTION: Ensure that Pmanual and other pressures are reduced to safe levels prior to connecting to patient. | -fail to reduce pressure -reduced wearing time -inaccurate pressure reading | -insufficient flow or underinfusion -detachment of device or device component -Failure to adhere or bond
System standby / pause | The Start / Stop button can be used to interrupt Mechanical ventilation for two minutes: In nCPAP, DUOPAP and O2 Therapy mode ventilation could be interrupted indefinitely, In both cases the built-in gas mixer delivers a minimal flow to prevent heat build-up inside the Ventilatory Gas Humidifier, Ventilation resumes following the two-minute pause, or the button is pressed again. | -ventilation interrupted -delayed button -delivery interrupted | -medication error -break -Blockage within Device or Device Component
System standby / pause | Risk of Oxygen undersupply: The Standby function is NOT intended for suctioning, A disconnection or reconnection is NOT recognized. | -loss of function -loss of sensing -loss of connection | -use of device issue -dose calculation error due to software problem -unstable
System standby / pause | The Standby – stopping / pausing Ventilation must NOT be used while ventilator is connected to the patient. | -inappropriate ventilation support -fail to maintain ventilation mode -incorrect ventilation mode | -no flow -user used incorrect product for intended use -tidal volume fluctuations
System standby / pause | Any FOT measurements taken during Standby are NOT valid | -measurement data corrupted -measurement range out of specification -low measurement accuracy | -device remains activated -Product Quality Issue -unstable capture threshold
System standby / pause | For device pause or standby: Remove patient from ventilator; Press and hold the Start / Stop button; A menu will appear on the screen stating, “Do you want to pause ventilation for two minutes?”; Press YES to pause the ventilator for 2 minutes. | -device paused -fail to activate pause -fail to disable ventilation | -Device Alarm System Issue -difficult to interrogate -no display or display failure
System standby / pause | The Information bar will display the pause in-progress message “Stop Ventilation”. When the ventilation has stopped the Information bar will display the following message “Ventilation stopped”. | -fail to display message -inadequate message display -incorrect message display | -incorrect display -failure to deploy -device remains activated
System standby / pause | The pause time will count down the two minutes and then the ventilator will resume ventilation | -fail to resume -fail to disable ventilation -fail to update time remain | -failure to deploy -device operates differently than expected -decrease in pressure
System standby / pause | Standby is only available in these modes: NCPAP, DUOPAP, O2 Therapy | -incorrect mode -inadequate operating mode -inadequate evaluation | -reset issue -use of incorrect control settings -defective item
System shutdown | Switching OFF the fabian: Remove patient from the ventilator; Power OFF any SpO2 or CO2 devices connected to the ventilator; Press and hold the ON / OFF button (1) for 6 seconds; After 6 seconds, a menu will appear on the screen stating, “Do you want to shut down the device?”; Press YES to power OFF the deice; After the device has powered OFF, a notice tone will sound. | -fail to turn on device -fail to shut-down device -sticky button | -Connection Issue -Detachment of Device Component -Disconnection
System shutdown | Press the Alarm Silence button to confirm system switch OFF. Hold the button for at least three seconds until the WARNING LED stops. | -fail to activate alarm -fail to indicate button status -button stuck on | -device alarm system issue -break -improper alarm
Emergency shutdown | WARNING: The emergency shutdown procedure should only be used during an actual emergency. This procedure disconnects all power to the ventilator and any other devices connected to the ventilator. Before performing this procedure, ensure that the alternate ventilation method (for example: resuscitation) is present, ready and standing by. | -unable to power off -incorrect procedure -fail to activate alarm | -loss of power -device difficult to setup or prepare -failure to deploy
Emergency shutdown | Emergency shutdown: Remove patient form the ventilator; Press and hold the ON / OFF button (1) for 1 second to 2 seconds, then release; After releasing, press and hold the ON / OFF button (1) again for five seconds; After the device has powered OFF, a notice tone will sound. | -fail to indicate button status -release delayed -sticky button | -Loose or Intermittent Connection -human-device interface issue -user used incorrect product for intended use
Emergency shutdown | Press the Alarm Silence button to confirm system switch OFF. Hold the button for at least three seconds until the WARNING LED stops. | -fail to activate alarm -fail to indicate button status -button stuck on | -device alarm system issue -break -improper alarm
Configurations menu | The Configuration menu can be accessed by pressing the Menu / Calibration button on the keypad and selecting the Menu push button. | -fail to configure infusion parameters -configuration confused -fail to setup configuration | -inadequate user interface -human-device interface issue -failure to deploy
Calibration | The Calibration screen is the first screen that comes up when the Calibration key is pressed. Then you can proceed through the menu options to select the Calibration menu again | -fail to calibrate manually -incorrect screen display -fail to program dosage | -No Display or Display Failure -aborted charge -excess flow or over-infusion
Calibration | The Calibration menu can be accessed through the Menu / Calibration button on the keypad. | -fail to access -confused button -fail to program dosage | -human-device interface issue -unable to obtain readings -failure to deploy
Calibration | Upon ventilator startup, you will be automatically directed to the Calibration menu. The Calibration menu is displayed as follows. | -fail to request calibration -fail to guide alignment -table incrementation incorrect | -incorrect display -device operates differently than expected -Power Source Issue
Calibration | From the Calibration screen, you can calibrate the following Sensors: Neonatal and Pediatric Flow sensors; O2 sensor; EtCO2 module; SpO2 module | -fail to calibrate manually -calibration failed -fail to recover sensor | -incorrect measurement -unstable -Inflation Issue
Calibration | Flow sensors can be calibrated in the upper half of the Calibration menu. | -fail to calibrate manually -out of calibration -calibration failed | -imprecision -Pressure Issue -defective item
Calibration | Only use the Neonatal Flow sensor for this calibration. To calibrate a Neonatal Flow sensor, select the Neonatal button. | -incorrect sensor orientation -inaccurate sensor measurement data -inaccurate measurement | -Pressure Issue -Inflation Issue -Failure to Calibrate
Calibration | Only use the Pediatric Flow sensor for this calibration. To calibrate a Pediatric Flow sensor, select the Pediatric button. | -incorrect sensor orientation -inaccurate sensor measurement data -inaccurate measurement | -Pressure Issue -Failure to Calibrate -Device Operates Differently than Expected
Calibration | The Calibration procedure is identical for both types of flow sensors and described below: Ensure that gas will NOT flow through the Flow sensor during Calibration; The sensor can be held occluding either one or both ends with a sterile glove to ensure no gas flow. | -fail to calibrate manually -incorrect test procedure -fail to establish test criteria | -Pressure Issue -incorrect measurement -use of device issue
Calibration | Press the Flow Cal button; First “Checking” will appear, then “Calibration running”; The Red lines in the diagram move from left to right, until they are centered in this diagram. When centered, these lines will turn Green; Following successful Calibration, the Date of the last calibration will be displayed in the corresponding area. | -fail to flow check -fail to indicate button status -indistinct button | -defective item -defective component -device maintenance issue
Calibration | The Flow Sensor Calibration needs to be performed each time: A new sensor is put in place; After device start-up; After enabling a Flow sensor (manual and automatic); After patient range change; After reconnection of a flow sensor when resolving a disconnection alarm; We recommend cleaning the Flow Sensor once daily. | -fail to calibrate manually -fail to request calibration -fail to reset sensor | -Pressure Issue -output issue -device maintenance issue
Calibration | If Zero Flow and Tidal Volume measurements are detected in numeric and graphs after: Disconnection; High Leaks; Low-level Rain-out; Over-breathing. Flow measurement will be automatically re-started after 15 seconds. | -fail to set tidal volume limit -fail to update tidal volume -fail to calculate tidal volume | -use of device issue -Pressure Issue -insufficient flow or underinfusion
Calibration | If Flow and Tidal Volume measurements are not regained, consider the following: Checking ET tube position (high leaks at the patient) and blockage; Flow Sensor Leakage (loose flow insert or leak); Flow Sensor Contamination (single-use sensor recommended); Rain-out (contact ACUTRONIC Medical Systems sales for recommended Patient Circuits and Humidifier to prevent Rain-out) | -fail to update tidal volume -fail to deliver tidal volume -loss of measurement | -fluid leak -Pressure Issue -improper flow or infusion
Calibration | The O2 sensor calibration can be accessed in the Configurations menu, by pushing the O2 button. | -fail to request calibration -fail to access -fail to setup configuration | -dose calculation error due to software problem -medical gas supply problem -no display or display failure
Calibration | The O2 sensor is automatically calibrated when the machine starts up and every 24 hours when in use. During Calibration procedure, the FiO2 concentration of the fresh gas to the patient is NOT altered. In case of a “supply Gas failure”, the oxygen sensor calibration is automatically disabled to avoid false sensor calibration. | -fail to calibrate manually -out of calibration -calibration failed | -pacing inadequately -unstable -calibration error
Calibration | Manual Calibration: Press the 21% button to perform One Point Calibration at 21% O2; Press the 21 + 100% button to perform a Two Point Calibration at 21% and 100% O2. After one of these buttons has been pressed, the Calibration procedure starts. If the procedure has been successful, a Green check mark will appear. | -fail to calibrate manually -fail to indicate button status -button stuck on | -output issue -failure to calibrate -human-device interface issue
Calibration | The O2 sensor can be turned OFF in the event that the O2 sensor reaches end of life during use. You are then expected to use a secondary respiratory gas monitor until the ventilator can safely be taken out of service to replace the O2 sensor. | -fail to turn on device -fail to turn on automatically -loss of sensor data | -device expiration issue -use of device issue -Pressure Issue
Calibration | After a successful calibration, the Date of the last calibration is displayed in the corresponding area. | -incorrect message display -incorrect display format -fail to display message | -no display or display failure -incorrect display -reset issue
Calibration | Automatic Sensor Calibration runs every time the equipment Restarts and every 24 hours. The gas concentration to patient is not altered during this calibration procedures. | -fail to request calibration -fail to calibrate manually -fail to check motor currect periodically | -Product Quality Issue -Failure to Sense -failure to recalibrate
Calibration | You can set the patient’s body weight in the Calibration menu. This setting is used to calculate and display the VTe BW [mL / kg] Numeric measurement in IPPV and SIPPV modes. | -fail to calibrate manually -loss of measurement -inadequate setting option | -use of incorrect control settings -incorrect display -failure to deploy
Display | Adjust screen brightness, four levels; Lock Touch screen. The screen will automatically be unlocked if an alarm is triggered or the Rotary Pulse Encoder is pressed; The Touch screen automatically locks after a set time; Display Graphs as filled or regular lines, Auto scale ON / OFF; Display Trends as filled or regular lines, Auto scale ON / OFF | -fail to reset alarm automatically -fail to unlock -insufficient brightness | -No Display or Display Failure -no display or display failure -incorrect display
Display | Lock Touch screen: 1. In the Touch area of the Display screen, press “lock” and the screen is locked. NOTE: To deactivate the lock, press the Rotary push knob. 2. Program automatic lock: slide bar to the right and the Time will be displayed. 3. Select Time and you can choose the appropriate time from the auto screen lock list. | -unable to lock screen -fail to lock -knob stuck | -no display or display failure -incorrect display -loss of data
Display | The touch screen can also be locked by holding the Rotary push knob for 3 seconds. Lock is automatically cancelled in the event of an alarm or by pressing the Rotary push knob. | -unable to lock screen -fail to lock -lock mechanism failed | -use of device issue -inadequate user interface -failure to sense
Display | Trend / graph display: Curve plotting in Graphics is automatically adjusted. Trend patterns are automatically scaled. | -inaccurate trend result -lack of trend analysis -incorrect plot | -device markings issue -user used incorrect product for intended use -air leak
Ventilation parameter settings | Automatic O2 Calibration: 21 Vol% / 21 and 100 Vol.%. In cases where oxygen supply source does NOT provide 100% oxygen, this should be set to 21% only to avoid calibration at false value. | -fail to compatible with oxygen supply -fail to update oxygen volume -fail to request calibration | -medical gas supply problem -excess flow or over-infusion -no flow
Ventilation parameter settings | NIV Leak Compensation: Set either Low, Middle or High level of NIV leak compensation (≤ ~15, 30, 40%). | -fail to compensate -improper compensation -low rate | -fluid leak -overheating of device or device component -improper device output
Ventilation parameter settings | Unit for Pressure: Selection of Pressure unit. | -inaccurate pressure reading -inappropriate pressure level -fail to adjust pressure | -incorrect measurement -Difficult to Position -application program issue
Ventilation parameter settings | NIV Tube Set: Selection of the nCPAP system used (Infant Flow, Infant Flow LP, Medijet®). | -inadequate flow -selection failed -incorrect tubing placement | -use of device issue -application network issue -compliance issue
Ventilation parameter settings | Maximum time for Manual breath: Inspiration time for Manual breath (2 to 30 seconds). | -long response time -incorrect breath type -long charge time | -improper device output -failure to deliver -misassembled
Ventilation parameter settings | Alarm Delay: After visual reset of alarm, the reactivation of alarm is delayed for the set time period. This avoids an immediate re-alarm in case Operational parameters have not yet stabilized. | -alarm delayed -fail to reset alarm automatically -fail to disable alarm | -device alarm system issue -use of device issue -Device Alarm System Issue
Ventilation parameter settings | Manual Breath HFO: The Manual Breath in HFO mode can be used to perform Lung Recruitment Maneuver (sustained lung inflation) if clinically accepted by internal guideline. Length and Pressure level are adjustable. | -fail to initiate breath -inadequate breath type -incorrect breath type | -use of device issue -failure to deploy -improper device output
Ventilation parameter settings | Trigger: Flow Trigger / Volume Trigger / Pressure Trigger. Trigger mode Configuration for assisted breathing | -inappropriate pressure level -incorrect device orientation -incorrect limit setting | -deployment issue -programming issue -excess flow or over-infusion
Ventilation parameter settings | E-flow: E-flow parameter setting | -incorrect parameter -incorrect settings -inappropriate flow pattern | -use of incorrect control settings -device difficult to setup or prepare -failure to deploy
Ventilation parameter settings | Ratio of I-time: I-time Display Configuration. Setting I time or frequency (Frequency - I-time / I-time / E-time). | -inadequate configuration -inaccurate settings -incorrect settings | -use of incorrect control settings -incorrect display -poor quality image
Ventilation parameter settings | Ppsv Parameter: Select Ppsv parameter to be absolute value or above PEEP Parameter. | -incorrect parameter -inadequate evaluation -invalid input | -use of incorrect control settings -pressure issue -inaccurate delivery
Ventilation parameter settings | Use BTB for Vt: Select if Tidal Volume measurement should be updated for every breath. | -fail to update tidal volume -fail to compensate tidal volume -lack of update | -use of device issue -device markings issue -output issue
Ventilation parameter settings | Hospital Settings: Resetting the device to hospital default settings. Optional: The “Hospital Settings” button is only active if it is enabled in SERVICE mode. | -device reset incorrectly -fail to restore to factory default state -incorrect settings | -human-device interface issue -use of incorrect control settings -user used incorrect product for intended use
Ventilation parameter settings | Factory Defaults: Resetting the device to the factory defaults. Optional: The “Factory Defaults” button is only active if it is enabled in SERVICE mode. | -fail to restore factory defaults -fail to restore to factory default state -device reset incorrectly | -use of incorrect control settings -human-device interface issue -user used incorrect product for intended use
Ventilation parameter settings | WARNING: Only reset to hospital and factory settings when the device is NOT connected to patients. | -device reset incorrectly -fail to restore factory defaults -inappropriate reset | -use of incorrect control settings -failure to deliver -insufficient flow or underinfusion
Patient data | Entering Patient data. 1. Reset : Clearing the actual patient data; 2. Apply : will store patient data from now; 3. Change data : is used to add new patient data. Deleting all stored data from the previous patient | -undesired patient data changes -fail to clear data -fail to retrieve patient diagnostics data | -patient data issue -data issue -application security issue
Patient data | For data input, just push “Change data” and choose the field for input. For capital letters, push Accept data with “Apply”. | -invalid input -incorrect data -inappropriate data flow | -data issue -problem with software installation -application security issue
Language | Language: 1. Selecting the operator Language. 2. Directly choose a language by touching the language. 3. For scrolling through the list, touch the Up and Down arrows. 4. Restart the device after selecting a new language. | -selection failed -inappropriate interface design -fail to incorporate | -failure to sense -failure to analyze rhythm -human-device interface issue
Date / Time | Changing the Date and Time: 1. Use the Up and Down arrows to set the date and time. 2. After the time or date has been changed, the confirmation check mark changes colors. 3. The new information is applied after confirmation. | -inappropriate color scheme -inappropriate material color -inappropriate inspiratory time | -use of device issue -use of incorrect control settings -failure to deliver
Tools | Only available with USB stick connected. | -incompatible connection -loss of connection -fail to connect | -therapeutic or diagnostic output failure -Component Incompatible -component or accessory incompatibility
Tools | Trend Data will be output in CSV format. Log Files, Alarm and System log, Device Information, SW and HW Configuration with licenses. | -inadequate log events -incorrect output -inadequate data | -data issue -data back-up problem -incompatibility problem
Tools | export of trends and log files are only allowed when patient is NOT connected | -lack of trend analysis -fail to trend data or parameter -unable to connect | -user used incorrect product for intended use -insufficient flow or underinfusion -patient data issue
Information | Displays the Information screen (System Information) containing Equipment data. | -incorrect screen display -fail to display data -fail to display message | -incorrect display -No Display or Display Failure -no display or display failure
Service mode | The Service mode is “password” protected. | -inadequate service -lack of protective measure -lack of use environment | -use of incorrect control settings -instruction for use issue -not audible alarm
Service mode | Service mode provides access to Nurse Call activation / deactivation, etc. | -fail to enter mode -unable to access -loss of call | -use of device issue -Difficult to Position -human-device interface issue
Service mode | This section is described in the Service Manual and is only for trained Service technicians. | -inadequate service -inadequate training -lack of training | -instruction for use issue -therapeutic or diagnostic output failure -use of incorrect control settings
Service mode | After using service mode and before connecting fabian to a patient, restart the device. | -device reset incorrectly -fail to reset -energy deactivation failed | -use of incorrect control settings -instruction for use issue -inadequate training
Service mode | service mode only allowed when patient is NOT connected | -fail to enter mode -fail to connect -unable to connect | -device operates differently than expected -user used incorrect product for intended use -no flow
Alarm limits menu | Alarm limits menu: 1. Automatically sets individual alarm limits. 2. Alarm volume (three levels). 3. Open Log file. 4. Display measurements with alarm limits. | -fail to set limit correctly -incorrect limit setting -fail to reset alarm automatically | -device alarm system issue -Device Alarm System Issue -no audible alarm
Alarm limits menu | Access the Alarm Limits menu by pressing the Alarm Limits (1) button. | -fail to limit flow -fail to raise low priority alarm -fail to silence alarm | -human-device interface issue -device alarm system issue -failure to sense
Alarm limits menu | The Alarm Limits menu allows all alarm limits to be manually or automatically adjusted. | -fail to adjust manually -fail to raise low priority alarm -incorrect limit setting | -device alarm system issue -Device Alarm System Issue -break
Alarm limits menu | In the event of an alarm, a notice appears in the Information bar and a crossed out bell symbol indicates the alarm cause. | -false alarm -invisible indication -unnoticeable alarm | -fluid leak -improper or incorrect procedure or method -device markings issue
Alarm limits menu | In addition, an audible alarm sounds, which can be silenced for two minutes by pressing the Alarm Silence button. | -loss of audible alarm -fail to silence alarm -sound mute failed | -improper alarm -Device Alarm System Issue -alarm not visible
Alarm limits menu | The Alarm Limit can be adjusted manually (1). The Alarm Limits menu automatically focuses on the corresponding parameter. | -fail to reset alarm automatically -fail to adjust manually -fail to raise low priority alarm | -device difficult to setup or prepare -break -device alarm system issue
Alarm limits menu | The limit can also be automatically adjusted by pressing the Autoset button (2). There is a slight delay before the Autoset alarm appears | -incorrect limit setting -unable auto adjustment -adjustment failed | -inadequate user interface -human-device interface issue -not audible alarm
Alarm limits menu | After the Alarm Limit has been adjusted or there are no more active alarms, the bell symbol will turn Grey. The Visual Alarm message must be reset by pressing the Alarm Silence button. | -fail to set or reset alarm manually -fail to set or reset alarm automatically -fail to sort alarm message | -device alarm system issue -not audible alarm -break
Alarm limits menu | Alarm conditions: Each displayed breath data is checked for violations of the current alarm limits. Additionally, the measured pressure will be checked against the limit every 20 milliseconds with the following conditions: | -incorrect breath type setting -incorrect breath type -fail to set or reset alarm automatically | -incorrect display -no display or display failure -Device Alarm System Issue
Alarm limits menu | Upper Alarm Limit of Pressure: • CPAP: with a delay of 300 milliseconds in case of a manual or mechanic breath -> pressure will be checked against upper limit, if the pressure is still above the limit after 5 seconds the alarm will be signalled, otherwise it is cleared. • nCPAP: with a delay of 3 seconds in case of a manual or mechanic breath -> pressure will be checked against upper limit, if the pressure is still above the limit after 5 seconds the alarm will be signalled, otherwise it is cleared. • All other ventilation modes: pressure will be checked against the upper limit, if the limit is exceeded the alarm will be signaled. • A breath with peak pressure below the upper limit will reset the delay of 5 seconds. | -incorrect limit setting -fail to set limit correctly -fail to check voltage periodically | -use of incorrect control settings -pressure issue -defective item
Alarm limits menu | Lower Alarm Limit of Pressure: • A breath with peak pressure above the lower limit will reset the delay set in the ventilation menu. • nCPAP: with a delay of 3 seconds in case of a manual or mechanic breath -> pressure will be checked against lower limit, if the pressure is still below the limit after the alarm delay set by the operator in the Ventilation menu, the alarm will be signaled, otherwise it is cleared. • DUOPAP: pressure will be checked against lower limi t, if the pressure is still below the limit after the Alarm delay set by the operator in the Ventilation menu, the alarm will be signaled, otherwise it is cleared. • If an Alarm Condition has been detected, a dedicated Alarm Array is written, otherwise it is cleared | -fail to set or reset alarm manually -fail to reset alarm automatically -fail to set limit correctly | -use of incorrect control settings -pressure issue -defective item
Alarm limits menu | Automatic alarm presets: The following alarm limits autosets are used as defaults. HFO: mean airway pressure Lower limit: 5 mbar above set mean Airway Pressure; Upper limit: 5 mbar above set mean Airway Pressure | -fail to set limit correctly -fail to adjust settings -fail to activate occlusion alarm | -use of device issue -use of incorrect control settings -device alarm system issue
Alarm limits menu | Configurable alarms: Following alarm limits are displayed and can be configured in the Alarm Limits menu. • High SpO2 alarm : 2 to 99%, off | -fail to configure vibrating alarm -fail to view alarm settings -inappropriate safe limit | -device alarm system issue -Device Alarm System Issue -break
Alarm limits menu | Alarm log: Switch to Alarm Limit menu. Press the Show Log key to access the Alarm Overview. Press the Hide Log key to return to the Alarm Limits menu | -inadequate log events -fail to activate alarm -fail to access | -unauthorized access to computer system -device alarm system issue -defective item
Alarm limits menu | The Alarm List is sorted in chronological order. The parameters shown are: • Time • Message • Priority | -fail to sort alarm message -inadequate message display -inadequate alarm states | -reset issue -Reset Issue -device alarm system issue
Alarm limits menu | Switching OFF the device will erase the Alarm Log. | -fail to update log file -inadequate log events -fail to sort alarm message | -device operates differently than expected -Device Alarm System Issue -Reset Issue
Alarm causes and solutions | Alarms are categories by three main priority levels; High, Medium and Low. They differ visually and acoustically according to priority. | -inadequate alarm states -false alarm -confused warning | -device alarm system issue -decreased sensitivity -solder joint failure
Alarm causes and solutions | Alarms table: This table is intended to help you determine and resolve the cause of an alarm message. Alarm Text(Cooling fan defect), Alarm Type(System Failure), Cause(Fan not moving), Solution(Contact your local ACUTRONIC Medical Systems Distributor), Main Priority(I), Sub Priority(1) | -fail to sort alarm message -fail to generate alarm -fail to clear alarm | -device alarm system issue -Device Alarm System Issue -use of device issue
Alarm causes and solutions | Alarm Text(Tube Occlusion), Alarm Type(Breathing Circuit alarms), Cause(• Expiration tube blocked/kinked • Inspiration tube blocked/kinked • Excessive deviation between inspiration and expiration pressure sensor.), Solution(• Check ventilation tubes. • Check pressure sensor connection (proximal). • Contact your local ACUTRONIC Medical Systems Distributor), Main Priority(I), Sub Priority(2) | -inadequate alarm states -false alarm -inappropriate alert type | -device alarm system issue -low audible alarm -defective item
Alarm causes and solutions | Alarm Text(O2 value out of range), Alarm Type(Sensor alarms), Cause(O2 level is out of range), Solution(Calibrate O2 sensor and / or Contact your local ACUTRONIC Medical Systems Distributor.), Main Priority(I), Sub Priority(2) | -inadequate alarm states -false alarm -low alarm volume | -failure to sense -Device Alarm System Issue -capturing issue
Alarm causes and solutions | Alarm Text(Apnea alarm), Alarm Type(Patient alarms), Cause(No spontaneous patient breathing detected.), Solution(• Check patient, • Switch to controlled ventilation.), Main Priority(I), Sub Priority(6) | -false alarm -fail to activate alarm -inadequate alarm states | -device alarm system issue -prompts will not clear -defective item
Alarm causes and solutions | Alarm Text(Low Pmean), Alarm Type(Patient alarms), Cause(Leak or disconnection), Solution(• Check patient connection. • Check tubing system for leaks. • Check setting of Pmean.), Main Priority(I), Sub Priority(7) | -false alarm -inappropriate alert type -inadequate alarm states | -output issue -unauthorized access to computer system -low battery
Alarm causes and solutions | Alarm Text(High PIP), Alarm Type(Patient alarms), Cause(Pressure rise in tubing system; mechanical inspiration was reduced to relieve the system.), Solution(• Check patient, • Check tubing system. • Replace patient system.), Main Priority(I), Sub Priority(7) | -inadequate alarm states -lack of audible alarm -false alarm | -break -Device displays error message -defective item
Alarm causes and solutions | Alarm Text(High SpO2), Alarm Type(Patient alarms), Cause(SpO2 measured above set limit), Solution(• Check patient, • Check SpO2 sensor for correct contact with patient • Adjust settings), Main Priority(II), Sub Priority(8) | -fail to set limit correctly -inadequate alarm states -false alarm | -output issue -inaudible voice prompts -improper alarm
Alarm causes and solutions | Alarm Text(High Pulse Rate), Alarm Type(Patient alarms), Cause(Pulse rate measured above set limit), Solution(• Check patient, • Check SpO2 sensor for correct contact with patient • Adjust settings), Main Priority(II), Sub Priority(8) | -fail to set or reset alarm automatically -inadequate alarm states -low rate | -output issue -inaudible voice prompts -improper alarm
Alarm causes and solutions | Alarm Text(Charge battery (<60min)), Alarm Type(Electrical power), Cause(• Ventilator is not connected to mains supply. • Power failure. Internal battery remaining charge less than 60 minutes.), Solution(Switch to mains supply. Alarm can be silenced with Alarm Silence button. The device continues in battery mode without interruption.), Main Priority(II), Sub Priority(8) | -fail to activate alarm -false alarm -low battery | -defective component -device-device incompatibility -failure to charge
Alarm causes and solutions | Alarm Text(High tube leak), Alarm Type(Patient alarms), Cause(Upper alarm limit exceeded), Solution(• Check patient, • Check tubing system for leaks, • Adjust limit value.), Main Priority(II), Sub Priority(8) | -inadequate alarm states -low alarm volume -alarm notified | -device alarm system issue -low audible alarm -output issue
Alarm causes and solutions | Alarm Text(Tidal Volume limited), Alarm Type(System limits), Cause(Upper alarm limit exceeded), Solution(Adjust settings.), Main Priority(III), Sub Priority(9) | -fail to set or reset alarm automatically -inadequate alarm states -false alarm | -device alarm system issue -Device Alarm System Issue -low audible alarm
Alarm causes and solutions | Pressure release behavior: In case of detected obstruction of the breathing tubes, the ventilator alarms and also ceases gas flow to the outlet and opens the exhalation valve to reduce the pressure in the breathing circuit to ambient pressure (ZEEP). The tube obstruction alarm is terminated once the airway pressure measured by the proximal pressure sensor has fallen below set PEEP and the time of the next breath trigger is reached. | -fail to reduce pressure -fail to activate alarm -valve stuck | -incorrect measurement -pressure issue -Pressure Issue
Alarm causes and solutions | This pressure relief occurs: • 12 mbar above the set PIP in IPPV, SIPPV, SIMV, SIMV+PSV, PSV and CPAP modes. • 10 mbar above the set PIP in NCPAP and DUOPAP modes | -inaccurate pressure reading -low pressure -pressure control failed | -use of incorrect control settings -pressure issue -programming issue
Alarm causes and solutions | In case of detected high inspiratory pressure, the ventilator alarms and switches to the expiratory phase – the exhalation valve pressure control is reduced to PEEP, and the outlet gas flow is switched from I-Flow to E-Flow. Only if this subsequently transitions to a tube obstruction alarm, the outlet flow is switched off and exhalation valve opened as above | -fail to reduce pressure -pressure control failed -fail to raise high priority alarm | -difficult to open or close -material integrity issue -Failure to Calibrate
Alarm causes and solutions | Application error: If the device software encounters an unexpected ERROR, a window appears in the middle of the screen with the following text in English: “Application fabianHFO.exe encountered a serious Error and must shut down.” | -error message -incorrect error code -unexpected reset | -no display or display failure -reset issue -material disintegration
Alarm causes and solutions | Application error: If there is an audible alarm, follow these steps: 1. Immediately start ventilation of patient with an Alternative Ventilation method 2. Shut down the device. 3. Report the incident to your local distributor 4. Use the device only if it is inevitable | -loss of audible alarm -device alarmed -alarm notified | -leak -no audible alarm -fluid leak
Alarm causes and solutions | Application error: If there is no accompanying audible alarm, follow these steps: 1. Stop interacting with the device 2. Prepare an alternative treatment for the patient while the device is kept under supervision 3. Start ventilation of patient with an Alternative Ventilation Method 4. Shut down the device. 5. Report the incident to your local distributor 6. Use the device only if it is inevitable | -loss of audible alarm -inadequate alarm states -fail to clear alarm | -user used incorrect product for intended use -failure to sense -incorrect display
Alarm causes and solutions | Watchdog alarms: There is a watchdog circuit, which monitors for failures of the following components: • NetDCU (WDST#4 error) • Control PIC (WDST#16 error) • Monitor PIC (WDST#32 error) • Alarm PIC (WDST#64 error) | -fail to detect error -fail to maintain failure records -data back-up failed | -protective measure issue -device maintenance issue -defective component
Alarm causes and solutions | Watchdog alarms: If a system failure occurs, the watchdog opens the pressure relief valves and stops power to the blender module to prevent and/or relieve pressure in the patient circuit while alarming. | -valve stuck open -fail to close valve -fail to activate alarm | -loss of power -calibration issue -no pressure
Alarm causes and solutions | Watchdog alarms: In case of an occurrence, follow these steps: 1. Immediately start ventilation of patient with an Alternative Ventilation method 2. Shut down the device. 3. Report the incident to your local distributor 4. Use the device only if it is unavoidable | -device alarmed -inappropriate ventilation support -fail to activate alarm | -Imprecision -imprecision -unstable
Alarm causes and solutions | If a WDST 4/16/32/64 error is resolved (was only a temporary error) WDST#8 will be triggered and displayed and the device can be further used. NOTE: WDST#8 is not an alarm and only shows that all watchdog circuits are working correctly | -incorrect message display -fail to resolve problem -fail to activate alarm | -Device Alarm System Issue -user used incorrect product for intended use -device difficult to maintain
Battery operation | The device features an Internal Battery. Battery Run Time in HFO mode on full charge : approx. 1 hour; Battery Run Time in Conventional mode on full charge : approx. 2.5 hours | -fail to charge battery to full -low battery -device burned | -charging issue -failure to charge -battery issue
Battery operation | These values may vary depending on the parameter settings. The device continuously calculates the remaining Run Time and displays it in the Information bar. The Battery charge is continuously monitored. | -incorrect display value -incorrect parameter -incorrect value | -no display or display failure -battery issue -failure to reset
Battery operation | WARNING: A reduction in Battery life when changing the Operating mode can cause an Automatic shutdown of the ventilator! | -reduced battery life cycles -fail to provide low battery indicator -internal battery short | -improper or incorrect procedure or method -delivery system failure -inaccurate delivery
Battery operation | Changing the Operating mode can greatly shorten battery life. • Monitor the remaining Battery life when changing the settings. • NEVER leave the device and patient unattended during battery operation. • Establish mains power supply in good time! | -reduced battery life cycle -reduced battery life cycles -low battery | -Battery Issue -battery issue -premature discharge of battery
Battery operation | In the event of Power Failure, the pneumatic system will automatically open against atmospheric pressure so that pressure can’t build in the breathing system, allowing spontaneous patient breathing. | -fail to open circuit -fail to turn on automatically -unable auto adjustment | -no pressure -device operates differently than expected -Device Alarm System Issue
Battery operation | When operating the device on the Internal Battery, the HFO amplitude is limited to 60 mbar | -fail to limit flow -low readings -inappropriate interval settings | -device operates differently than expected -failure to read input signal -electronic property issue
Battery operation | WARNING: If you are using HFO mode with higher amplitudes, disconnecting from the main power supply will automatically decrease the amplitude to 60 mbar. When the main power supply is reconnected, the amplitude remains at 60 mbar. | -inaccurate amplitude measurement -fail to activate alarm -fail to update pressure | -Power Source Issue -detachment of device component -failure to power-up
Battery operation | When storing the ventilator for long term storage, please remove the Batteries. | -fail to run on battery -inappropriate storage environment or condition -stored inappropriately | -self-activation or keying -improper alarm -Power Source Issue
Battery operation | In the event of a Power Failure, the device automatically switches to Battery Power without interruption. A visual and audible message is triggered. “Power Failure” This message can be confirmed with the Alarm Silence button. | -fail to turn on automatically -fail to run on battery -fail to shut-down device | -device maintenance issue -no display or display failure -component or accessory incompatibility
Battery operation | The device will continuously monitor the battery run time. | -fail to monitor battery temperature -fail to run on battery -fail to measure battery voltage | -material integrity issue -Battery Issue -failure to shut off
Battery operation | Additional messages will be triggered at a remaining Run Time of: • 15 minutes • 30 minutes • 60 minutes, only if the remaining run time is less than 60 minutes during a Power Failure | -fail to activate message -incorrect message display -loss of message | -device alarm system issue -electrical power problem -data issue
Battery operation | Depleted Battery: If the remaining Run Time is less than one minute, a warning will appear prompting to immediately switch to a Primary Power Source. After a total of 60 minutes on Battery power, the Battery is drained, and fabian will power OFF. | -fail to run on battery -fail to generate warning -warning displayed | -device alarm system issue -mechanical issue -no display or display failure
Operating on External Power Source | The device can also be operated from an external power source. | -power source incompatibility -power source detection issue -backup power activation failed | -Difficult to Position -Power Source Issue -failure to power-up
Operating on External Power Source | In this case connect the external 24 Volt power source at the back of the equipment. Conventional mode: 24 Volt DC, 1 A; HFO mode: 24 Volt DC, 1.5 A | -improper power configuration -loss of connection -loss of power | -failure to power-up -spark -Power Source Issue
Operating on External Power Source | External power connection on the back panel. 1/2 minus; 3/4 plus | -power consumption out of range -loss of power -power out of range | -use of incorrect control settings -Difficult to Position -battery issue
Operating on External Power Source | • NO charging when connected to 24V DC external power supply • NO monitoring of remaining time on external power supply | -fail to monitor power event -power supply malfunction -loss of power | -Power Source Issue -failure of device to self-test -aborted charge
Ventilation parameters | Ventilation Parameters: Parameter(Backup), Mode(CPAP), Description(Backup Frequency; In the event of Apnea, spontaneous breathing is stimulated after the apnea duration with a default number of mechanical breaths. Flow Sensor must be connected and activated. Backup Trigger depends on Apnea Period Setting.) | -incorrect parameter -parameter confused -parameter setting error | -use of incorrect control settings -failure to sense -improper device output
Ventilation parameters | Ventilation Parameters: Parameter(Backup Rate), Mode(• PSV • DUOPAP), Description(Backup Rate in PSV and DUOPAP modes) | -incorrect parameter -parameter confused -inappropriate method | -improper or incorrect procedure or method -use of incorrect control settings -programming issue
Ventilation parameters | Ventilation Parameters: Parameter(CPAP), Mode(• DUOPAP • CPAP • NCPAP), Description(Continuous Positive Airway Pressure; Continuous flow for producing Airway pressure at the CPAP Level.) | -incorrect parameter -parameter confused -inappropriate method | -fluid leak -Device Operates Differently than Expected -Blockage within Device or Device Component
Ventilation parameters | Ventilation Parameters: Parameter(E-flow), Mode(• IPPV • SIPPV • SIMV • SIMV + PSV • PSV), Description(Expiratory flow (base flow) The continuous expiratory flow is variable separate from the inspiratory flow. • Flushing dead space in the Y-piece • Maintaining the PEEP) | -incorrect parameter -parameter confused -parameter setting error | -fluid leak -pacing inadequately -Output Issue
Ventilation parameters | Ventilation Parameters: Parameter(O2), Mode(All modes), Description(Oxygen Concentration, Inspiratory Oxygen Concentration Setting.) | -parameter setting error -incorrect parameter -parameter confused | -no flow -medical gas supply problem -medication error
Ventilation parameters | Ventilation Parameters: Parameter(O2 Flush), Mode(All modes), Description(Oxygen Flush, Perform O2 Flush / Oxygen spray (Preoxygenation for maximum of two minutes)) | -incorrect parameter -parameter confused -parameter setting error | -medical gas supply problem -air leak -excess flow or over-infusion
Ventilation parameters | Ventilation Parameters: Parameter(Trigger), Mode(• SIPPV • SIMV • SIMV + PSV • PSV • DUOPAP), Description(There are three Trigger functions: Volume trigger, Flow trigger, Pressure trigger) | -incorrect parameter -parameter confused -parameter setting error | -fluid leak -use of device issue -Pressure Issue
Ventilation parameters | Parameters in HFO mode: Parameter(Flow), Description(Constant Flow / Bias, Adjustable from 5 to 20 L/min, Default value: 8 L/min) | -incorrect parameter -parameter setting error -parameter confused | -use of incorrect control settings -use of device issue -device difficult to setup or prepare
Ventilation parameters | Parameters in HFO mode: Parameter(Freqrec), Description(Frequency Lung Recruitment* Adjustable 5 to 20 Hz.) | -incorrect parameter -parameter setting error -content out of specification | -use of incorrect control settings -device difficult to setup or prepare -use of device issue
Ventilation parameters | Parameters in HFO mode: Parameter(P mean), Description(Mean Pressure, Continuous positive Airway Pressure.) | -incorrect parameter -parameter confused -current out of range | -reset issue -use of incorrect control settings -programming issue
Ventilation modes | The ventilator is intended for the following Ventilation modes: IPPV, SIPPV, SIMV, SIMV+PSV, PSV, CPAP, NCPAP, DUOPAP, HFO, O2 Therapy | -fail to maintain ventilation mode -incorrect ventilation mode -inappropriate ventilation support | -reset issue -device operates differently than expected -pacing inadequately
Ventilation modes | In Intermittent Positive Pressure Ventilation (IPPV) mode Ventilation is performed at the specified patterns set on the equipment without consideration to possible spontaneous patient respiration. The Set Rate is the Delivered Rate. | -incorrect ventilation mode -inappropriate ventilation support -fail to disable ventilation | -No Known Device Problem -detachment of device component -human-device interface issue
Ventilation modes | IPPV: Set following parameters: • E‒flow • Frequency (Rate) • I‒flow • I‒time • PEEP • Pinsp • Positive pressure is created here during Inspiration with Passive Expiration. • This mode should only be used if no spontaneous breathing from patient is expected. | -incorrect parameter -incorrect settings -expiration date error | -use of incorrect control settings -deployment issue -device operates differently than expected
Ventilation modes | IPPV can be applied also in Non-Invasive mode (Nasal IPPV or NIPPV), for example: with oronasopharynx tube or RAM Cannula; in this case you must turn OFF and remove the flow sensor from the Y-piece of the dual limbed breathing circuit and connect the proper interface. A higher I-flow can be requested during nIPPV mode, because of potential higher leakages. | -fail to produce pressure -lack of specification -loss of sensing | -device difficult to setup or prepare -positioning issue -use of device issue
Ventilation modes | Synchronized Intermittent Positive Pressure Ventilation (SIPPV) each spontaneous patient inspiratory effort triggers a mechanical breath of the ventilator according to the ventilator parameters set for Inspiratory Period and Pressure. | -incorrect breath type -pressure out of range -fail to initiate breath | -device operates differently than expected -failure to cycle -delivery system failure
Ventilation modes | SIPPV: The number of breaths per minute supported by the ventilator is controlled by the patient. In this mode the patient breathing frequency must be closely monitored to prevent hyperventilation. | -incorrect breath type -frequency out of range -inadequate breath types | -inaccurate delivery -calibration issue -use of device issue
Ventilation modes | SIPPV: Set the Alarm for high Frequency according clinical guidelines. Set following parameters: • E‒flow. • Frequency (Rate) • I‒flow • I-Time • PEEP • Pinsp • Trigger Sensitivity | -fail to set or reset alarm automatically -improper alarm -inadequate alarm states | -device alarm system issue -use of device issue -therapeutic or diagnostic output failure
Ventilation modes | SIPPV: To prevent Auto-Triggering, another breath cannot be triggered for 180 milliseconds following a breath. | -activation delayed -commission time long -installation time long | -no flow -device operates differently than expected -Deployment Issue
Ventilation modes | SIMV: Ventilation with specified pattern and frequency synchronous to the patient’s independent breathing. The patient can spontaneously breathe in between breaths but receives no pressure support. | -inappropriate ventilation support -incorrect breath type -incorrect ventilation mode | -failure to deliver -pressure issue -therapeutic or diagnostic output failure
Ventilation modes | SIMV: For weaning from ventilation. If Apnea is detected, ventilation will commence with the specified TI and TE frequency. The synchronization window for the next mechanical breath is ½ Te. | -inappropriate ventilation support -incorrect breath type -inaccurate synchronization | -use of incorrect control settings -fluid leak -medication error
Ventilation modes | SIMV + PSV: In Synchronized Intermittent Mandatory Ventilation + Pressure Support Ventilation (SIMV + PSV) mode, the patient can initiate a PSV breath in between the mandatory SIMV breaths. | -inappropriate ventilation support -incorrect ventilation mode -incorrect breath type | -human-device interface issue -unstable -failure to deploy
Ventilation modes | SIMV + PSV: The machine breath (SIMV) is synchronized with the patient’s breathing pattern. The amount of mechanical breaths per minute is same as the Preset Frequency. To prevent Auto-Triggering, another breath cannot be triggered for 180 milliseconds following a breath. | -incorrect breath type -incorrect breath type setting -confused breath type | -excess flow or over-infusion -calibration issue -incorrect software programming calculations
Ventilation modes | SIMV + PSV: If patient stops spontaneous breathing the ventilator will deliver the amount of mechanical breath set with the Frequency (Rate) button. Every spontaneous Inspiratory effort of the patient is supported with the Preset Ppsv Pressure Level. It ends when flow termination criteria are met. | -incorrect breath type -fail to initiate breath -fail to indicate button status | -device operates differently than expected -improper device output -tidal volume fluctuations
Ventilation modes | SIMV + PSV: Each SIMV breath can be triggered by the patient. Between SIMV breaths, the patient can trigger a supportive PSV breath. • Minimum pressure for the Ppsv is 2 mbar (trigger breath) / 5 mbar (untriggered breath) above PEEP. • Maximum PSV Level Ppsv = Pinsp | -incorrect breath type -fail to trigger -fail to generate indication | -use of device issue -medication error -human-device interface issue
Ventilation modes | SIMV + PSV: A triggered breath is colored Green. The non-triggered breaths are Grey. | -inappropriate color scheme -incorrect breath type -fail to initiate breath | -improper device output -inadequate user interface -defective item
Ventilation modes | SIMV + PSV: The Inspiratory Pressure for the PSV breath is set using the Ppsv button. The PSV breath is terminated at one of the following criteria: • Airway Pressure exceeds a pressure Setting PPSV • Inspiratory Flow has dropped to 1 to 85 % of the maximum value of Inspiratory Flow. | -incorrect breath type setting -inappropriate pressure setting -incorrect breath type | -use of incorrect control settings -fluid leak -pressure issue
Ventilation modes | PSV: The Pressure Support Ventilation (PSV) option is used to support the pressure of insufficient spontaneous breathing in Triggered Ventilation modes. The breathing frequency is determined by the spontaneously breathing patient, whereas the ventilator assumes an adjustable portion of the breathing effort. | -insufficient pressure -incorrect breath type setting -incorrect breath type | -deployment issue -device operates differently than expected -unstable
Ventilation modes | PSV: To prevent Auto-Triggering, another breath cannot be triggered for 180 milliseconds following a breath. | -fail to trigger -commission time long -activation delayed | -no flow -defective item -capturing issue
Ventilation modes | PSV: Thus, inspiration is controlled by the patient, and occurs according to the parameters set in Ventilation type SIMV. However, the ventilator will now control Expiration. This will occur when one of the following criteria is met: • Airway Pressure exceeds a PPSV pressure Setting • Inspiratory Flow has dropped to 1 to 85% of the Maximum Inspiratory Flow Setting (indicating a virtually filled lung) | -expiration date error -loss of control -incorrect settings | -user used incorrect product for intended use -programming issue -use of incorrect control settings
Ventilation modes | PSV: If the patient stops breathing spontaneously the ventilator will assume alternate breathing with the Ventilation parameter settings. Apnea Backup Ventilation will start after preset Apnea Time. If this is set to OFF, Backup starts after E-time. | -inappropriate ventilation support -incorrect ventilation mode -incorrect settings | -device operates differently than expected -failure to deploy -user used incorrect product for intended use
Ventilation modes | PSV: The Inspiratory Pressure for the PSV breath is set using the Pinsp button. | -inappropriate pressure setting -inaccurate pressure reading -insufficient pressure | -use of incorrect control settings -pressure issue -fluid leak
Ventilation modes | CPAP: In CPAP Ventilation the patient breathes spontaneously, the ventilator does NOT provide breathing. This mode will merely produce a positive pressure during inspiration and expiration, noticeably reducing the patient’s breathing effort. The maximum automatic leak compensation is 100% of the Inspiratory Flow. | -incorrect breath type -inappropriate ventilation support -fail to initiate breath | -use of device issue -device operates differently than expected -user used incorrect product for intended use
Ventilation modes | CPAP: In the event of an interruption in the patient’s breathing following the default Apnea Period, the ventilator performs a default number of mechanical breaths to stimulate spontaneous breathing. After breathing commences, stimulation stops and only commences with the next Apnea event. | -incorrect breath type -fail to initiate breath -inadequate breath type | -device expiration issue -improper device output -no display or display failure
Ventilation modes | CPAP: CPAP can also be applied in Non-Invasive mode to the patient (Nasal CPAP), using a device with a 15-millimeter adaptor, for example: oronasopharynx tubes or RAM Cannula; in this case you must turn OFF the ventilator and remove the flow sensor from the Y-piece of the dual limbed breathing circuit and connect the proper interface with a 15-millimeter adaptor. | -inappropriate method -inappropriate user -unable to view display | -improper or incorrect procedure or method -use of device issue -device operates differently than expected
Ventilation modes | NCPAP / DUOPAP: In NCPAP / DUOPAP mode the patient spontaneously breathes from a mask or nasal prongs. A NIV Trigger option can be additionally implemented to provide limb modes with Breath detection (Apnea monitoring with Alarms) and triggered supported breaths (or Synchronized Transition from low to high CPAP levels) in DUOPAP. | -incorrect breath type setting -fail to detect occlusion -inadequate breath types | -inaccurate delivery -medication error -unstable
Ventilation modes | NCPAP / DUOPAP: Depending on the system generator, fabian will then automatically select the proper NIV Trigger Sensor: • Flow Sensor (Neonatal) with Infant Flow / INSPIRE and Infant Flow LP generators • Pressure Sensor with Medijet (by Medin) generator Additionally, monitoring SpO2 and PCO2 is always required. | -incorrect sensor orientation -fail to recover sensor -fail to setup valve correctly | -fluid leak -Pressure Issue -device operates differently than expected
Ventilation modes | NCPAP: Supplies a positive Airway Pressure with automatic Leak Compensation. The maximum flow compensation is selectable by menu. | -fail to update leak flow -inaccurate pressure reading -inappropriate pressure level | -fluid leak -use of device issue -Pressure Issue
Ventilation modes | DUOPAP: Same as nCPAP but with the option of Positive Pressure Ventilation with adjustable frequency and Inspiratory Period. | -fail to adjust pressure -inappropriate ventilation support -fail to adjust settings | -pressure issue -device difficult to setup or prepare -use of incorrect control settings
Ventilation modes | NCPAP / DUOPAP: This mode requires a special NCPAP Patient Set with NCPAP generator. Before using the NCPAP / DUOPAP mode the correct system must be specified in the specifications menu. The following systems currently can be used: • Infant Flow • Infant Flow LP • MediJet | -incorrect settings -missing information -inappropriate user | -use of device issue -device operates differently than expected -device difficult to setup or prepare
Ventilation modes | NCPAP / DUOPAP: When ventilating a patient using a nasal interface in nCPAP or DuoPAP modes, the clinician must make a compromise between providing additional flow to compensate for a leak at the nose, or providing an alarm to warn of a leak, particularly when the appropriate size prongs are small. For a patient who is capable of spontaneous breathing without support, leak compensation may be set to a high setting, in order to provide additional flow to compensate for a leak at the interface between the prongs and the nose. This will provide a more stable pressure support in the presence of variable leak (for example, with patient movement). However; in this scenario, the ventilator may fail to provide an alarm is the event of decannulation. | -fail to compensate -improper compensation -inadequate flow | -fluid leak -Device Handling Issue -device markings issue
Ventilation modes | NCPAP / DUOPAP: For a patient who requires an alarm in the event of decannulation, leak compensation should be set to a low setting. This will provide less stable pressure support in the presence of variable leak at the nose, but a reliable alarm for decannulation. | -low pressure -insufficient pressure -fail to activate occlusion alarm | -incorrect measurement -use of device issue -Device Handling Issue
Ventilation modes | NCPAP / DUOPAP: If Inspiratory Time is set lower than 1 second; the maximum Pressure can NOT be reached. Depending on the Tubing Set, NCPAP Generator and Humidifier. | -fail to set tidal volume limit -pressure out of range -incorrect breath type setting | -component missing -use of incorrect control settings -inadequate filtration process
Ventilation modes | O2 Therapy mode: O2 Therapy is an option which allows use of a Continuous Flow of blended gas, between 0 to 15 L/min in NEO and 0 to 30 L/min in PED mode. Nasal Cannulas of various makes like F&P, Atom or similar can be used. There are no Alarm functions active in this mode, except for the set FiO2 | -fail to control flow -inappropriate therapy -inappropriate flow pattern | -over-sensing -device sensing issue -Malposition of Device
Ventilation modes | O2 Therapy mode: This mode can also be used to put the Ventilator in Standby mode. By setting a Flow of 4L/min, the Humidifier dual servo temperature controls remain active, so no need to switch it OFF in case of short-term Standby mode. | -inappropriate therapy -fail to adjust settings -fail to deliver therapy | -use of device issue -incorrect measurement -defective item
Ventilation modes | O2 Therapy mode: There are no patient alarms active in this mode. | -fail to activate alarm -alarm notified -inappropriate patient type | -break -device alarm system issue -solder joint failure
Ventilation modes | HFO: Ventilation with High-frequency Pressure Oscillations allows a gas exchange inside the lung despite the very small Tidal Volumes (often on the scale of the Dead Space Volume). While pressure amplitudes inside the breathing tube system can be quite significant, only minor fluctuations around the Mean Pressure occur inside the lung. | -fail to update tidal volume -fail to calculate tidal volume -pressure out of range | -Device Alarm System Issue -failure to deploy -decrease in pressure
Ventilation modes | HFO: The mechanical strain due to periodic lung expansion and relaxation is minimal. High-frequency Ventilation about the PEEP/CPAP level as Mean Pressure Average. The High frequency pulses superimpose the Mean Pressure at the set frequency. | -inappropriate pulse period -high amplitude -low pressure | -programming issue -electronic property issue -use of device issue
Ventilation modes | HFOV: can be applied also in Non-Invasive mode (Nasal-HFO) with Nasal prongs, Mask or Cannula; in this case you must turn OFF and remove the Flow Sensor from the Y-piece of the dual limbed breathing circuit. | -fail to remove -fail to reset sensor -inappropriate method | -device difficult to setup or prepare -positioning issue -component incompatible
Ventilation modes | HFO: During HFO ventilation, the disconnection detection is only possible to a limited extent. Be aware that an external monitoring device is required in any case of ventilation. | -device disconnected -loss of detection -device detached | -human-device interface issue -failure to sense -use of incorrect control settings
Ventilation modes | HFO lung recruitment: Lung Recruitment is an optional setting in HFO. The Pmean is thereby increased cyclically to an adjustable value “Pmean rec.” The Repetition Frequency is selectable from one cycle / hour to four cycles / minute. The Inspiratory Time can be set from two to 60 seconds if the Expiratory Time is ≥two seconds. | -incorrect value -incorrect settings -inappropriate usage | -electronic property issue -failure to read input signal -use of device issue
Ventilation modes | Special HFO mode information: When operating the device using the internal battery, HFO amplitude is limited to 60 mbar; During HFO Ventilation the distinct recognition of disconnection and measurement of the Minute Volume is limited; Use of a heated Inspiratory and Expiratory Tubing system is strongly recommended in HFO mode. | -inaccurate amplitude measurement -low readings -inappropriate threshold setting | -use of device issue -unstable -detachment of device component
Ventilation Additives | To optimize Ventilation, the following additives can be combined with the selected mode. | -incorrect ventilation mode -inappropriate ventilation support -inadequate ventilation | -protective measure issue -device cleaning issue -human-device interface issue
Ventilation Additives | The Volume limit function automatically switches to expiration when the default Tidal Volume is reached. | -fail to set tidal volume limit -fail to use default pressure setting -incorrect volume calculation | -use of device issue -use of incorrect control settings -device operates differently than expected
Ventilation Additives | When changing the ventilation pattern or adjusting the volume limit, the Tidal Volume is automatically limited. The display will show the message: “Volume Limited”. | -fail to set tidal volume limit -fail to compensate tidal volume -incorrect limit setting | -incorrect display -tidal volume fluctuations -use of device issue
Ventilation Additives | Volume limit: In the event of Flow Measurement Outage (for example: changing the sensor) “V-limit” will automatically be deactivated. After Flow Measurement is available again “V-limit” is automatically reactivated. | -fail to deactivate -incorrect limit setting -fail to reset alarm automatically | -use of device issue -incorrect measurement -incorrect display
Ventilation Additives | Volume Guarantee is a pressure controlling function that adjusts the inspiratory pressure to achieve the operator set targeted tidal volume. | -fail to set tidal volume limit -fail to adjust pressure -fail to clear tidal volume | -pressure issue -use of incorrect control settings -device difficult to setup or prepare
Ventilation Additives | Volume Guarantee: In the event flow measurement becomes unavailable (i.e. bad flow sensor or changing the flow sensor) Volume Guarantee function will automatically be deactivated. After flow measurement is available again Volume Guarantee will automatically be reactivated. | -incorrect measurement -inaccurate measurement -fail to deactivate | -Pressure Issue -device operates differently than expected -fluid leak
Ventilation Additives | Volume Guarantee: In the event of possible tube occlusion condition (i.e. “Check ETT tube” alarm) Volume Guarantee function will automatically be deactivated. After the tube occlusion condition is removed Volume Guarantee will automatically be reactivated. | -fail to activate occlusion alarm -fail to deactivate -fail to activate alarm | -use of device issue -Leak -application network issue
Special functions | Manual inspiration (manual Breath): In virtually all Ventilation modes, a Manual breath (with the Ventilation parameters set in the current Ventilation mode) can be triggered by pressing the Manual Inspiration button. • The duration of the manual breath can be set under the Configurations menu ->Ventilation. • This could last from 2 to 30 seconds, then a termination of the manual breath will be forced. • The next manual breath is only allowed after a Block Period of 200 milliseconds. | -incorrect breath type -fail to enter mode -fail to initiate breath | -fluid leak -no pressure -inadequate instructions for healthcare professional
Special functions | A Manual breath however can lead to lung over distension. In HFO mode, Manual Inspiration can be deactivated. If activated, some clinicians use this for a manual sustained lung inflation after suctioning. | -fail to deactivate -incorrect breath type -confused breath type | -device cleaning issue -instruction for use issue -inadequate instructions for healthcare professional
Special functions | Nebulizing medications: WARNING: Flow Sensor wires are HOT. • If the Flow Sensor remains in the breathing system for extended periods during Nebulizing without being cleaned, deposits can form the medication aerosols and impair Flow Measurement. • In worst case scenarios these deposits could ignite. • Disconnecting the plug from the flow sensor does NOT prevent the risk of ignition from medication deposits. • Always remove the flow sensor before nebulizing medication. | -device remains activated -inadequate warning -fail to recover sensor | -incorrect measurement -Pressure Issue -particulates
Special functions | O2 Flush / preoxygenation: Press the O2 Flush button to trigger an O2 Flush. Short-term Oxygen Spray (O2 Flush) with an increased O2 concentration is permissible in all Ventilation modes. This will automatically end after a maximum of 2 minutes (also see defaults in menu: Configuration -> Ventilation -> Flush Duration); Press the O2 Flush button in a Ventilation mode to initiate the O2 Flush; The Flush Concentration can now be set with the Rotary Pulse Encoder; Pressing the O2 Flush button again will prematurely end the Flush. | -low oxygen concentration -fail to indicate button status -inadequate breathing mode | -no display or display failure -Sticking -insufficient flow or under infusion
Accessories and options | WARNING: The items shown in this list have been approved for use with the fabian HFO ventilator. Use of non-approved items is NOT recommended and will NOT be recognized or supported by the manufacturer. If a system malfunctions with non-approved parts and /or accessories, the user is responsible and liable for any and all issues associated with the system malfunction. | -fail to approve -inappropriate interface design -lack of specification | -device handling issue -dose calculation error due to software problem -device cleaning issue
Accessories and options | Accessories List: Flow sensor connector cable; Flow sensor, single use; Flow sensor, reusable; Flow sensor, MPU, Pediatric; Adapter NIST – DISS O2; Adapter NIST– DISS Air; Masimo SpO2 module USB HPL/MIN cable; Masimo SpO2 module LEMO HPL/MIN cable; Microstream MicroPod etCO2 Sensor; LoFlo Sidestream etCO2 Sensor; Breathing circuit set – Altera; NIV generator – Inspire InfantFlow; NIV generator – Vyaire InfantFlow LP; NIV generator – Medin Medijet; Humidifier – WILAmed, AIRcon; Humidifier – VADI; Trolley | -accessory missing -incompatible accessory -fail to support correct use | -device cleaning issue -imprecision -difficult to open or close
Accessories and options | Spare parts list: Membrane holder, 22mm OD / 15mm ID; Membrane, expiration; Oxygen sensor | -missing part number -incorrect model number -lack of tester | -failure to recalibrate -incorrect measurement -deployment issue
CO2 monitoring | CO2 Sensor module types and selection: The fabian HFO supports three different CO2 module types to provide CO2 monitoring: 1. Microstream, MicroPod™ External etCO2 Module from Oridion® (Covidien/Medtronic) 2. LoFlo Sidestream etCO2 Sensor from Respironics (Philips). | -incorrect sensor orientation -inaccurate sensor measurement data -inadequate device information | -device difficult to setup or prepare -positioning issue -use of incorrect control settings
CO2 monitoring | CO2 monitoring: Only one module at a time can be connected and used with the ventilator. | -loss of connection -incompatible connection -unable to connect | -failure to deploy -incorrect software programming calculations -medical gas supply problem
CO2 monitoring | CO2 monitoring: The module can be activated in the etCO2 submenu of Calibration view of fabian HFO | -fail to activate system -fail to activate alarm -fail to activate control | -no display or display failure -optical issue -incorrect measurement
CO2 monitoring | To access the CO2 Monitoring view: 1. Press the Graphics button (Figure 12-3) 2. Press the etCO2 button on the touch screen in the Graphics view | -unresponsive touch screen -fail to access -unable to access | -human-device interface issue -no display or display failure -inadequate user interface
CO2 monitoring | CO2 monitoring: After a module is selected, the following measurement related alarm limits can be set under the Alarm Limit view (Figure 12-6), which can be reached by pressing the Alarm Limits button. • etCO2 high • etCO2 low • FiCO2 high • FiCO2 low | -incorrect limit setting -fail to set limit correctly -inappropriate measurement unit | -device alarm system issue -not audible alarm -use of device issue
CO2 monitoring | Connect the CO2 module to fabian HFO: To connect the compatible CO2 module to the fabian HFO with mainboard 3.x or higher: 1. Attach the 7-pin Binder Connector from the MicroPod (Figure 12-8) to the Respective Binder Connector at the back of the fabian HFO labelled CO2. 2. Make sure the extrusion on the 7-pin Binder Connector slots in at the bottom of the Respective Binder Connector. 3. Fasten the connection by turning the Binder connector ring clockwise. | -incompatible connection -connector mismatched -loss of connection | -loose or intermittent connection -detachment of device component -blockage within device or device component
CO2 monitoring | Connect to fabian HFO with Main Board 2x: To connect any of the compatible CO2 module to the fabian HFO with Main Board 2.x, 1. Attach the 9-pin RS-232 Connector from the MicroPod to the Receptacle RS-232 Connector at the back of the fabian HFO labelled CO2. 2. Fasten the connection by turning the screws on either side of the pinned connector clockwise. | -connector mismatched -connector broken or damaged -loss of connection | -detachment of device component -loose or intermittent connection -Detachment of Device Component
CO2 monitoring | MicroPod® sensor module: The module is designed to be incorporated into a host monitoring system. The module allows for measurement of inspired /expired carbon dioxide and respiration rate on patients in the Operating Room, ICU, NICU, Sedation procedures, GI, General Floor, Transport and Emergency treatment. | -incorrect measurement -incorrect incorporated -fail to incorporate | -protective measure issue -incorrect software programming calculations -failure to sterilize
CO2 monitoring | MicroPod® sensor module: The method for measuring etCO2 is based on non-dispersive IR absorption of the CO2 in the breath sample using the Oridion IR source. | -sensor measurement interfered -fail to measure -incorrect measurement | -use of incorrect control settings -reset issue -use of device issue
CO2 monitoring | The major components of the MicroPod® sensor module are as follows: • CO2 sensor assembly, including the Oridion long-life source • Flow System, including pump • FRS to identify the correct usage of a Microstream® FilterLine • Input gas connector with the FRS interface • Power supply • Serial interface • μController + peripherals | -incorrect usage -fail to support correct use -inappropriate usage | -failure to service -protective measure issue -incorrect measurement
CO2 monitoring | MicroPod® sensor module: The capnography module is intended to provide professionally trained health care providers with continuous, non-invasive measurement and monitoring of carbon dioxide concentration of the expired and inspired breath and respiration rate. The capnography module is intended for use with neonatal, pediatric, and adult patients in hospitals, hospital -type facilities, transport, intra-hospital transport and home environments. | -lack of health and status monitor -inadequate training -insufficient training | -incorrect measurement -incorrect software programming calculations -data back-up problem
CO2 monitoring | MicroPod® sensor module: The MicroPod contains the Oridion microMediCO2 module and uses Microstream capnography. The MicroPod external CO2 module attaches to the host monitor with an external communication cable, and data from the module appears on the host monitor screen. | -sensor damaged -lack of tool -loss of electrical connection | -No Display or Display Failure -incorrect software programming calculations -compliance issue
CO2 monitoring | Microstream capnography modules use Microstream non-dispersive infrared (NDIR) spectroscopy to continuously measure the amount of CO2 during every breath, the amount of CO2 present at the end of exhalation (etCO2) and the respiratory rate. | -fail to measure battery voltage -fail to measure -inadequate markings | -incorrect measurement -calibration issue -invalid sensing
CO2 monitoring | The respiration rate is referred on the fabian HFO as Freq EtCO2. Infrared spectroscopy is used to measure the concentration of molecules that absorb infrared light. Because the absorption is proportional to the concentration of the absorbing molecule, the concentration can be determined by comparing its absorption to that of a known standard. | -incorrect measurement -inaccurate measurement -fail to measure | -no display or display failure -incorrect display -not audible alarm
CO2 monitoring | The Microstream etCO2 sampling lines deliver a sample of the inhaled and exhaled gases from the ventilator circuit or directly from the patient (through an oral /nasal cannula) into the monitor for CO2 measurement. Moisture and patient secretions are extracted from the sample, while maintaining the shape of the CO2 waveform. | -fail to obtain sample -incorrect sample size -fail to expel air | -output above specifications -pressure issue -unstable
CO2 monitoring | The 50 mL / min. sampling flow rate reduces liquid and secretion accumulation, decreasing the risk of obstruction in the sample pathway in humid ICU environments. This low sampling rate is ideally suited to measure a neonate’s low tidal volume. | -low rate -inaccurate flow rate -fail to reduce contamination | -use of device issue -fluid leak -use of incorrect control settings
CO2 monitoring | Once inside the Microstream CO2 sensor, the gas sample goes through a micro-sample cell (15 microliters). This extremely small volume is quickly flushed, allowing for fast rise time and accurate CO2 readings, even at high respiration rates | -excessive charge time -fail to purge -short fill | -use of device issue -dose calculation error due to software problem -reset issue
CO2 monitoring | The Micro Beam IR source illuminates the micro-sample cell and the reference cell. This proprietary IR light source generates only the specific wavelengths characteristic of the CO2 absorption spectrum. Therefore, no compensations are required when different concentrations of N2O, O2, anesthetic agents and water vapor are present in the inhaled and exhaled breath. The IR that passes through the micro-sample cell and the IR that passes through the reference cell are measured by the IR detectors | -fail to generate indication -fail to provide sufficient pressure -low volume | -use of device issue -device difficult to setup or prepare -detachment of device component
CO2 monitoring | The microprocessor in the board calculates the CO2 concentration by comparing the signals from both detectors. | -fail to generate periodic calculation -dose calculation incorrect -incorrect concentration | -failure to align -device difficult to setup or prepare -output below specifications
CO2 monitoring | MicroPod messages: The MicroPod supports the following messages and alarms: "check calibration", "check flow", "CO2 value over-range", "initialization", "invalid CO2 value", "pump off", "purging in progress", “Sensor off", "Sensor state: ok", "SFM in progress", "standby", "Temperature out of range". The messages are displayed on the information line of the CO2 Monitoring screen or in the CO2 Settings menu. | -inadequate message display -incorrect message display -invalid signal | -use of incorrect control settings -instruction for use issue -Device Handling Issue
CO2 monitoring | MicroPod calibration: Oridion supplies a Calibration kit (including a Calibration software CD and Calibration accessories) that enables calibration of the module on a PC. The fabian HFO displays a Calibration message that indicates on the host monitor CO2 screen that calibration is required at the designated interval | -fail to calibrate manually -calibration failed -out of calibration | -failure of device to self-test -human factors issue -Incorrect or Inadequate Result
CO2 monitoring | MicroPod calibration interval: The monitor should be calibrated by qualified service personnel after the first 1,200 operating hours of use or 12 months, whichever comes first. After 1,200 operating hours of use or 12 months, calibration should be performed every 12 months or after 4,000 operating hours, whichever comes first. | -calibration failed -incorrect interval calculation -fail to calibrate manually | -use of incorrect control settings -unauthorized access to computer system -pressure issue
CO2 monitoring | Basic principles for choosing Microstream® CO2 Sampling lines: When choosing Microstream CO2 sampling lines, the following should be considered: The condition of the patient (ventilated or not ventilated); If the patient is ventilated, whether ventilation is humidified or non-humidified; Patient’s size and weight; The probability that the patient will switch between oral and nasal breathing; Duration of use; For best results, for short term monitoring, use Microstream CO2 sampling lines with Orange connectors. For long term monitoring, use Microstream CO2 sampling lines with Yellow connectors. Products that include an “H” in the name are intended for long term use. | -insufficient sample size -inappropriate method -inadequate labeling | -use of device issue -device operates differently than expected -market issue
CO2 monitoring | In the etCO2 module menu you can: Disable or switch etCO2 module ON. Choose the displayed unit for measurement values (mmHg, kPa, Vol%). Specify the time until the pump automatically restarts (5, 10, 15 or 30 minutes or OFF). after it has been manually turned OFF. | -inappropriate measurement unit -fail to disable control -incorrect measurement | -incorrect display -incorrect measurement -electromagnetic interference (emi)
CO2 monitoring | In the etCO2 module, the following information on the module is displayed: Calibration state, Serial number Service state, Software release, The Sensor state | -inadequate message display -inadequate information -fail to request calibration | -no display or display failure -use of incorrect control settings -incorrect display
CO2 monitoring | The fabian HFO CO2 Monitoring View (Figure 12-18) displays real time CO2 data. The displayed data includes: Real time “etCO2“values along with selected unit “mmHg”, “kPa” or “Vol%”. Respiration rate (RR) named “Freq” in breaths-per-minute (bpm). EtCO2 Waveform. | -incorrect display format -loss of data -incorrect data format | -incorrect display -poor quality image -no display or display failure
CO2 monitoring | 1. To start the Pump OFF Mode, slide the Blue button under the Pump [ON / OFF] text. The Pump OFF timer will begin. If the auto‒restart of the CO 2 pump is enabled, the state is displayed in the information line of the CO2 Monitoring screen. 2. To restart the timer, press the Restart Timer button. On the top of the main screen a message (“! EtCO 2 pump will be restarted in 1min!”) appears one minute before the pump turns ON automatically. | -fail to indicate button status -fail to communicate with pump -incorrect message display | -no display or display failure -break -incorrect display
CO2 monitoring | Connecting a filter line: Before monitoring a patient with capnography, the appropriate filter line must be connected to the monitor and to the patient. | -loss of connection -fail to connect -connection leakage | -fluid leak -component missing -loose or intermittent connection
CO2 monitoring | To make the connections for filter line: 1. Slide open the filter line input connector shutter and connect the appropriate filter line. 2. Screw the filter line connector into the monitor clockwise until it can no longer be turned. 3. Connect the filter line to the patient as described in the Directions for Use supplied with the filter line. | -connector mismatched -connection broken or damaged -fail to connect | -loose or intermittent connection -component missing -difficult to open or close
CO2 monitoring | When the Filter Line is connected, the monitor will immediately begin to search for breaths, but it will NOT indicate a No Breath condition before any valid breaths have occurred. | -fail to initiate breath -loss of indication -invisible indication | -fluid leak -improper device output -decrease in pressure
CO2 monitoring | MicroPod® Functional testing: We recommend performing a functional test each time the MicroPod is used. After connection of the MicroPod, verify that the CO2 data is displayed on the monitor | -inadequate message display -loss of function -incorrect display format | -use of device issue -device maintenance issue -missing test results
CO2 monitoring | MicroPod Start-up: The time before CO2 measurements are available includes power-up time and initialization time. The initialization time includes module initialization and self-tests. Power-up time: Maximum 10 seconds Initialization time: Typically, 30 seconds, maximum 180 seconds | -insufficient time -inadequate time format -inappropriate time format | -reset issue -failure to deploy -programming issue
CO2 monitoring | System response time: The system response time of the microMediCO2 and fabian HFO together is typically 3.1 seconds. This includes the following timing sequences: 2.7 seconds delay time, 0.2 seconds rise time from the microMediCO2 and 0.2 seconds response time from the fabian HFO. | -long response time -insufficient time -insufficient discharge time | -deployment issue -data back-up problem -device sensing issue
CO2 monitoring | MicroPod Self-maintenance (SFM) interval: Self-Maintenance (SFM) is performed only during measurement mode. The module performs one or more of the following: Ambient pressure measurement Auto zero (AZ), Flow test | -fail to hold pressure -inappropriate interval settings -fail to generate periodic notification | -incorrect measurement -deployment issue -no pressure
CO2 monitoring | SFM is triggered: During the first hour after entering measurement mode, periodically for durations of 10 seconds at a rate which limits the total time consumed by SFMs to less than 2% of the time in which active measurements are taken. Following the first hour after entering measurement mode, periodically for durations of 10 seconds at a rate of once per hour. | -inappropriate measurement unit -fail to generate periodic notification -timing variation | -use of incorrect control settings -failure to align -device expiration issue
CO2 monitoring | SFM is triggered: A significant temperature change measured by the module (not less than 8°C from the temperature at the time that the last AZ was detected). | -temperature out of range -high temperature -over pressure | -incorrect measurement -device expiration issue -Failure to Sense
CO2 monitoring | SFM is triggered: A significant ambient pressure change (not less than 18 mmHg relative to the last ambient pressure measurement) for a period of 30 seconds | -pressure out of range -inappropriate pressure level -over pressure | -incorrect measurement -pressure issue -Low Battery
CO2 monitoring | The module prevents the triggering of an SFM in the following situations: In case of purging until the end of this state; During a breath absence period which follows a valid breath; While waiting a minimum of 20 seconds for host SFM enable command. (After the 20-second opportunity given to the host to schedule the SFM passes, the module schedules the SFM according to a priority determined by current conditions). | -circuit triggered without software requesting activation -request trigger failed -undesired contact | -unintended movement -device damaged prior to use -split
CO2 monitoring | Periodic service. Periodic maintenance is recommended according to operating hours: The CO2 pump should be replaced every 30,000 operating hours; A calibration should be performed as described in section “12.3.3.3: MicroPod calibration interval". | -fail to generate periodic notification -fail to clear periodic notification -lack of maintenance process | -use of device issue -device maintenance issue -delivery system failure
CO2 monitoring | Limited Operating life. The following parts of a MicroPod with a microMediCO2 capnography board will require replacement after a set period: The CO2 Pump should be replaced every 30,000 operating hours; The I.R. Source should be replaced every 30,000 operating hours. | -reduced life cycle -inadequate operating mode -low life cycle | -use of incorrect control settings -Device Handling Issue -device maintenance issue
CO2 monitoring | Respironics® etCO2 sensors: When using either of the following sensors from Respironics (Philips) you need to select the Respironics option in the etCO2 menu (Figure 12-21). | -incorrect sensor orientation -inappropriate program selected -fail to recover sensor | -use of device issue -use of incorrect control settings -unstable
CO2 monitoring | Respironics messages: "Check Airway Adapter", "Check Sampling Line", "CO2 out of Range", "No Message", "Sensor Faulty", "Sensor off", "Sensor Over Temperature", "Sensor state: ok", "Sensor Warm Up", "Standby", "Zero in Progress", "Zero Required” | -incorrect message display -inadequate display message -fail to vent air out | -capturing issue -device cleaning issue -fluid leak
CO2 monitoring | Respironics module settings and information. In addition, the following information on the module is displayed: The sensor state, Last service date, Sensor type | -incorrect settings -inadequate message display -inadequate information | -use of incorrect control settings -failure to deploy -incorrect display
CO2 monitoring | The fabian HFO CO2 Monitoring View (Figure 12-22) displays real time CO2 data. The displayed data includes: Real time “etCO2“values along with selected unit “mmHg”, “kPa” or “Vol%”. Respiration rate (RR) named “Freq” in breaths-per-minute (bpm). etCO2 Waveform. | -fail to display data -incorrect display format -loss of data | -incorrect display -poor quality image -no display or display failure
CO2 monitoring | Respironics CO2 sensor zeroing. The Respironics CO2 sensor is compatible with a variety of different airway adapters: Zeroing allows for the sensor to accommodate the optical characteristics of each of the different adapter types; A zero should be performed whenever the type of adapter being used with the Respironics CO2 sensor is changed; For optimal accuracy, a zero should also be performed whenever the sensor is connected to the host system. | -sensor zero offset out of range -fail to compatible with oxygen supply -incorrect placement | -incorrect measurement -failure to recalibrate -image resolution poor
CO2 monitoring | Before performing a zero calibration (Figure 12-23), the Respironics CO2 sensor should be removed from the patient circuit, and the airway adapter type to be used in the circuit should be inserted into the sensor: Care should be taken to ensure that the airway adapter is clear of any residual CO2 gas; The maximum elapsed time for a zero is 30 seconds. The typical time for a zero is 15 to 20 seconds. | -unable onehand to remove -fail to reset sensor -fail to remove | -incorrect software programming calculations -unable to obtain readings -failure to deploy
CO2 monitoring | Several conditions could also request that a zero be performed: These requests stem from changes in the airway adapter that could indicate that the sensor is not in optimal measuring condition. When this occurs, the airway adapter should be checked to ensure optical occlusions such as mucus have NOT obscured the adapter window; If occlusions are found, the airway adapter must be cleaned or replaced. | -fail to measure -fail to generate indication -fail to identify change | -failure to sense -deployment issue -electrical shorting
SpO2 module | Setting up the Masimo sensor: To enable the SpO2 measurement, the module must be enabled. This could be done in the menu of the SpO2 module. | -inappropriate measurement unit -incorrect measurement -fail to enter mode | -failure to deploy -device difficult to setup or prepare -material rigid or stiff
SpO2 module | From the SpO2 menu, you can choose different module settings described in the following sections. Here, information on the sensor state and software release of the used module is displayed. You can select a sensitivity mode, a Fast SAT mode, Alarm Delay and the SpO2 Averaging Time. | -inappropriate software release -incorrect software settings -inappropriate software interface | -use of device issue -deployment issue -use of incorrect control settings
SpO2 module | The information on the sensor state is displayed in the configuration menu, in the information bar at the bottom. In addition, the following general messages for the sensor state can appear in the information bar in the Configurations menu: "checking ", "processing active ", "pulse search " | -incorrect message display -fail to display message -inadequate message display | -no display or display failure -break -incorrect display
SpO2 module | Sensitivity mode: The Sensitivity mode can be Max, Normal or Adaptive Probe OFF Detection (APOD). | -inappropriate sensitivity setting -low sensitivity -sensitivity setting out of range | -device difficult to setup or prepare -use of incorrect control settings -positioning issue
SpO2 module | Maximum: This mode should be used for the most critical patients, where obtaining a reading is most difficult. Maximum Sensitivity is designed to interpret and display data for even the weakest of signals. This mode is recommended during procedures and when clinician and patient contact is continuous. | -inadequate readings -incorrect readings -low sensitivity | -use of device issue -no display or display failure -incorrect display
SpO2 module | Normal: This mode provides the best combination of sensitivity and probe-off detection performance. This mode is recommended for most patients. | -inappropriate sensitivity setting -low sensitivity -incorrect detection | -use of device issue -deployment issue -use of incorrect control settings
SpO2 module | Adaptive Probe OFF Detection (APOD): This mode is the least sensitive in picking up a reading on patients with low perfusion but has the best detection for probe-off conditions. This mode is useful for patients that are at risk of the sensor becoming detached (pediatric, combative, etc.). | -low sensitivity -incorrect readings -inappropriate detection criteria | -use of incorrect control settings -device operates differently than expected -failure to deploy
SpO2 module | Fast SAT mode. The Fast SAT mode can also be selected. The Fast Sat mode can either be ON or OFF. | -atrial tracking too fast -lack of fast acting -fail to support high speed | -use of device issue -use of incorrect control settings -deployment issue
SpO2 module | Fast SAT enables rapid tracking of arterial oxygen saturation changes by minimizing the averaging. This mode is clinically applicable during procedures when detecting rapid changes in oxygen saturation is paramount such as induction, intubation, and sleep studies. | -fail to adjust settings -incorrect measurement -fail to support high speed | -calibration issue -Pressure Issue -computer software issue
SpO2 module | Alarm delay: You can configure the alarm delay so that transient desaturations do not give an immediate audible alarm or visual display. You can also select alarm delay settings of 0, 5, 10, and 15 seconds. | -fail to configure vibrating alarm -configuration confused -incorrect configuration | -Device Alarm System Issue -alarm not visible -device alarm system issue
SpO2 module | SpO2 averaging time: Selectable averaging time enables clinicians to optimize SpO2 monitoring for various patient application areas such as the PACU, NICU, ICU, Telemetry, and Sleep. This reduces the number of alarms due to real, yet non-clinically actionable rapid desaturations. There are different averaging times that can be selected (2 to 4, 4 to 6, 8, 10, 12, 14, 16). | -inadequate time format -inadequate evaluation -inadequate hazard information | -use of device issue -programming issue -unauthorized access to computer system
PRICO module | General information on PRICO: Pulse oximetry is a continuous and non-invasive method of measuring the level of arterial oxygen saturation in blood. The measurement is taken by placing a sensor on a patient, usually on the fingertip for pediatric, adult patients, and the hand or foot for neonates. The sensor is connected to the fabian with a patient cable. The sensor collects signal data from the patient and sends it to the ventilator. | -incorrect measurement -inaccurate measurement -inappropriate measurement unit | -protective measure issue -decreased sensitivity -calibration issue
PRICO module | The fabian displays the calculated data additional to the ventilation data in four ways: 1. As a Percent value (%) for arterial oxygen saturation (SpO2), 2. As a Pulse Rate (bpm), 3. The Perfusion Index (PI) and 4. As a plethysmographic waveform | -incorrect data -incorrect data format -received data incorrectly | -incorrect display -data issue -patient data issue
PRICO module | USB HPL/MIN: The USB HPL/MIN is a patient cable with an integral MSLP-1 Board contained in an enclosure that connects to Masimo Pulse Oximetry sensors and provides functional oxygen saturation (SpO2), pulse rate (PR), perfusion index (PI), Plethysmograph Waveforms, and other information through a serial digital interface. | -inadequate connection -incompatible connection -fail to connect | -user used incorrect product for intended use -therapeutic or diagnostic output failure -electrical shorting
PRICO module | The USB HPL/MIN contains a 9-pin (DB9) LNCS-compatible sensor connector with latch at the sensor end, and a standard USB Type A connector at the opposite end so that it is compatible with the USB device port from the fabian ventilator at the back side. The total length of the USB HPL/MIN, patient cable, and sensor cannot exceed 5 meters. | -fail to support ethernet small form-factor pluggable (sfp) -connector mismatched -adapter mismatched | -fitting problem -component incompatible -detachment of device component
PRICO module | The LEMO HPL/MIN is a patient cable with an integral Masimo MSLP-1 Board contained in an enclosure that connects to Masimo Pulse Oximetry sensors and provides functional oxygen saturation (SpO2), pulse rate (PR), perfusion index (PI), Plethysmograph Waveforms, and other information through a serial digital interface. | -inadequate connection -fail to connect -loss of connection | -user used incorrect product for intended use -use of device issue -Failure of device to Self-Test
PRICO module | The LEMO HPL/MIN contains a 9-pin (DB9) LNCS-compatible sensor connector with latch at the sensor end, and a 7-pin LEMO connector at the opposite end so that it is compatible with the 7-pin LEMO device port on the rear panel of the fabian ventilator. The total length of the LEMO HPL/MIN, patient cable, and sensor cannot exceed 5 meters. | -fail to latch -connector mismatched -jammed latch | -detachment of device component -failure to deploy -incorrect software programming calculations
PRICO module | PRICO can be activated in the SpO2 section of the Graphics menu. PRICO will adjust the FiO2, to keep the SpO2 within the selected range. | -fail to adjust settings -adjustment failed -fail to activate | -device difficult to setup or prepare -failure to deploy -use of incorrect control settings
PRICO module | PRICO has the following parameters: Min FiO2, Max FiO2, SpO2 low target, SpO2 high target, PRICO ON / OFF | -incorrect parameter -inappropriate method -inappropriate environment | -break -material integrity issue -device maintenance issue
PRICO module | Dependencies on Alarm Limits: • “SpO2 low target” must be greater than or equal to the low SpO2 alarm limit. • “SpO2 high target” must be less than or equal to the high SpO2 alarm limit. • PRICO can only be “ON” when the measured SIQ (Signal Quality) is higher than the SIQ alarm limit. | -low alarm volume -inappropriate safe limit -fail to clear alarm | -improper alarm -break -device alarm system issue
PRICO module | warning: Due to the dependencies of the PRICO alarm limits, it is critical that you first select the appropriate alarm limits for the patient, before switching on PRICO. | -inappropriate safe limit -alarm delayed -fail to set limit correctly | -failure to sense -no display or display failure -device alarm system issue
PRICO module | Dependencies on each other: • “Min FiO2” must be less than “Max FiO2”. • “SpO2 low target” must be less than the “SpO2 high target” | -low resistancy -inappropriate usage -inappropriate sensing | -infusion or flow issue -break -detachment of device or device component
PRICO module | Press the PRICO parameter on the touch screen to select: 1.1. Adjust the parameter with the Rotary Pulse encoder. The value will be shown in the button. 1.2. The Green bars depicting the range of the FiO2 and the SpO2 will be adjusted accordingly. 1.3. The current SpO2 and FiO2 values are indicated with vertical Blue lines. | -incorrect parameter -parameter confused -inappropriate variable value | -no display or display failure -device difficult to setup or prepare -failure to deploy
PRICO module | Press the Rotary Pulse encoder again to confirm the value as a PRICO parameter: 2.1. SIQ (Signal quality) is depicted as the red to green vertical scale. 2.2. The Red line on the SIQ scale indicates the SIQ alarm limit | -fail to initiate pulse -fail to confirm -incorrect display value | -defective item -incorrect display -no display or display failure
PRICO module | warning: The “FiO2” parameter in the main screen will be used as the Back -up “O2” in case the PRICO is turned OFF. The value can be adjusted also when PRICO is ON. | -parameter confused -incorrect parameter -fail to adjust settings | -no display or display failure -incorrect display -inadequate user interface
PRICO module | When all PRICO parameters are set appropriately for the patient, PRICO can be switched ON. The PRICO algorithm works as outlined in the diagram above. After every 30 seconds, an FiO2 adjustment is made based on the current SpO2 and its position in one of the four regions. • Outside the SpO2 target range: the FiO2 step size (1 to 10%) is determined by current SpO2, trend of SpO2 data and an extrapolation of SpO2 data. • Inside the SpO2 target range: lower half +1, if FiO2 is in lower half. Upper half decrease FiO2 with 1%. • FiO2 adjustments are made up to the pre-set FiO2 ranges. | -switch invoked inappropriately -incorrect parameter -fail to set limit correctly | -use of incorrect control settings -electromagnetic interference (emi) -device difficult to setup or prepare
PRICO module | PRICO is available on the fabian ventilators in the following modes: • Continuous Positive Airway Pressure (CPAP) • High and Low Flow Oxygen Therapy HFNC (O2 Therapy) • High-Frequency Oscillation ventilation (HFO) • Intermittent Positive Pressure Ventilation (IPPV) • NIV Bi-levels positive airway pressure (DUOPAP) • NIV Continuous Positive Airway Pressure (NCPAP) • Pressure Support Ventilation (PSV) • Synchronized Intermittent Mandatory Ventilation (SIMV) • Synchronized Intermittent Mandatory Ventilation with PSV support (SIMV+PSV) • Synchronized Intermittent Positive Pressure Ventilation (SIPPV) | -incorrect mode -inappropriate interface design -inadequate evaluation | -Pressure Issue -Occlusion within Device -failure to recalibrate
PRICO module | The “O2 Flush” can be activated during PRICO, this will deactivate PRICO. • To re-enable PRICO after an O2 Flush, push the PRICO “ON” button in the PRICO menu. • To disable the O2 Flush function, set the Flush Time to 0s | -fail to deactivate -fail to activate -sticky button | -leak -unstable -use of device issue
PRICO module | The “Start / Stop” button could be used to interrupt mechanical ventilation for 2 minutes. In nCPAP, DUOPAP and O2 Therapy mode ventilation can be interrupted indefinitely. In both cases PRICO will be switched OFF. | -ventilation interrupted -button stuck on -fail to disable ventilation | -medication error -break -material rigid or stiff
PRICO module | When PRICO is ON, this is indicated by the PRICO icon in the Status bar and the Green field with the current FiO2 labelled PRICO above the O2 parameter. | -inappropriate indication -incorrect labeling -invisible indication | -no display or display failure -device difficult to maintain -protective measure issue
PRICO module | PRICO disabling alarms, PRICO is turned OFF and cannot be re-enabled while any of the following alarms are active: • Air supply pressure • Check ET tube • Clean flow sensor • Check SpO2 sensor • Flow sensor defect • Flow sensor not connected • Input pressure blender • Low SIQ • O2 sensor calibration error • O2 value out of range • Oxygen sensor defect • Oxygen sensor used up • Oxygen supply pressure • Patient disconnected • SpO2 module disconnected • SpO2 sensor failure • Tube Occlusion | -fail to activate alarm -fail to activate air-in-line alarm -fail to disable alarm | -Pressure Issue -device maintenance issue -failure to zero
PRICO module | PRICO re-enabling cases. PRICO automatically turns back on in case the following alarms are not active anymore: • Check ET tube • Check SpO2 sensor • Clean flow sensor • Flow sensor defect • Flow sensor not connected • Low SIQ • O2 value out of range • Patient disconnected • Tube Occlusion | -fail to turn on automatically -fail to turn on device -fail to enter auto mode | -Pressure Issue -imprecision -insufficient flow or underinfusion
PRICO module | PRICO errors. Inaccurate measurements can be caused by: • Excessive patient movement • Exposure to excessive illumination, such as surgical lamps (especially those with a xenon light source), bilirubin lamps, fluorescent lights, infrared heating lamps, or direct sunlight (correct exposure to excessive illumination by covering the sensor with a dark or opaque material) • Incorrect sensor application or use • Intravascular dyes such as indocyanine green or methylene blue • Placement of a sensor on an extremity with a blood pressure cuff, arterial catheter, or intravascular line. • Significant levels of dysfunctional hemoglobins (for example: carboxyhemoglobin or methemoglobin) • Venous pulsation | -inaccurate measurement -incorrect measurement -error message | -inaccurate delivery -use of device issue -pressure issue
PRICO module | Loss of pulse signal can occur in any of the following situations: • The patient has hypotension, severe vasoconstriction, severe anemia, or hypothermia • The patient is in cardiac arrest or shock • The sensor is too tight • There is arterial occlusion proximal to the sensor • There is excessive illumination from light sources such as a surgical lamp, a bilirubin lamp, or direct sunlight | -loss of signal -inappropriate pulse delivered -pulse enhanced | -failure to read input signal -device sensing issue -retraction problem
PRICO module | An alarm of alarm limits is displayed with a Red alarm clock at the corresponding alarm limit and as a flashing text with Yellow background in the Information bar. | -incorrect message display -inadequate message display -loss of alarm type indication | -device alarm system issue -defective item -no display or display failure
PRICO module | The actual alarm limits and state (OFF) are also displayed in the graphical view next to the numeric values with the High / Low alarm limit symbol. | -incorrect limit setting -inadequate message display -loss of state indication | -incorrect display -no display or display failure -Device Alarm System Issue
FOT module | FOT is available as a submenu under Waves / loops in fabian HFO ventilators in the following modes: • CPAP : Continuous Positive Airway Pressure • HFO : High-Frequency Oscillation ventilation • IPPV : Intermittent Positive Pressure Ventilation • PSV : Pressure Support Ventilation • SIMV : Synchronized Intermittent Mandatory Ventilation • SIMV + PSV : Synchronized Intermittent Mandatory Ventilation with PSV support • SIPPV : Synchronized Intermittent Positive Pressure Ventilation | -inadequate deliverable -incorrect configuration -inappropriate classification | -instruction for use issue -device remains activated -incorrect measurement
FOT module | FOT is only available if: • There is an FOT license installed on the device • Flow sensor is connected and turned ON • Any of the previously mentioned ventilation modes are active | -fail to control flow -loss of connection -fail to install | -user used incorrect product for intended use -Failure to Calibrate -instruction for use issue
FOT module | There are two types of FOT distinguished in fabian HFO: • FOT-HFO: The forced oscillation is created by the HFO module. The measurement does not change the ventilation mode as it is available only in HFOV. • FOT-conventional: The forced oscillation is created by the exhalation membrane. During the measurement, ventilation mode is changed to CPAP. | -incorrect measurement -inappropriate measurement unit -fail to adjust settings | -not audible alarm -use of device issue -protective measure issue
FOT module | FOT-HFO and FOT-conventional shares: • The same layout • The same operation steps | -inappropriate display size -incorrect placement -incorrect configuration | -inadequate user interface -Device displays error message -size incorrect for patient
FOT module | FOT-HFO and FOT-conventional differ in: • Parameter names • Parameter ranges • Forced signal creation | -incorrect parameter -incorrect program parameter -parameter confused | -use of device issue -Device Packaging Compromised -human factors issue
FOT module | The following sensors and ventilation additives can be used together with FOT: • Volume Guarantee (VG) • Volume Limit (VL) • SpO2 sensor • Predictive Intelligent Control of Oxygenation (PRICO) The following option is disabled during the FOT recruitment / derecruitment maneuver: • Lung recruitment at HFOV | -inadequate information -fail to convert to back-up -loss of indication | -use of device issue -protective measure issue -incorrect measurement
FOT module | FOT general layout: FOT graph, FOTsteps setting, Pmean low/ PEEPlow setting, Pmean high/ PEEPhigh setting, Step size information field, Information field, PIPmax information field, Control buttons, Information field. | -inadequate information -incorrect graph format -incorrect settings | -device markings issue -size incorrect for patient -use of incorrect control settings
FOT module | Erasing the FOT graph. The FOT Graph is cleared if any of the following actions are performed: • a new FOT sequence is started, • the device is restarted, • the patient range is changed, • device date / time is changed | -incorrect graph format -incorrect graph -fail to clear data | -not audible alarm -reset issue -incorrect display
FOT module | FOT disabling conditions. FOT is turned OFF and cannot be re-enabled while any of the following actions occurs or alarms are active: Any system alarm: • Air supply pressure • Check Blender • Exhalation calibration • Oxygen supply pressure • Tube Occlusion | -disable failed -fail to disable ventilation -fail to disable control | -Device Alarm System Issue -decrease in pressure -Pressure Issue
FOT module | Any flow sensor alarm: • Clean flow sensor • Flow sensor defect • Flow sensor not connected | -fail to activate alarm -sensor damaged -loss of alarm | -Pressure Issue -imprecision -incorrect measurement
FOT module | Disconnection / Tube blocked alarm: • Check ET tube • Patient disconnected | -tubing blocked -fail to disable alarm -device disconnected | -output issue -compliance issue -application network issue
FOT module | Any system failure: • Checksum conductor PIC • Checksum monitor PIC • COM interface • Cooling fan defect • DIO interface • I2C interface • Input pressure blender • Low physical memory • parallel interface • Safety relay defect • SPI interface • Voltage monitoring | -system failure -fail to monitor power event -fail to check voltage periodically | -inadequate user interface -no pressure -loss of power
FOT module | Alarm actions: • Applying Manual breath • Applying or pre-set another ventilation mode • Changing Date / time • Changing Patient range • Changing the Bias Flow setting • Changing the value of any Ventilation parameters, except O2, O2 Flush and Trigger sensitivity • Changing Trigger type • Changing VL or VG state • Turning OFF Flow sensor • Turning ON lung recruitment | -incorrect breath type -improper alarm -lack of breath type | -imprecision -tidal volume fluctuations -Pressure Issue
FOT module | warning: The “Manual Breath” can be activated during FOT, but this will deactivate FOT. | -incorrect breath type -confused breath type -fail to activate | -leak -particulates -disconnection
FOT module | warning: To re-enable FOT after a Manual Breath /Sustained Inflation, you must restart FOT from the beginning. | -fail to initiate breath -fail to adjust manually -fail to replace manually | -output issue -unauthorized access to computer system -no flow
FOT module | warning: The Start / Stop button can be used to interrupt mechanical ventilation for two minutes, but this will deactivate FOT. | -button stuck on -fail to indicate button status -fail to deactivate | -Blockage within Device or Device Component -user used incorrect product for intended use -medication error
FOT module | warning: Any FOT measurements taken during Standby are NOT valid | -measurement range out of specification -measurement data corrupted -measurement high | -device remains activated -Product Quality Issue -unstable capture threshold
FOT module | warning. Inaccurate measurements may be caused by: • Flow sensor not calibrated or defective; • Incorrect pressure measurement line application or use; • Excessive patient spontaneous breathing (strong active respiratory drive); • Excessive leakage at the patient interface. | -inaccurate measurement -incorrect measurement -fail to measure | -programming issue -Pressure Issue -insufficient flow or underinfusion
FOT module | Perform the following steps to complete an FOT procedure: Terminate any ongoing EtCO2 measurements during FOT procedure as it can affect the results. | -fail to terminate process completely -fail to avoid incorrect connection terminal -incorrect bg measurement data | -use of device issue -not audible alarm -incorrect measurement
FOT module | Ventilate in one of the compatible ventilation modes with an active flow sensor | -incorrect ventilation mode -inappropriate ventilation support -fail to activate | -imprecision -Pressure Issue -defective item
FOT module | Verify that the flow sensor in the selected mode is calibrated. | -inaccurate sensor measurement data -fail to calibrate manually -fail to recover sensor | -use of device issue -incorrect measurement -programming issue
FOT module | If Flow sensor is NOT calibrated, calibrate the Flow sensor and continue with the procedure. | -inaccurate sensor measurement data -fail to recover sensor -calibration failed | -Pressure Issue -Failure to Calibrate -incorrect measurement
FOT module | Go to the Curves / Loops menu by pressing the Graphics button on the front panel and select FOT submenu. | -fail to indicate button status -inappropriate interface design -fail to log in | -human-device interface issue -inadequate user interface -application security issue
FOT module | Set the three adjustable parameters (FOTsteps, low and high-pressure settings) according to the needs (FOT contains a recruitment and a derecruitment phase starting at the set low pressure level and turning back at the high-pressure level). | -fail to adjust settings -insufficient pressure -inappropriate stimulation rate | -use of incorrect control settings -pressure issue -fluid leak
FOT module | Adjust the pressure related alarms to avoid unnecessary patient alarms during the maneuver. | -unnoticeable alarm -improper alarm -fail to activate alarm | -device alarm system issue -Defective Alarm -improper alarm
FOT module | When all FOT parameters are set appropriately for the patient’s respiratory system, FOT is ready to be initiated by pressing the ON / OFF button, which will turn Green and a confirmation message will appear on the display. | -incorrect parameter -fail to indicate button status -incorrect message display | -failure to deploy -use of incorrect control settings -instruction for use issue
FOT module | After confirmation (clicking on Yes) the status bar is updated to “active mode + FOT” and the ventilation remains stable to the set low pressure level and a timer starts to count back from 180 seconds (stabilization period). | -fail to update indicator -incorrect limit setting -fail to indicate pump status | -use of incorrect control settings -pressure issue -use of device issue
FOT module | At the end of the 180 seconds or by pushing the Next button, FOT period is started by either changing to CPAP mode and applying forced oscillation with the FOT parameters (in FOT-conventional) or changing HFOV to the FOT parameters (in FOT-HFO). | -sticky button -fail to enter mode -force out of range | -inaccurate delivery -incorrect software programming calculations -failure to deliver
FOT module | If the conditions are acceptable (pressure is reached, and leak is low) the three seconds measurement takes place. | -fail to measure -inappropriate measurement unit -inadequate measure unit | -use of device issue -fluid leak -deployment issue
FOT module | If the measurement passes the Quality check, the calculated reactance measurement is displayed as a squared dot on FOT graph and the Ventilation mode is changed back to the recent Ventilation mode. | -incorrect measurement -incorrect message display -inaccurate measurement | -incorrect display -fluid leak -no display or display failure
FOT module | If the measurement does NOT pass the Quality check, or other problems were identified during FOT period, the measurement must be repeated. This can be performed by using the Repeat button. | -fail to pass check -fail to measure -fail to generate periodic notification | -use of device issue -human-device interface issue -unstable
FOT module | After a successful measurement you can proceed to the next step by using the Next step button. The next PEEP/ Pmean values are displayed in the information field (PEEPnext / Pmean next and PIPnext). | -incorrect measurement -fail to incorporate step sequence -inadequate message display | -fluid leak -inadequate instructions for healthcare professional -incorrect display
FOT module | If the patient shows signs of discomfort or decreased cardiac output during the recruitment phase because of the higher pressures, the upcoming higher -pressure level measurements can be skipped by pressing the Reverse button and the whole process can be continued at the previous step as a first step of the derecruitment. | -fail to cancel -loss of indication -loss of conduction | -pressure issue -improper device output -incorrect measurement
FOT module | After completing all the desired steps in the FOT procedure, normal ventilation resumes at FOT low PEEP and the FOT graph displays the reactance behaviour during the complete recruitment-derecruitment maneuver. | -fail to resume -fail to maintain ventilation mode -inadequate ventilation | -device expiration issue -fluid leak -no display or display failure
FOT module | The last calculation date and time is also updated with the last measurement’s time | -incorrect time and date -fail to update time -incorrect interval calculation | -use of incorrect control settings -reset issue -use of device issue
FOT module | You can then assess PEEP (CPAP) or MAP (CDP) accordingly if there is a need for readjustment and optimization. | -inadequate readings -inadequate markings -fail to document feedback | -failure to reset -fluid leak -medication error
FOT module | NOTE: If this situation occurs, the FOT procedure can be completed but with having less measurements than planned. | -loss of measurement -lack of procedure -failure to perform | -defective item -improper or incorrect procedure or method -environmental control or utility issue
FOT module | NOTE: In case FOT is terminated, ventilation continues the last / current pressure levels and the whole FOT procedure must be restarted from the first measurement (step 4). Use only recommended HFO breathing circuit. | -fail to maintain ventilation mode -fail to activate system -fail to replace manually | -use of incorrect control settings -deployment issue -fluid leak
FOT module | The following graphs illustrate a complete FOT maneuver in real time showing pressure levels in FOT Conventional mode with the following settings and how the results are displayed: | -incorrect graph -inaccurate pressure reading -incorrect graph format | -failure to deploy -environmental control or utility issue -reset issue
FOT module | FOT post data analysis: The exported system log files can be used for post data analysis. The log files contain the FOT settings and the measurement results (Xrs and Resistance) at each pressure level. | -lack of log file -inadequate log events -insufficient log space | -patient data issue -data issue -compliance issue
service and maintenance | WARNING: Maintenance and safety inspections must be performed by ACUTRONIC Medical Systems trained personnel with access to suitable testing and measuring equipment. | -fail to validate basic safety -inappropriate safety factor -lack of test tool | -human-device interface issue -compliance issue -device cleaning issue
service and maintenance | WARNING: The user of the device is responsible for performing routine maintenance when scheduled, and is also responsible for notifying ACUTRONIC Medical Systems in a timely basis when service needs to be performed. | -fail to perform inspection -lack of maintenance process -inadequate warning | -failure to sense -Difficult to Position -therapeutic or diagnostic output failure
service and maintenance | Maintenance Interval Indicator. The symbol will appear in the Information bar 30 days before the next maintenance has to be performed. | -loss of indication -missing symbol -lack of maintenance and repair document | -defective component -device maintenance issue -device expiration issue
service and maintenance | Always clean and disinfect equipment or equipment components before any maintenance – including when returning the equipment for repair. | -improper disinfection -device damaged by inappropriate cleaning -material degradation by inappropriate cleaning | -fitting problem -Material Integrity Issue -no device output
service and maintenance | Every 12 months: Perform the following work: • Check alarm and limit value functions • Check Electrical connections • Check Pressure connections • Check Safety Shutdown Calibration • Device check as specified by manufacturer | -fail to check voltage periodically -fail to set or reset alarm manually -fail to activate alarm | -output issue -device maintenance issue -patient data issue
service and maintenance | Every 12 months: Replace the following components: • Air and Oxygen Input Filter • Cooling Air Filter (casing bottom) • Exhalation membrane • Flow sensor • Flow sensor cable • O2 sensor | -fail to replace component -component missing -missing replacement tool | -Pressure Issue -Break -device clogged
service and maintenance | Every 4 years: Replace the following components: • Air and Oxygen Input Filter • Battery Pack • Calibration Valves • Cooling Air Filter (casing bottom) • Exhalation membrane • Flow sensor • Input pressure regulator pressurized Air / Oxygen • Proportional Valves pressurized Air / Oxygen mixer • Lithium Battery2 • Membrane HFO module • O2 sensor | -fail to replace component -component incompatible -missing replacement tool | -defective component -fail-safe mechanism issue -device clogged
Sterilization / cleaning / disinfection Ventilator | WARNING: The device must NOT be sterilized under any circumstances | -inappropriate sterilization cycle -sterility compromised -sterilization cycle failed | -device cleaning issue -Failure to Power-Up -programming issue
Sterilization / cleaning / disinfection Ventilator | Preparation for disinfection: Switch OFF fabian before cleaning and unplug at the mains in order to prevent damage due to penetration by liquid | -device damaged by inappropriate cleaning -failed to be disinfected -liquid ingress | -mechanical issue -disassembly -difficult to open or close
Sterilization / cleaning / disinfection Ventilator | To ensure that no contamination gets into the gas path, make sure all gas connectors (O2 and Air hoses on the backwall, FG / nCPAP port, Insp port and Prox port on frontside) are covered (e.g. with caps from transport or circuits) before cleaning. | -gas leakage -contamination during use -connector edge leakage | -use of device issue -leak -data back-up problem
Sterilization / cleaning / disinfection Ventilator | Remove any visible soiling on the fabian enclosure with the cleaning agent based on the option selected from Section 14.1.2. “Manual disinfection options” | -improper disinfection -inadequate disinfection -material degradation by inappropriate cleaning | -Device displays error message -device sensing issue -failure to align
Sterilization / cleaning / disinfection Ventilator | Manual disinfection options. Option A: CaviWipes Surface Disinfectant Wipes; Option B: 10% bleach (Clorox EPA REG. NO.: 5813-1); Option C: 5% H2O2, Hydrogen Peroxide (EPA REG. NO.: 335-1) | -inadequate disinfection -improper disinfection -reduced washable | -improper device output -device cleaning issue -not audible alarm
Sterilization / cleaning / disinfection Ventilator | Disinfectant Wipes: Unfold the CaviWipe and wipe the surface of the enclosure completely; Wipe the enclosure with a clean lint-free cloth dampened with purified water for one minute; Allow to dry in the air | -device damaged by inappropriate cleaning -improper disinfection -material degradation by inappropriate cleaning | -device cleaning issue -failure to prime -defective item
Sterilization / cleaning / disinfection Ventilator | NOTE: The surface treated must remain visibly moist for three minutes. Use more wipes if necessary, to keep the surface moist for three minutes | -dry paper -moisture -loss of pressure | -use of device issue -device difficult to setup or prepare -mechanical issue
Sterilization / cleaning / disinfection Ventilator | 10% bleach: 1. Prepare a 10% bleach solution using purified water. Use a clean lint-free cloth dampened with the bleach solution. 2. Wipe the enclosure completely. Remove any heavy soiling with additional dampened cloths if necessary. 3. Wipe the enclosure with a clean lint-free cloth dampened with purified water for one minute. 4. Allow to dry in the air. | -device damaged by inappropriate cleaning -material degradation by inappropriate cleaning -water proof failed | -device cleaning issue -use of device issue -Inflation Issue
Sterilization / cleaning / disinfection Ventilator | 5% H2O2, Hydrogen Peroxide: 1. Prepare a 5% H2O2 solution using purified water. Use a clean lint-free cloth dampened with the H2O2 solution 2. Wipe the enclosure completely. Remove any heavy soiling with additional dampened cloths if necessary. 3. Wipe the enclosure with a clean lint-free cloth dampened with purified water for one minute. 4. Allow to dry in the air. | -water resistant failed -water ingress -device damaged by inappropriate cleaning | -device cleaning issue -use of device issue -Inflation Issue
Sterilization / cleaning / disinfection Ventilator | NOTE: The surface treated must remain visibly moist for five minutes. Use more dampened cloths if necessary, to keep the surface moist for five minutes. | -sheath surface compromised -jacket wet -dry paper | -break -use of device issue -market issue
Sterilization / cleaning / disinfection Flow sensor cable | Flow sensor cable. Preparation for disinfection: 1. Switch OFF fabian before cleaning and unplug at the mains in order to prevent damage due to penetration by liquid. 2. Remove the flow sensor from the flow sensor cable. 3. Remove any visible soiling on the flow sensor cable with the cleaning agent using a manual disinfection option listed in section 14.2.1.2. “Manual disinfection options” | -cable broken or damaged -difficult to remove -inadequate disinfection | -incorrect measurement -difficult to open or close -mechanical jam
Sterilization / cleaning / disinfection Flow sensor cable | Flow sensor cable. Manual disinfection options: Option A: CaviWipes Surface Disinfectant Wipes (EPA REG. NO.: 46781-14) 1. Unfold the CaviWipe and wipe the surface of the enclosure completely. NOTE: The surface treated must remain visibly moist for three minutes. Use more wipes if necessary, to keep the surface moist for three minutes. 2. Wipe the cable with a clean lint-free cloth dampened with purified water for one minute. 3. Allow to dry in the air. | -improper disinfection -inadequate disinfection -part scratched | -device cleaning issue -protective measure issue -failure to deploy
Sterilization / cleaning / disinfection Flow sensor cable | Flow sensor cable. Manual disinfection options: Option B: 10% bleach (Clorox EPA REG. NO.: 5813-1) 1. Prepare a 10% bleach solution using purified water. Use a clean lint-free cloth dampened with the bleach solution. 2. Wipe the cable completely. Remove any heavy soiling with additional dampened cloths if necessary. NOTE: The surface treated must remain visibly moist for five minutes. Use more dampened cloths if necessary, to keep the surface moist for five minutes. 3. Wipe the cable with a clean lint-free cloth dampened with purified water for one minute. 4. Allow to dry in the air. | -inadequate disinfection -material degradation by inappropriate cleaning -improper disinfection | -device cleaning issue -instruction for use issue -Pressure Issue
Sterilization / cleaning / disinfection Flow sensor cable | Flow sensor cable. Manual disinfection options: Option C: 5% H2O2, Hydrogen Peroxide (EPA REG. NO.: 335-1) 1. Prepare a 5% H2O2 solution using purified water. Use a clean lint-free cloth dampened with the H2O2 solution 2. Wipe the cable completely. Remove any heavy soiling with additional dampened cloths if necessary. NOTE: The surface treated must remain visibly moist for five minutes. Use more dampened cloths if necessary, to keep the surface moist for five minutes. 3. Wipe the cable with a clean lint-free cloth dampened with purified water for one minute. 4. Allow to dry in the air | -inadequate disinfection -improper disinfection -material degradation by inappropriate cleaning | -device cleaning issue -Pressure Issue -human factors issue
Sterilization / cleaning / disinfection Flow sensor cable | Flow sensor: NOTE: Prior to sterilization, ensure that the flow sensor has been replaced during the 12 month maintenance or after 30 Sterilization cycles, whichever comes first. | -sensor damaged -loss of sensor data -fail to reset sensor | -Inflation Issue -Pressure Issue -failure to capture
Sterilization / cleaning / disinfection Flow sensor cable | WARNING: Conduct hygienic preparation of flow sensor prior to each use. | -inadequate flow -inappropriate flow pattern -fail to prepare | -Inflation Issue -Pressure Issue -excess flow or over-infusion
Sterilization / cleaning / disinfection Flow sensor cable | Preparation for sterilization by automated cleaning: 0. Remove the flow sensor from the flow sensor cable. 1. Rinse under running cold tap water (<35°C / 95°F) to remove excess soil. 2. Prepare a detergent bath using Neodisher Mediclean Forte solution at the manufacturer’s recommendation of 10 mL per Liter using utility (tap) water 20°C / 68°F - 30°C / 86°F. 3. Transfer the flow sensor on to the 4-Level manifold rack accessory (or other appropriate rack system) contained inside the washer for processing. 4. Select the appropriate cycle as listed below: 5. After the drying process package immediately for sterilization in packaging suitable for vapor sterilization. Ensure that packaging stays dry before autoclaving. | -water resistant failed -unable onehand to remove -fail to remove | -Blockage within Device or Device Component -Inflation Issue -incorrect measurement
Sterilization / cleaning / disinfection Flow sensor cable | NOTE: Articles must be individually single pouched in a qualified pouch ( for example: Cardinal Health self-sealing pouch CAT# 92713). Each of the non-critical devices inscope are thermostable and have been validated for steam sterilization. | -single member adjustment failed -single use violated -single member uninstallation failed | -device difficult to setup or prepare -component incompatible -device packaging compromised
Sterilization / cleaning / disinfection Flow sensor cable | Sterilization options for flow sensor: The recommended sterilization method is steam sterilization by autoclave | -inappropriate sterilization cycle -inappropriate sterilization -sterilization cycle failed | -Inflation Issue -Device Disinfection or Sterilization Issue -Failure to Sterilize
Sterilization / cleaning / disinfection Flow sensor cable | Sterilization option A for flow sensor. Set the autoclave with the following parameters: 4 preconditioning pulses, Steam cycle at 134°C / 273°F for 3 minutes, Dry cycle for 30 minutes, Individually single pouch the cassette articles, Autoclave the cassette articles using the cycle parameters | -incorrect sensor orientation -inaccurate sensor measurement data -use incorrect control settings | -failure to cycle -Device Handling Issue -Blockage within Device or Device Component
Sterilization / cleaning / disinfection Flow sensor cable | Sterilization option B for flow sensor. Set the autoclave with the following parameters, 4 preconditioning pulses, Steam cycle at 134°C / 273°F for 4 minutes, Dry cycle for 30 minutes, Individually single pouch the cassette articles, Autoclave the cassette articles using the cycle parameters in Step 1 | -inaccurate sensor measurement data -fail to adjust settings -use incorrect control settings | -failure to cycle -Blockage within Device or Device Component -output issue
Sterilization / cleaning / disinfection Flow sensor cable | Sterilization option B for flow sensor. Set the autoclave with the following parameters: 4 preconditioning pulses, Steam cycle at 134°C / 273°F for 5 minutes, Dry cycle for 30 minutes, Individually single pouch the cassette articles, Autoclave the cassette articles using the cycle parameters in Step 1 | -incorrect sensor orientation -inaccurate sensor measurement data -use incorrect control settings | -failure to cycle -Blockage within Device or Device Component -Device Handling Issue
Sterilization / cleaning / disinfection Flow sensor cable | Duration of use: The flow sensor must be replaced after 30 sterilization cycles and also in case that any detected defect is affecting the accessory performance. | -fail to replace component -fail to reset sensor -fail to replace manually | -Pressure Issue -mechanical jam -device contamination with chemical or other material
Sterilization / cleaning / disinfection | WARNING: Conduct hygienic preparation of exhalation membrane and exhalation membrane holder prior to each use. | -fail to prepare -inappropriate technique -inadequate process | -device cleaning issue -improper device output -defective component
Sterilization / cleaning / disinfection exhalation membrane | Disassembly. Remove the exhalation membrane holder and membrane from the ventilator: Disassemble the membrane holder from the expiratory limb port by rotating the holder a quarter of turn in counterclockwise direction and then pull it out from the port. | -difficult to remove -fail to remove -inefficient position | -no pressure -bent -arcing
Sterilization / cleaning / disinfection exhalation membrane | NOTE: Containers are recommended to transport used accessories between the NICU and the cleaning/sterilization area. | -inappropriate transportation configuration -inappropriate transportation environment or condition -container mishandled | -device markings issue -Material Integrity Issue -user used incorrect product for intended use
Sterilization / cleaning / disinfection exhalation membrane | Before cleaning the membrane holder and the membrane, disassemble the exhalation membrane from the membrane holder by pulling it out from the second circumference of the exhalation valve membrane. | -missing symbol -loose component -fail to insert | -no pressure -unstable -fluid leak
Sterilization / cleaning / disinfection exhalation membrane | Preparation for sterilization by automated cleaning: Rinse under running cold tap water (<35°C / 95°F) to remove excess soil. | -fail to clean adequately -insufficient discharge time -inappropriate usage environment | -Infusion or Flow Issue -Use of Device Issue -Blockage within Device or Device Component
Sterilization / cleaning / disinfection exhalation membrane | Prepare a detergent bath using Neodisher Mediclean Forte solution at the manufacturer’s recommendation of 10 mL per Liter using utility (tap) water 20°C / 68°F - 30°C / 86°F. | -water resistant failed -water proof failed -water ingress | -fluid leak -Leak -Difficult to Position
Sterilization / cleaning / disinfection exhalation membrane | Transfer the cassette items on to the 4-Level manifold rack accessory (or other appropriate rack system) contained inside the washer for processing. | -cassette leakage -missing packaging -part mismatched | -break -Defective Component -Burn of Device or Device Component
Sterilization / cleaning / disinfection exhalation membrane | Select the appropriate cycle as listed below: Phase, Recirculation Time (Minutes), Temperature, Detergent Type and Concentration | -fail to cycle -inappropriate program selected -fail to set limit correctly | -medication error -use of incorrect control settings -patient data issue
Sterilization / cleaning / disinfection exhalation membrane | After the drying process, package immediately for sterilization in packaging suitable for vapor sterilization. Ensure that packaging stays dry before autoclaving. | -insufficient package -missing packaging -package broken or damaged | -device packaging compromised -market issue -Size Incorrect for patient
Sterilization / cleaning / disinfection exhalation membrane | NOTE: Articles must be individually single pouched in a qualified pouch ( for example: Cardinal Health self-sealing pouch CAT# 92713). Each of the non-critical devices inscope are thermostable and have been validated for steam sterilization. | -single member adjustment failed -single use violated -single member uninstallation failed | -device difficult to setup or prepare -component incompatible -device packaging compromised
Sterilization / cleaning / disinfection exhalation membrane | Sterilization options. The recommended sterilization method is steam sterilization by autoclave. The following validated cycles are suitable for sterilization: | -inappropriate sterilization cycle -inappropriate sterilization -sterilization inefficacy | -Device Disinfection or Sterilization Issue -market issue -therapy delivered to incorrect body area
Sterilization / cleaning / disinfection exhalation membrane | Option A: 1. Set the autoclave with the following parameters, 1.1. 4 preconditioning pulses, 1.2. Steam cycle at 134°C / 273°F for 3 minutes, 1.3. Dry cycle for 30 minutes 2. Individually single pouch the exhalation membrane and membrane holder separately. 3. Autoclave the exhalation membrane and membrane holder using the cycle parameters in Step 1 above. | -fail to adjust settings -incorrect settings -inappropriate pacing mode | -packaging issue -component or accessory incompatibility -device difficult to setup or prepare
Sterilization / cleaning / disinfection exhalation membrane | Option B: 1. Set the autoclave with the following parameters 1.1. 4 preconditioning pulses 1.2. Steam cycle at 134°C / 273°F for 4 minutes 1.3. Dry cycle for 30 minutes 2. Individually single pouch the exhalation membrane and membrane holder separately. 3. Autoclave the exhalation membrane and membrane holder using the cycle parameters in Step 1 above. | -fail to dry -fail to adjust settings -improper isolation | -device difficult to setup or prepare -component or accessory incompatibility -image resolution poor
Sterilization / cleaning / disinfection exhalation membrane | Option C: 1. Set the autoclave with the following parameters: 1.1. 4 preconditioning pulses 1.2. Steam cycle at 134°C / 273°F for 5 minutes 1.3. Dry cycle for 30 minutes 2. Individually single pouch the exhalation membrane and membrane holder separately. 3. Autoclave the exhalation membrane and membrane holder using the cycle parameters in Step 1 above. | -fail to adjust settings -inappropriate pacing mode -inappropriate pacing rate | -device difficult to setup or prepare -component or accessory incompatibility -failure to convert to back-up
Sterilization / cleaning / disinfection exhalation membrane | Duration of use: The exhalation membrane and membrane holder must be replaced after 30 sterilization cycles and in case of any detected defect that is affecting the accessory performance. | -fail to replace component -degradation of essential performance -upgrade failed | -Tip-over -failure to sterilize -mechanical jam
Sterilization / cleaning / disinfection exhalation membrane | Use of cleaning solutions other than recommended in section 14 or use of nebulizing drugs may reduce the duration of use. | -reduced wearing time -fail to reduce risk -reduced washable | -free or unrestricted flow -Inflation Issue -unintended movement
Volume Guarantee (VG) | Mechanical Ventilation is required to manage Neonates with severe respiratory failure. Pressure-limited Ventilation (PLV), delivering an operator set Peak Inspiratory Pressure (PIP), has traditionally been used to control the arterial Carbon dioxide (PaCO2). However the Tidal Volume (VT) may fluctuate due to the patient's breathing effort, changes in lung mechanics and variable Endotracheal Tube (ETT) Leak. | -fail to maintain ventilation mode -inappropriate ventilation support -fail to hold pressure | -Pressure Issue -Increase in Pressure -use of incorrect control settings
Volume Guarantee (VG) | For some patients, a strategy of targeting an operator set exhaled tidal volume (VTe) while maintaining pressure within a prescribed range can maintain consistent delivered tidal volume. | -fail to clear tidal volume -tidal volume fluctuation -fail to hold pressure | -delivery system failure -use of incorrect control settings -pressure issue
Volume Guarantee (VG) | Strategy for targeting expired VTe can reduce variability from ETT Leaks unless they are very large. In case of an improvement of lung compliance, the inspiratory pressure is automatically reduced. | -fail to reduce risk -timing variation -fail to follow plan | -incorrect measurement -increase in pressure -use of device issue
Volume Guarantee (VG) | The fabian series of neonatal/pediatric microprocessor-controlled ventilators allow volume targeted ventilation (VTv). Tidal volume measurements are done with the flow sensor placed at the patient circuit wye piece. The flow sensor measures inspired, expired VT, and ETT Leak percentage. | -fail to set tidal volume limit -fail to calculate tidal volume -fail to update tidal volume | -imprecision -Pressure Issue -incorrect measurement
Volume Guarantee (VG) | Volume Guarantee (VG) Ventilation on fabian is a VTv-mode which targets the operator’s set VTe. | -inappropriate ventilation support -inadequate ventilation -incorrect ventilation mode | -use of incorrect control settings -use of device issue -fluid leak
VG: Conventional ventilation | There are three key operator settable parameters related to VG operation: 1. Vguarant: Targeted/guaranteed patient's expired volume (VTe). 2. Pinsp: Initial inspiratory pressure for VG first breath (also known as “test breath”). A test breath is needed to determine the patient’s lung mechanic at a starting target pressure. Targeted inspiratory pressure for the test breath is at the operator set PEEP + 75% of (set Pinsp – set PEEP). 3. Pmax: Maximum inspiratory pressure VG can adjust. | -insufficient pressure -low pressure -loss of pressure | -use of incorrect control settings -device difficult to setup or prepare -use of device issue
VG: Conventional ventilation | The VG function can be enabled once all VG related settings and alarm limits are appropriately set per the patient’s condition. Pushing the VG ON/OFF push-button (between the “Pmax” and “Vguarant” buttons) enables the VG function. | -fail to disable alarm -incorrect settings -fail to activate | -Device Alarm System Issue -break -medication error
VG: Conventional ventilation | The ventilator delivers a test breath after VG function is enabled. It will then automatically increase or decrease the inspiratory pressure for each subsequent breath so that the measured VTe reaches the set Vguarant. The maximum targeted inspiratory pressure change between two consecutive breaths of the same triggering category (i.e. time-triggered or patienttriggered) is limited to no more than 3 mbar. PSV breaths are not volume guaranteed. | -fail to activate test -incorrect breath type -fail to update pressure | -improper device output -failure to deploy -device operates differently than expected
VG: Conventional ventilation | The ventilator will activate the “VTe not reached / check settings” alarm if the set Pmax or set I-Flow is insufficiently high for the measured VTe to reach 94% of the set Vguarant for a predetermined duration (5 consecutive breaths for non-HFO modes; 10 seconds for HFO mode). An upward arrow will be displayed in each of the Pmax and I-Flow setting area. When this happens, increase one or both settings (make sure this is safe for the patient) to allow the ventilator to increase delivered volume and therefore more VTe. | -fail to activate alarm -fail to activate alert -fail to activate | -failure to cycle -Blockage within Device or Device Component -decrease in pressure
VG: Conventional ventilation | The minimum pressure difference between PEEP and target inspiratory pressure is 2 mbar. | -pressure out of range -inaccurate pressure reading -inappropriate pressure level | -use of incorrect control settings -pressure issue -Increase in Pressure
VG: Conventional ventilation | If the flow sensor measurements become unavailable, the VG function will be disabled and the pre-set Pinsp is used for ventilation. VG function will restart once flow sensor measurements become available again. | -loss of function -disable failed -fail to shut-down device | -Pressure Issue -imprecision -decrease in pressure
VG: Conventional ventilation | If the “Check ETT tube” alarm is activated, the VG function will be disabled and the pre-set Pinsp is used for ventilation. VG function will restart once the alarm is removed. | -fail to activate alarm -fail to disable alarm -disable failed | -break -no display or display failure -device alarm system issue
VG: Conventional ventilation | The following procedure describes the use of VG in SIPPV mode. The same procedure can be followed in the other conventional ventilation modes such as IPPV, SIPPV, SIMV+PSV and PSV where VG function is available. | -incorrect procedure -inappropriate usage environment -incorrect usage | -use of incorrect control settings -protective measure issue -failure to deploy
VG: Conventional ventilation | Step 1: Setup Ventilator in SIPPV mode with VG disabled and start ventilation. Initial parameter settings: | -inappropriate ventilation support -incorrect settings -inadequate configuration | -use of device issue -use of incorrect control settings -improper or incorrect procedure or method
VG: Conventional ventilation | NOTE: The suggested parameters and alarm settings are for explanation purposes. These must be carefully examined and appropriately set for each patient. | -incorrect parameter -incorrect settings -parameter confused | -use of device issue -device markings issue -medication error
VG: Conventional ventilation | Step 2:Select the appropriate VTe to be guaranteed by setting Vguarant to a target volume. A typical infant tidal volume target is 4 to 6mL/kg. | -fail to set tidal volume limit -fail to clear tidal volume -fail to deliver tidal volume | -device difficult to setup or prepare -pressure issue -use of incorrect control settings
VG: Conventional ventilation | Step 3: Adjust the “Pinsp” setting until the measured VTe (on the right side of the display) reaches the set Vguarant. | -adjustment failed -incorrect pin number -incorrect material origin | -device difficult to setup or prepare -use of incorrect control settings -failure to deploy
VG: Conventional ventilation | IMPORTANT: To ensure the VG function works appropriately, a small plateau pressure in the Pressure wave needs to be present. Plateau pressure means the patient lung pressure and the circuit pressure reached the equilibrium state and are stabilized. Sufficient I-Time and I-Flow are required for the inspiratory pressure to rise to the selected Pinsp and form a small plateau pressure, DO NOT set the inspiratory time to a value too short to allow the plateau pressure to be reached. | -fail to produce pressure -inappropriate pressure level -low pressure | -pressure issue -decrease in pressure -calibration issue
VG: Conventional ventilation | Step 4: Set in the Pmax area, the maximal pressure value to which the VG function can regulate. Pmax should be set somewhere between Pinsp setting and High PIP alarm setting. | -high flux rate -low pressure -over pressure | -pressure issue -use of incorrect control settings -failure to deploy
VG: Conventional ventilation | NOTE: To avoid unintentional breath abortions by reaching / exceeding High PIP, make sure that the High PIP alarm is set to at least +2mbar above Pmax | -fail to raise high priority alarm -incorrect breath type -fail to raise low priority alarm | -device alarm system issue -improper device output -defective item
VG: Conventional ventilation | Triggered and non-triggered breaths are independently supported based on lung compliance. Normally, a triggered breath has a lower Pinsp than a non-triggered one due to the effort provided by the patient. | -incorrect breath type -fail to initiate breath -confused breath type | -inaccurate delivery -mechanical issue -improper device output
VG: Conventional ventilation | Step 5: VG can now be activated by selecting the ON/OFF button. When VG function is enabled, this button will turn Green. A Red line will appear in the Pressure Wave diagram to mark the Pmax. | -fail to indicate button status -confused button -fail to activate | -failure to deploy -decrease in pressure -device operates differently than expected
VG: Conventional ventilation | NOTE: If the Targeted Volume cannot be reached with the set Pmax, “VTe not reached / check settings” alarm will appear. An upward pointing arrow will appear in the “Pmax” and the “I-Flow” parameter fields to indicate the maximum pressure has been reached. Increase the Pmax or I-Flow if necessary. | -fail to activate alarm -incorrect parameter -incorrect settings | -particulates -use of device issue -incorrect display
VG: Conventional ventilation | If Volume Guarantee is added to PSV, the Ventilator automatically is adjusting the PPSV level necessary to maintain the preset Target Volume. In case of an Apnea, the Ventilator will start cycling at a preset rate and PBackup. As soon as spontaneous activity restarts, the backup stops. | -fail to adjust volume -fail to adjust settings -low volume | -leak -use of device issue -unstable
VG: Conventional ventilation | The minimum pressure difference between PEEP and PPSV is always 2 mbar and PPSV is limited not to exceed the set Pinsp. | -pressure out of range -over pressure -inaccurate pressure reading | -use of incorrect control settings -improper flow or infusion -device contamination with biological material
VG: Conventional ventilation | PSV breaths and Apnea backup breaths are Volume Targeted Breaths in this mode. | -incorrect breath type -inappropriate flow volume -fail to adjust volume | -failure to deliver -improper device output -parameter calculation error due to software problem
VG: Conventional ventilation | The delay to start backup Ventilation is set with the Apnea time in Alarm Limit screen. If Apnea is set to OFF, backup starts after E-time. | -fail to set or reset alarm automatically -inappropriate timeout setting -incorrect settings | -use of incorrect control settings -defective item -deployment issue
VG: Conventional ventilation | In PSV+VG mode, the patient controls the onset of Inspiration as well as the end of Inspiration, offering a total synchrony with patients breathing pattern. | -fail to control flow -loss of control -pressure control failed | -reset issue -incorrect measurement -defective item
VG: Conventional ventilation | If patient lung compliance is improving, the Ventilator automatically is using the Lowest Pressure, necessary to maintain the Preset Target Volume. | -fail to hold pressure -low pressure -fail to reduce pressure | -user used incorrect product for intended use -unstable -device operates differently than expected
VG: Conventional ventilation | In case of a Respiratory fatigue and a ceasing Inspiratory effort, the ventilator automatically provides an Apnea Backup Ventilation with the preset parameters for I-time, Frequency and Pbackup. | -inadequate ventilation -inappropriate ventilation support -incorrect ventilation mode | -use of device issue -improper or incorrect procedure or method -use of incorrect control settings
VG: Conventional ventilation | The delay for the onset of the Apnea Backup Ventilation is set in the Alarm screen, by setting the Apnea Delay Time. If Apnea Backup Ventilation is set to OFF, the ventilator starts Apnea Backup Ventilation immediately after E-time, otherwise after Preset Apnea Time. | -fail to set or reset alarm automatically -fail to set or reset alarm manually -positioning delayed | -use of incorrect control settings -defective item -not audible alarm
VG: Conventional ventilation | Using a delay time will determine the apnea period before backup ventilation begins. The Apnea Backup breaths are Volume Guaranteed breaths. | -incorrect breath type -low volume -delayed monitoring | -reset issue -improper device output -device expiration issue
VG operation: HFO ventilation | fabian HFO+VG operates in the same principle as conventional ventilation mode with VG except with a different set of parameters and monitor values. The ventilator automatically adjusts oscillating pressure amplitude (peak-to-peak) to achieve the operator set targeted VTe.There are three key operator adjustable parameters related to VG operation: 1. Vguarant: Targeted/guaranteed patient's exhaled tidal volume (VTe). 2. Amp: Initial oscillating pressure amplitude at start of VG. 3. Ampmax: Maximum oscillating pressure amplitude VG can adjust. | -fail to adjust settings -adjustment failed -incorrect settings | -device difficult to setup or prepare -component or accessory incompatibility -component incompatible
VG operation: HFO ventilation | The VG function can be enabled once all VG related settings and alarm limits are appropriate ly set per the patient’s condition. Pushing the VG ON/OFF push-button (between the “Ampmax” and “Vguarant” buttons) enables VG function. | -incorrect settings -fail to disable control -fail to activate system | -break -failure to deploy -no display or display failure
VG operation: HFO ventilation | Once VG is enabled, the ventilator will automatically increase or decrease its oscillating pressure amplitude so that the measured VTe reaches the set Vguarant. | -fail to adjust settings -fail to update pressure -incorrect settings | -failure to deploy -deployment issue -programming issue
VG operation: HFO ventilation | The ventilator will activate the “VTe not reached / check settings” alarm if the set “Ampmax” is insufficiently high for the measured VTe to reach 94% of the set Vguarant for 10 seconds. An upward arrow will be displayed in the “Ampmax” setting area. When this happens, increase “Ampmax” (make sure this is safe for the patient) to allow the ventilator to increase delivered volume and therefore more VTe. | -fail to activate alarm -fail to activate -fail to activate alert | -improper alarm -Blockage within Device or Device Component -Device Alarm System Issue
VG operation: HFO ventilation | If the flow sensor measurements become unavailable, the VG function will be disabled, and the pre-set “Amp” is used for ventilation. VG function will restart once flow sensor measurements become available again. | -fail to disable alarm -fail to shut-down device -fail to initiate pulse | -Pressure Issue -imprecision -decrease in pressure
VG operation: HFO ventilation | If the “Check ETT tube” alarm is activated, the VG function will be disabled, and the pre-set “Amp” is used for ventilation. VG function will restart once the alarm is removed. | -fail to activate alarm -fail to disable alarm -disable failed | -no display or display failure -break -device alarm system issue
VG operation: HFO ventilation | Follow similar concept described in section 16.2: How to start the VG function for conventional ventilation to setup HFO+VG. It is important to note ventilation parameter and alarm settings for HFO are considered instead of for conventional ventilation. | -parameter confused -incorrect parameter -incorrect settings | -use of incorrect control settings -device difficult to setup or prepare -failure to analyze rhythm
suction system | Suction can be performed in the following Ventilation modes: • CPAP • HFO • IPPV • PSV • SIMV • SIMV+PSV • SIPPV | -incorrect ventilation mode -fail to maintain ventilation mode -inappropriate ventilation support | -improper or incorrect procedure or method -Device Operates Differently than Expected -component incompatible
suction system | The suction catheter system must be placed between the Flow sensor and the patient. (See the following picture). | -inappropriate suction setting -inappropriate patient type -inappropriate indication | -compliance issue -component missing -material rigid or stiff
suction system | 1. Tube set coming from the fabian +nCPAP evolution. 2. Flow sensor must be placed on the device side of the suction system. 3. Closed Suction System. 4. Patient side. | -incorrect tubing placement -fail to retain closed -tubing blocked | -compliance issue -mechanical jam -Blockage within Device or Device Component
suction system | warning: Before suctioning, do a preoxygenation of the patient. | -inappropriate suction setting -undesired patient data changes -inadequate breathing mode | -use of device issue -device unsafe to use in environment -missing test results
suction system | Warning: Choose the right catheter for the suctioning, otherwise, a “Vacuum too High” alarm is possible. | -catheter misplacement -catheter obstructed -catheter incompatible with liquid | -device alarm system issue -defective component -device maintenance issue
suction system | Silence the alarms before suction (two minutes) for not disturbing the suctioning. After the suction reactivate the alarm again. | -fail to silence alarm -alarm notified -alarm interrupted | -Device Alarm System Issue -device alarm system issue -improper alarm
Technical specifications | Ambient conditions: During operation; During transportation; During storage | -inappropriate environment condition -inadequate conditions -inappropriate pressure level | -temperature issue -Device Markings Issue -parameter calculation error due to software problem
Technical specifications | IMPORTANT: For moving the device between ambient and controlled temperatures, removing the device from storage, and before use; acclimate the fabian for 12 hours. | -moving speed slow -fail to move -inappropriate indication | -failure to sense -use of incorrect control settings -Biocompatibility Issue
Technical specifications | The minimum and maximum ambient temperatures in the clinical area will not affect the accuracy of the flow, pressure or oxygen measurement. | -temperature out of range -low pressure -incorrect measurement | -incorrect measurement -Medical Gas Supply Problem -Battery Issue
Technical specifications | Monitoring: Acoustic Pressure of Alarm tones. Alarm tone sequence IEC 60601-1-8: High Priority, Medium Priority, Low Priority. A weighted sound pressure level in 1m distance: 46 dB | -sound pressure level out of range -sound pressure level out of specification -low pressure | -improper alarm -Device Alarm System Issue -device alarm system issue
Technical specifications | Measuring: Airway measurement Range, Breath volume measurement Range, Volume Controlled Breaths, Pressure Controlled Breaths, Inspiratory Oxygen Concentration, Max. Error delivered vs. Set FiO2, Max. Error delivered vs. Set FiO2, Response Time to change in FiO2, Drift, Tube Leak Range, Dynamic Compliance Range | -measurement range out of specification -measurement high -inaccurate measurement | -use of device issue -parameter calculation error due to software problem -delivery system failure
Technical specifications | Resistance values: System Resistance, Inspiratory Resistance, Expiratory Resistance | -high resistance -incorrect value -resistance inadequate | -Connection Issue -Insufficient Flow or Underinfusion -Delivery System Failure
Technical specifications | Ventilation menu settings: Alarm delay, Automatic O2 calibration, E-flow, Manual breath in HFO, Max. time Manual breath, NIV Leak compensation, NIV Tube set, Ppsv parameter, Ratio of I-time, Termination criteria PSV, Trigger, Unit for Pressure, Use BTB for Vt | -fail to adjust settings -confused settings -incorrect settings | -failure to deliver -improper device output -no pacing
Technical specifications | Ratings: Device Fuse, Internal Battery, IP Protection, Oxygen Supply, Power Rating, Power Supply, Pressurized Air Supply, Safety Classification (Applied Parts), Electrical Safety Classification, Sound Pressure | -loss of electrical connection -lack of fuse -incorrect electrode placement | -failure to power-up -defective component -Failure to Power-Up
Technical specifications | Data storage: Alarms, log, trend | -fail to store data -insufficient storage space -fail to view alarm settings | -patient data issue -unauthorized access to computer system -data issue
Technical specifications | Applied parts for the device are as follows: • Flow Sensor • Nebulizer • SpO2 Module (type BF) • Tube set (type B) | -missing parts -part mismatched -inadequate configuration | -Contamination During Use -incorrect display -failure to recalibrate
Technical specifications | The following internal device checks are performed on Startup: • COM : Communication between GUI and Conductor PIC. • DIO : Communication between GUI and Digital I/O. • I2C : Communication between GUI to AccuPIC, FRAM and Multiplexer. • PIF : Communication between GUI and Parallel Interface • Relias : realis Check from the SPI • SPI : Communication between GUI and Monitor PIC. • SPICks : SPI Checksum • Voltage : Voltage Error Check from the SPI | -fail to pass check -fail to schedule check -fail to clear message | -communication or transmission issue -instruction for use issue -wireless communication issue
Technical specifications | The Air / Oxygen Blender provides a Gas blend with adjustable proportions of Oxygen and Air. Gas from the central gas supply enters the device through the gas input connections. | -fail to adjust volume -adjustment failed -fail to adjust settings | -use of device issue -no display or display failure -failure to convert to back-up
Technical specifications | Blending of the gases occurs through two proportional valves. The Proportional Valves, which also act as return valves, prevent the return flow of a gas into the supply line of the other gas. | -valve stuck -valve blocked -valve leakage | -Leak -Fluid Leak -Malposition of Device
Technical specifications | An Oxygen sensor measures the Inspiratory Oxygen Concentration | -low oxygen concentration -measurement high -fail to update oxygen volume | -Pressure Issue -imprecision -medical gas supply problem
Technical specifications | Delivery of ventilation at the patient-connection port within the alarm limits set by the operator; or the following alarm conditions shall occur as appropriate: • Low battery as internal electrical power source nears depletion • Low and High Minute Volume • Low and High PEEP • Obstruction/Occlusion • Low and High Air/Oxygen Inlet Pressure • Air/Oxygen Gas Supply Failure or a technical alarm condition shall occur. | -fail to set limit correctly -low battery -inappropriate ventilation support | -not audible alarm -device alarm system issue -power conditioning issue
Technical specifications | Measurement accuracy of the oxygen respiratory gas monitor (RGM) is maintained within the IFU specification, ± (3% absolute); or Low and High Oxygen Level alarm conditions shall occur, as appropriate; or a technical alarm condition shall occur. | -low measurement accuracy -loss of measurement -fail to display accuracy | -incorrect measurement -device alarm system issue -calibration issue
Technical specifications | Measurement accuracy of the etCO2 RGM is maintained within the IFU specification: 0 to 38 mmHg ±2 mmHg 39 to 150 mmHg ±(5% of reading + 8% of (reading – 39 mmHg)); or Low and High CO2 Level alarm conditions shall occur, as appropriate; or a technical alarm condition shall occur. | -low measurement accuracy -orientation accuracy out of specification -measurement high | -incorrect measurement -use of device issue -volume accuracy issue
Technical specifications | Measurement accuracy of the SpO2 monitor is maintained within the IFU specification only within the range 70 to 100%: ± 2% pediatric patients, no motion conditions; ± 3% pediatric patients, motion conditions, and all neonatal patients; or Low and High SpO2 Level alarm conditions shall occur, as appropriate; or a technical alarm condition shall occur. | -measurement range out of specification -low measurement accuracy -orientation accuracy out of specification | -incorrect measurement -unintended system motion -calibration issue
Technical specifications | Measurement accuracy of the pulse rate from the SpO2 monitor is maintained within the IFU specification within the range 25 to 240 bpm: ± 3 bpm (no motion conditions); or ± 5 bpm (motion conditions); or Low and High Pulse Rate alarm conditions shall occur, as appropriate; or a technical alarm condition shall occur. | -orientation accuracy out of specification -low measurement accuracy -rate out of range | -incorrect measurement -use of device issue -unstable
EMC statement | fabian HFO is a MEDICAL APPLIANCE subject to specific precautionary measures regarding EMC and must be installed and started up according to the notices and the instructions in this manual. | -fail to install -installation difficult -inappropriate technique | -use of device issue -installation-related problem -Failure to Cycle
EMC statement | WARNING: Portable medical HF communication devices could impact fabian HFO. | -device vibration -device collided -loss of communication | -failure to pace or properly pace -wireless communication issue -Failure to Cycle
EMC statement | WARNING: NEVER use fabian HFO directly adjacent to, or stack with other apparatuses. If unavoidable, be sure to monitor the equipment for proper operation with this set-up. | -incorrect tool identification -fail to update -inappropriate environment | -fitting problem -compliance issue -use of device issue
EMC statement | Using non-manufacturer recommended accessories, cables or converters with fabian HFO could result in increased electromagnetic interference and reduce the immunity of fabian HFO. | -input originating from electromagnetic interference -increased sensitivity -magnetic interference | -therapeutic or diagnostic output failure -material integrity issue -chemical issue
EMC statement | The key performance characteristics of the fabian HFO ventilator are: • fabian HFO must operate within the defined specifications and the medical purpose. Failure to meet these specifications will result in respiration being stopped. • For this reason, a second stand-alone ventilation unit must always be available (for example: Resuscitation Bag). | -lack of characteristics -inappropriate program selected -fail to address deliverable | -installation-related problem -device difficult to setup or prepare -calibration issue
EMC statement | RF Emission CISPR 11: The fabian HFO device uses RF Energy solely for internal operation. Its RF emission is thus very low and interference with adjacent apparatuses is unlikely. | -fail to activate energy -device malfunction -fail to run on battery | -protective measure issue -communication or transmission issue -device maintenance issue
EMC statement | Harmonic Distortion IE 61000-3-2: The emissions characteristics of this equipment make it suitable for use in industrial areas and hospitals (CISPR 11 class A). If it is used in a residential environment (for which CISPR 11 class B is normally required) this equipment might not offer adequate protection to radio-frequency communication services. The user might need to take mitigation measures, such as relocating or re-orienting the equipment. | -inappropriate sound pressure -inappropriate environment condition -noise emitted | -use of device issue -protective measure issue -user used incorrect product for intended use
EMC statement | Electrostatic Discharge (ESD) IEC 61000-4-2: Floors should be wooden or concrete or tiled with ceramic tiles. For flooring made from synthetic materials, the Relative Air Humidity must be no less than 30%. | -material degradation by inappropriate cleaning -inappropriate floor level setting -inadequate architecture | -electrical issue -device-device incompatibility -medication error
EMC statement | Electric Fast Transients & Burst IEC 610004-4: The Mains quality should correspond to typical business or hospital environments. | -atrial tracking too fast -burst vessel -high flux rate | -electrical issue -fitting problem -defective item
EMC statement | Electromagnetic Immunity testing: Electrostatic discharge (ESD) IEC 61000-4-2; Quick electric transients/burst IEC 610004-4; Surges IEC 61000-4-5; Voltage drops, temporary power failures and fluctuations IEC 61000-4-11; Magnetic field at supply frequency (50/60 Hz) IEC 61000-4-8; HF interference currents IEC 61000-4-6; HF interference radiation IEC 61000-4-3 | -voltage dropped -power frequency out of specification -illumination power out of specification | -battery issue -electrical issue -power source issue
EMC statement | Voltage Dips, Interruptions & Voltage Variations IEC 61000-4-11: If the operator of the fabian HFO requires continued use in the event of a power failure, we recommend connecting the fabian HFO to an uninterruptible Mains supply or a Battery. | -voltage dropped -voltage out of range -induced voltage | -electrical issue -improper or incorrect procedure or method -no display or display failure
EMC statement | NEVER use portable or mobile radio devices closer to the fabian HFO, including cables, than the recommended safety distance calculated using the equation applicable to the transmission frequency. | -display distance out of specification -distance out of range -inappropriate safety margin | -incorrect display -material integrity issue -use of device issue
EMC statement | fabian HFO is intended to be operated in an electromagnetic environment with controlled HF interferences. The operator of the fabian HFO can help avoid electromagnetic interference by maintaining the minimum distance between portable and mobile HF telecommunication devices (transmitters) and the device fabian HFO – depending on the output rating of the communication device, as listed below. | -input originating from electromagnetic interference -interference with monitoring device -fail to maintain status | -no display or display failure -defective item -sticking

## System_Architecture
Index | category | Inputs | Failure_Modes | Hazard | Source
FAB-SAD-001 | system function | The Fabian HFO is intended for premature infants, new-born as well as children weighing up to 30 kg, “in-patient use” in professional healthcare environment including intra-hospital patient transport. |  |  | SAD
FAB-SAD-002 | system function | The Fabian HFO Ventilator is an innovative, compact, microprocessor-based device for accurate control and monitoring of respiration. The pneumatic circuits are based on a novel control concept using two fast micro controllers for monitoring and control. The user interface is controlled by a single board computer based on a RISC-Processor that also supervises the pneumatic controller circuits. For monitoring flow the Fabian HFO uses a small, low dead space plastic sensor that is adapted to the ET tube close to the patient. |  |  | SAD
FAB-SAD-003 | system function | Oxygen is metered by the integrated Air/O2 blender. The oxygen concentration is measured internally with a galvanic oxygen sensor. |  |  | SAD
FAB-SAD-004 | system overview | This section of the System Architecture Definition document includes design output as well as design specifications. Tags in this section do not denote requirements that are intended to be tested at the system level. |  |  | SAD
FAB-SAD-005 | Theory of Operation | The blender module consists of 1 proportional valve for air and 1 proportional valve for oxygen. Each valve can deliver a flow from 0 to the maximum. The highest flow through one valve is 32 LPM. The delivered flow of each valve is measured by a flow sensor and a local control loop maintains the correct value of flow. |  |  | SAD
FAB-SAD-006 | Theory of Operation | The air input pressure and oxygen input pressure are read by pressure transducers on the blender module board. |  |  | SAD
FAB-SAD-007 | Theory of Operation | The respiratory gas delivered to the patient is a mixture of medical air and O2. Based on the clinical need, the O2 concentration in the blend is programmable from 21% to 100%. The relationship between O2 and air concentrations are described in the equations below: |  |  | SAD
FAB-SAD-008 | Theory of Operation | The mechanism with which the Ventilator components achieve the target blend ratio is described in section 7.2 Pneumatics Architecture. |  |  | SAD
FAB-SAD-009 | Theory of Operation | The total outlet flow target is servo controlled to achieve a desired pattern of ventilation, which may either be based on a target pressure waveform, a target flow waveform, or dual control (breaths which switch between control target modalities). |  |  | SAD
FAB-SAD-010 | Theory of Operation | The calibration of the oxygen sensor is performed at two points. The sensor is calibrated at 21% (room air) and at 100% oxygen. It is also possible to activate the automatic calibration of only 21% or 21% and 100% oxygen. The automatic O2 Sensor calibration will be accomplished at start-up and every 24 hours. |  |  | SAD
FAB-SAD-011 | Theory of Operation | The Fabian HFO Ventilator is an innovative, compact, microprocessor-based device for accurate control and monitoring of respiration. The pneumatic circuits are based on a novel control concept using two fast micro controllers for monitoring and control. |  |  | SAD
FAB-SAD-012 | Theory of Operation | The user interface is controlled by a single board computer based on a RISC-Processor that also supervises the pneumatic controller circuits. For monitoring flow the Fabian HFO uses a small, low dead space plastic sensor that is adapted to the ET tube close to the patient. |  |  | SAD
FAB-SAD-013 | Theory of Operation | The output pressure of the blender delivered into the tube system is monitored. There are two security functions realized by this sensor. 1. Inspiratory pressure sensor – blender status 2. Inspiratory pressure sensor – airway pressure compare |  |  | SAD
FAB-SAD-014 | Theory of Operation | The airway pressure sensor monitors the pressure close to the patient at the y-connector. Its value is used to regulate pressure conditions through the exhalation valve by a closed loop regulation circuit. |  |  | SAD
FAB-SAD-015 | Theory of Operation | The total outlet flow target (i.e. Vout,target) is servo-controlled to achieve a desired pattern of ventilation, which may be based on the desired pressure waveform. The Ventilator is programmed to achieve a target pressure waveform by controlling the opening pressure of the Exhalation Valve. |  |  | SAD
FAB-SAD-016 | Theory of Operation | Pressure breaths have a controlled pressure waveform, which may be comprised of components that are preset by the operator (e.g. pressure rise time and plateau pressure) and empirically-derived values (e.g. leakage of air at the patient mask or nasal prongs in NIV mode). |  |  | SAD
FAB-SAD-017 | Theory of Operation | In the Fabian HFO Ventilator, pressure of delivered breaths are controlled by the Exhalation Valve. |  |  | SAD
FAB-SAD-018 | Ventilator | The Fabian HFO ventilator will deliver gas to the patient. The Fabian HFO ventilates with excess pressure based on the continuous-flow principle. (Time cycled, pressure/volume limited or volume guaranteed) Oxygen is metered by the integrated Air/O2 blender. The oxygen concentration is measured internally with a galvanic oxygen sensor. |  |  | SAD
FAB-SAD-019 | Ventilator | The ventilator incorporates connections to high pressure gas inlet; the patient circuit (inspiratory and expiratory ports); a shared connector port; and dedicated ports for the capnometry, (optional) pulse oximetry, and nebulizer accessory devices. |  |  | SAD
FAB-SAD-020 | Ventilator | The shared connector port provides isolated DC power and signal interface for: • Hot wire flow sensor • TPP module (Esophageal pressure measurement system) |  |  | SAD
FAB-SAD-021 | Ventilator | The capnometer port provides isolated DC power and signal interface for a mainstream or sides-tream capnometer module. |  |  | SAD
FAB-SAD-022 | Ventilator | The nebulizer port provides isolated drive signals for an electronic nebulizer. |  |  | SAD
FAB-SAD-023 | Ventilator | The ventilator incorporates an external alarm connector for a remote nurse call. |  |  | SAD
FAB-SAD-024 | Ventilator | The ventilator incorporates RS232 port for RS232 port for PDMS to ICU monitors. Connectivity via this port is defined in the Connectivity section of the PRS and includes VOXP protocol. |  |  | SAD
FAB-SAD-025 | Ventilator | The ventilator incorporates a USB port for data output, Software update and SpO2 Module (MASIMO) |  |  | SAD
FAB-SAD-026 | Ventilator | The ventilator incorporates a non-isolated HDMI output for connection to an external display monitor or a projector. This functionality is labeled as not for clinical use. |  |  | SAD
FAB-SAD-027 | Ventilator | The ventilator incorporates an internal rechargeable battery pack. |  |  | SAD
FAB-SAD-028 | Ventilator | The ventilator incorporates a variable speed fan to provide ambient air flow to inside the ventilator for cooling ventilator components. There will be a gross particulate filter in front of the fan. This filter can be accessed and cleaned or replaced by the end-user without the need to remove any ventilator covers. Note: the primary protection against oxygen buildup within the device enclosure will be dual seals on all oxygen-carrying parts. |  |  | SAD
FAB-SAD-029 | Ventilator | The ventilator mobile stand (cart) provides a secure, wheeled mounting for the ventilator and its accessories. The mobile stand incorporates a mechanical attachment point for the ventilator, a circuit support arm, a mechanical attachment point for the optional humidifier, a basket or tray for storing at least one patient circuit and patient interface set and a holder for two ‘E” size gas cylinders. |  |  | SAD
FAB-SAD-030 | Ventilator | The ventilator will have an AC power connection. The ventilator will support the use of an IEC 60320 terminated power cordset for connection to wall AC power. |  |  | SAD
FAB-SAD-031 | Ventilator | The ventilator will include a graphical user interface to allow access to all normal operating controls and functions. The graphical user interface will include a color TFT-LCD display assembly, touch screen panel, ambient light sensor. LCD screen provides real-time graphic visualization of pressure, flow, and volume on the LCD screen |  |  | SAD
FAB-SAD-032 | Ventilator | The LCD panel with touchscreen will have a size 26 cm (10.4”) will have a resolution of 800x600 pixels connect the LCD panel and backlight cables to the Mainboard. |  |  | SAD
FAB-SAD-033 | Ventilator | The LCD display incorporates a display backlight. The backlight brightness is programmable via the GUI. |  |  | SAD
FAB-SAD-034 | Ventilator | The ventilator incorporates two high-pressure gas inlets, for connecting to a high-pressure oxygen/air source and a high-pressure medical air both at nominal 2.0 – 6.5 bar. The connector adaptor plate incorporates magnets to provide a code that identifies the gas type. The magnetic code is read through Hall-effect switches. |  |  | SAD
FAB-SAD-035 | Ventilator | In Fabian HFO ventilator, the high-pressure gas inlets will each use a gas-specific connector in one of the following types: male DISS, female DISS, SIS, NIST to prevent connection to an incorrect gas pipeline. |  |  | SAD
FAB-SAD-036 | Ventilator | In Fabian HFO ventilator, the high-pressure gas inlets can identify, at least, the following gas types: • Oxygen • Air (note: Air is normal use, but identification is required to prevent foreseeable misuse) |  |  | SAD
FAB-SAD-037 | Ventilator | For use with high-pressure air, oxygen, a gas specific hose and inlet manifold will be required. This gas hose will be color coded with gas specific connectors at the supply end, to match the encoded ID. |  |  | SAD
FAB-SAD-038 | Ventilator | This gas hose will be available in various configurations applicable to country of sale. |  |  | SAD
FAB-SAD-039 | Ventilator | The ventilator will include a filter at the oxygen/air inlet to minimize the possibility of contaminants getting into the breath delivery system and causing failures of the flow control valve and flow sensors. |  |  | SAD
FAB-SAD-040 | Ventilator | The ventilator will include a filter at the air inlet to minimize the possibility of contaminants getting into the breath delivery system and causing failures of the flow control valve and flow sensors. |  |  | SAD
FAB-SAD-041 | Ventilator | The high-pressure gas inlet filters will be capable of replacement as a planned maintenance. This filter (not shown in Figure 1) resides inside the Flow Blender subsystem and the ventilator back cover needs to be removed in order to access this filter. NOTE: A BFE/VFE or HEPA filter is not required here since the gases are from a piped or bottled source |  |  | SAD
FAB-SAD-042 | Ventilator | The ventilator will include a coalescing water trap before the air inlet to remove the liquid water from the compressed air. |  |  | SAD
FAB-SAD-043 | Ventilator | The ventilator incorporates a HEPA filter and water trap at the return gas (expiratory) port of the ventilator. This water trap and exhalation filter will be a replaceable accessory; it can be accessed and replaced by the end-user without the need to remove any ventilator covers. |  |  | SAD
FAB-SAD-044 | Ventilator | The ventilator incorporates a heated exhalation manifold, downstream of the exhalation filter and water trap. This manifold is upstream of the exhalation flow sensor, to ensure that the flow sensor is not subjected to liquid water (condensate). |  |  | SAD
FAB-SAD-045 | Ventilator | The exhalation manifold, flow sensor and exposed parts of the exhalation valve are captured within thermal insulation. |  |  | SAD
FAB-SAD-046 | Ventilator | To maximize runtime with the battery, the exhalation manifold heater is operational only when the ventilator is running off AC mains power. |  |  | SAD
FAB-SAD-047 | Ventilator | The ventilator incorporates a solid state (non-depleting) oxygen sensor. The oxygen sensor will be used to monitor the fraction of oxygen in the gas mixture output from the ventilator. Note: the oxygen sensor is for monitoring only and is not used for servo control. |  |  | SAD
FAB-SAD-048 | Ventilator | The oxygen sensor will be used for both air/oxygen mixture. |  |  | SAD
FAB-SAD-049 | Ventilator | A single use external inspiratory bacterial filter will be installed between the inspiratory outlet of the ventilator and the humidifier inlet or circuit if no humidifier is used. |  |  | SAD
FAB-SAD-050 | Ventilator | The ventilator incorporates the use of a hot wire flow sensor for neo/infant patients for invasive ventilation. The ventilator will be designed to provide isolated DC power and isolated signal interface to enable this function to be added. The HWFS will be of the constant temperature (resistance) type where the temperature (and hence the resistance) of the heated wire is maintained at a constant value. The temperature of the heated wire changes initially as the flow rate varies, but the electrical power supplied by the ventilator to the heated wire will be automatically adjusted to compensate for the change in temperature and bring the temperature of the wire back to its original value. The change in the supplied electrical power to the heated wire is a measure of flow rate. |  |  | SAD
FAB-SAD-051 | Ventilator | The ventilator is required to connect to an external esophageal pressure module. The ventilator will be designed to provide DC power and serial data signals to enable this function to be added. |  |  | SAD
FAB-SAD-052 | Ventilator | The ventilator is required to connect to an internal OEM pulse oximeter module. The ventilator will be designed to provide isolated DC power and serial data signals to enable this function to be added. |  |  | SAD
FAB-SAD-053 | Ventilator | The ventilator may be normally used with a heater / humidifier; in this use scenario, the ventilator mobile stand will provide for secure mounting and transport of the humidifier together with the ventilator. The humidifier will interface mechanically to the patient breathing circuit. The humidifier will connect to AC power directly to the wall receptacle. |  |  | SAD
FAB-SAD-054 | Ventilator | The ventilator will be required to connect to an electronic nebulizer, for delivery of aerosolized medications during inspiration. The ventilator will be designed to provide the drive signals for the nebulizer. |  |  | SAD
FAB-SAD-055 | Ventilator | The ventilator will be required to connect to an optional external OEM capnometer module. The ventilator will be designed to provide DC power and serial data signals to enable this function to be added. |  |  | SAD
FAB-SAD-056 | Ventilator | The ventilator incorporates a proximal pressure sensor. The ventilator will be designed to provide a proximal pressure port for monitoring airway pressure at the wye for infant flow circuit only. |  |  | SAD
FAB-SAD-057 | Pneumatic Architectural View | The pneumatic architecture of Fabian HFO ventilator is shown in Figure 2. The Fabian HFO ventilator is an electromechanical device, with all pneumatic functions implemented within the ventilator module, or within the breathing circuit and associated accessories. |  |  | SAD
FAB-SAD-058 | Pneumatic Architectural View | The pneumatic architecture will comprise of three functional subsystems, the Gas Delivery, Safety, and Patient Monitoring. While all three of these subsystems interact to perform their intended functions, they provide distinctly different operations for care of the patient. |  |  | SAD
FAB-SAD-059 | Pneumatic Architectural View | The Gas Delivery subsystem will provide a high pressure gas delivery for delivery of compressed ambient air, a high pressure gas delivery for delivering oxygen , a gas mixing subsystem, and an exhalation subsystem. |  |  | SAD
FAB-SAD-060 | Pneumatic Architectural View | The primary component of the Safety subsystem will be 1) the exhalation valve that maintains and controls circuit pressure in conjunction with the gas delivery subsystem and provides an exit path for patient exhaled gas when the ventilator declares a Safety or Fail State condition, 2) the pressure discharge solenoid valve that acts as backup for the exhalation valve and opens under software control in order to relieve the circuit pressure 3) the sub-ambient valve that provides patient access to ambient air when the airway pressure falls below the ambient pressure, and 4) a passive pressure relief valve that opens when the outlet pressure exceeds a certain value in order to reduce the gas pressure applied to the patient. |  |  | SAD
FAB-SAD-061 | Pneumatic Architectural View | The Patient Monitoring subsystem will include sensors for measuring barometric pressure, inspiratory pressure (via inspiratory pressure sensor), proximal airway pressure (via proximal pressure sensor only for infants flow), delivered oxygen, and airway flow (only for neo/infants in invasive ventilation using a hot wire flow sensor). |  |  | SAD
FAB-SAD-062 | Pneumatic Architectural View | The Gas Delivery System will have the capability to provide the required flow rates and mixtures of air, oxygen to enable the ventilator to deliver a target pressure or target flow at the patient interface as required by control system based on the user settings. The Gas Delivery system consists of the two high pressure gas cassettes (Flow Cassette subsystem, one for oxygen and the other for air), inspiratory pressure sensor, and the exhalation valve. |  |  | SAD
FAB-SAD-063 | Pneumatic Architectural View | A Flow Cassette subsystem will control the gas flow from a high-pressure gas source and deliver a controlled flow based on a commanded value (30) sent every millisecond from the main breath control servo residing on the main PCBA (VCM). |  |  | SAD
FAB-SAD-064 | Pneumatic Architectural View | The high-pressure gas supply will be a piped medical oxygen supply (100% oxygen), a high-pressure gas cylinder of pure oxygen and a piped/compressed medical air. |  |  | SAD
FAB-SAD-065 | Pneumatic Architectural View | The Flow Cassette delivery system consists of a high-pressure gas inlet, gas mixture identification, a check valve, a filtration system, a gas temperature sensor, a pressure transducer, a flow sensor, and a flow control valve. |  |  | SAD
FAB-SAD-066 | Pneumatic Architectural View | The Flow Cassette subsystem incorporates an ISO 5359 high pressure gas inlet connector and is responsible for identification (1) of the gas mixture (see XLC-SAD-1080). |  |  | SAD
FAB-SAD-067 | Pneumatic Architectural View | The Flow Cassette is provided with coefficients for gas corrections by the VCM based on the identified gas mixture. |  |  | SAD
FAB-SAD-068 | Pneumatic Architectural View | A check valve will be used to prevent air or other gas contaminating the hospital gas pipeline or venting to atmosphere in single point failure (no pressure at inlet). An inlet filter will be used (see XLC-SAD-1086). |  |  | SAD
FAB-SAD-069 | Pneumatic Architectural View | A gas temperature sensor, placed in the gas path, will be used to allow appropriate compensation for the flow sensor. A pressure transducer will measure the inlet pressure of the gas provided to the Flow Cassette. The pressure will be used for gas corrections and as an input to the flow servo for closed-loop control of the flow control valve. |  |  | SAD
FAB-SAD-070 | Pneumatic Architectural View | The inlet pressure transducer will also be part of the safety system and be used to detect the loss of the high pressure gas source. Upon loss of high pressure gas the ventilator will annunciate an alarm to notify the clinician to take action, and also reverts to operation on the alternative source alone. |  |  | SAD
FAB-SAD-071 | Pneumatic Architectural View | An internal flow sensor will be provided to monitor, and provide control for, the flow of gas to the patient outlet. This flow sensor will consist of a flow restrictor, with a differential pressure transducer across the restrictor element. The pressure transducer incorporates a die (substrate) temperature monitoring capability to allow for appropriate compensation of flow and pressure values. A micro-controller on the Flow Cassette PCBA will use the die temperature, the raw pressure measurements, and the coefficients stored on an EEPROM on the pressure transducer to calculate corrected differential pressure and absolute pressure values, which will be then accurate over temperature and changing common mode pressure. The manufacturing calibration process of the flow sensor will allow for operation with oxygen, air. |  |  | SAD
FAB-SAD-072 | Pneumatic Architectural View | A flow control valve will be used to control the flow of gas to the device outlet. This will be capable of controlling gas flow over a turn-down ratio of greater than 60,000:1. The flow control valve is controlled by a set of algorithms that calculate the correct flow of oxygen based on the clinician set FiO2 and patient demand. The control system uses the flow sensor to measure flow and provide a feedback mechanism to ensure precise control of the oxygen delivered. When used with an oxygen gas source, the flow control valve regulates the amount of oxygen delivered to the gas mixing system based on the total patient flow required, and the set oxygen fraction (see Section 6.3). |  |  | SAD
FAB-SAD-073 | Pneumatic Architectural View | The inlet pressure transducer reading will be used to correct the gas flow from measured conditions to SLPM conditions. The internal flow servo loop will modulate the flow control valve and maintain accurate volumetric gas delivery. The delivered flow from the Flow Cassette (3, 8) SLPM, the gas temperature (2, 7), and the inlet pressure (4, 9) will be communicated every millisecond to the upper level system (the main processor on VCM board). These measurements plus other sensor inputs such as the barometric (14) pressure will be used on by the VCM processor to convert the measured flow in SLPM to corrected conditions (BTPS) for the patient. |  |  | SAD
FAB-SAD-074 | Pneumatic Architectural View | The outlet flow from the air Flow Cassette and the outlet flow from the oxygen Flow Cassette are connected to a manifold incorporating a sub-ambient valve, connections for a safety valve (comprised of a pressure discharge solenoid valve and a passive pressure relief valve), connections for an inspiratory pressure sensor (including a inspiratory autozero solenoid valve), connections for a proximal pressure sensor (including a proximal autozero solenoid valve), an oxygen sensor, and then to the ventilator outlet (inspiratory port). |  |  | SAD
FAB-SAD-075 | Pneumatic Architectural View | The flows from the air Flow Cassette and from the oxygen Flow Cassette flow are flow-metrically proportioned to deliver the set oxygen concentration and real time combined output flow value (see Section 6.3). |  |  | SAD
FAB-SAD-076 | Pneumatic Architectural View | The pressure relief and safety subsystem consists of a pressure discharge solenoid valve, a passive pressure relief valve, a sub-ambient valve, and the exhalation valve. |  |  | SAD
FAB-SAD-077 | Pneumatic Architectural View | The ventilator will have an exhalation valve at its expiratory port. The exhalation valve will be opened under software control to relieve circuit pressure. |  |  | SAD
FAB-SAD-078 | Pneumatic Architectural View | The exhalation valve will ‘fail-safe’ open in case of loss of electrical power, software control, or loss of all inlet gases, and whenever the ventilator is turned off. |  |  | SAD
FAB-SAD-079 | Pneumatic Architectural View | The ventilator will have a pressure discharge solenoid valve at the outlet that will be opened under software control to relieve the circuit pressure. The pressure discharge solenoid valve is intended as a backup to the exhalation valve and it will be activated in cases when opening the exhalation valve alone is not able or sufficient to relieve the circuit pressure. |  |  | SAD
FAB-SAD-080 | Pneumatic Architectural View | To prevent applying excessive gas pressures to the patient, the gas outlet incorporates a passive pressure relief valve that will open if the outlet pressure exceeds 100 cmH2O and will prevent the outlet pressure from exceeding 125 cmH2O. |  |  | SAD
FAB-SAD-081 | Pneumatic Architectural View | The gas outlet incorporates a passive sub-ambient valve that will open if the outlet pressure falls below the ambient pressure in order to allow ambient air to be available to patient. |  |  | SAD
FAB-SAD-082 | Pneumatic Architectural View | During a ventilator Fail-Safe state condition, the exhalation valve, the sub-ambient valve, and the mask sub-ambient relief port (if applicable) will work in conjunction to: • Relieve pressure from the circuit down to ambient pressure conditions • Allow ambient air to be available to the patient for breathing • Minimize re-breathing of gases (function of patient circuit and application) |  |  | SAD
FAB-SAD-083 | Pneumatic Architectural View | The exhalation subsystem consists of a heated manifold incorporating an exhalation valve and an X-Flow Sensor, and an unheated filter and water trap. The filter and water trap are as described in XLC-SAD-1092. |  |  | SAD
FAB-SAD-084 | Pneumatic Architectural View | The exhalation valve is a software-controlled valve that is used to adjust the flow of gas passed through the expiratory port of the ventilator to outside environment. The exhalation valve is controlled by a software control signal (34) and works in conjunction with the gas delivery subsystems to maintain user set pressure control levels. In exhalation, in CPAP therapy, in TCPL ventilation, and in APRV the exhalation valve is maintained at a set pressure, and outlet flow is servo controlled to a specified target bias flow rate. Also, additional (demand) flow is provided to maintain the pressure in the event of patient inspiratory flow exceeding the bias flow. |  |  | SAD
FAB-SAD-085 | Pneumatic Architectural View | The exhalation valve operates on the principle of a force balance across the control diaphragm, which is a disposable valve membrane. A linear magneto-mechanical actuator controls the force on the diaphragm, which in turn controls the circuit pressure. The force generated by the actuator is based on a command (34) from the software closed-loop controller. The exhalation valve poppet position is measured and provided as an output (10). |  |  | SAD
FAB-SAD-086 | Pneumatic Architectural View | When a proximal flow sensor is not in use, airway flow is estimated from measurement of flow in the exhalation subsystem, circuit compliance volume change, and the total outlet flow (minus the known flow to the oxygen sensor). Gas flows in the exhalation subsystem are measured under conditions which vary depending on the breathing circuit configuration. With an active humidifier, gas reaching the exhalation flow sensor remains saturated; the BTPS correction is only required for gas temperature, which is measured within the X-Flow Sensor. |  |  | SAD
FAB-SAD-087 | Pneumatic Architectural View | The exhalation valve diaphragm will be capable of replacement as a planned maintenance. This exhalation valve diaphragm can be accessed after removing the covers. |  |  | SAD
FAB-SAD-088 | Pneumatic Architectural View | The exhalation X-Flow Sensor is primarily used to determine the patient flow, which is calculated by subtraction of exhalation flow, leak flow, and circuit compliance volume change from the outlet flow. |  |  | SAD
FAB-SAD-089 | Pneumatic Architectural View | The Exhalation X-Flow Sensor measures the gas temperature, the pressure, and the flow at the exhalation valve inlet. Effects of back pressure between the exhalation inlet and the exhalation X-Flow Sensor will be considered in the measurement. |  |  | SAD
FAB-SAD-090 | Pneumatic Architectural View | The Exhalation X-Flow Sensor is provided with coefficients for gas corrections by the VCM, based on the current output mixture of air and high pressure inlet gas and the gas type for the high pressure inlet (air and Oxygen mixture). |  |  | SAD
FAB-SAD-091 | Pneumatic Architectural View | The STPD21 corrected flow (14), the gas temperature (15), and the exhalation pressure (13) will be reported from the X-Flow Sensor, as requested by the upper level system (the main processor on VCM board). |  |  | SAD
FAB-SAD-092 | Pneumatic Architectural View | A heater is provided downstream of the water trap and filter, and upstream of the exhalation flow sensor. The heater power (35) is controlled to warm the exhalation gas in order to prevent condensation within the flow sensor or exhalation valve. |  |  | SAD
FAB-SAD-093 | Pneumatic Architectural View | The Patient Monitoring system will include sensors for measuring barometric pressure, machine outlet and exhalation port pressure, exhalation port flow, airway flow (with infant hot-wire flow sensor only), and the delivered oxygen percent. |  |  | SAD
FAB-SAD-094 | Pneumatic Architectural View | A barometric pressure sensor (12) will provide measurements to the main breath control system or any other device components requiring this data. Both pressure measurements are used to correct gases to the required patient conditions. |  |  | SAD
FAB-SAD-095 | Pneumatic Architectural View | The proximal and Machine pressure sensor (13) is integral to the operation of the exhalation valve and the safety systems. |  |  | SAD
FAB-SAD-096 | Pneumatic Architectural View | A hot wire proximal flow sensor (17) is used only for neo/infant patients for invasive ventilation. The hot wire flow sensor (HWFS) resides at the interface to the patient and measures flow bi-directionally. |  |  | SAD
FAB-SAD-097 | Pneumatic Architectural View | Inspiratory pressure transducer in the inspiratory manifold and the exhalation X-Flow Sensor provide for redundancy of pressure monitoring, and detection of circuit occlusion faults. |  |  | SAD
FAB-SAD-098 | Pneumatic Architectural View | The oxygen sensor will be used to monitor the fraction of oxygen (11) in the gas mixture output from the ventilator and it is used for both air/oxygen mixture. |  |  | SAD
FAB-SAD-099 | Pneumatic Architectural View | Measured oxygen values should be processed in software to correct for pressure and for vibration. |  |  | SAD
FAB-SAD-100 | Pneumatic Architectural View | In order to maintain the performance of the oxygen sensor within specification, this paramagnetic oxygen sensor requires calibration using the known supply gas and air. At appropriate intervals, the ventilator software will ask the operator to run an extended self test (EST), which will include the calibration of the oxygen sensor. The ‘calibration’ defined here is the process of obtaining the zero (21% O2) and span (100% O2) values of the oxygen sensor and storing these values in the ventilator for use in normal operation. |  |  | SAD
FAB-SAD-101 | Pneumatic Architectural View | The patient circuit forms an integral part of the pneumatic architecture, incorporating components for humidification, pressure relief, and pressure and flow measurement of gas to and from the patient. |  |  | SAD
FAB-SAD-102 | Pneumatic Architectural View | The breathing circuit consists of the gas-wetted components of an optional dry gas conduit, an optional humidifier chamber, a gas conduit to convey breathing gas from the ventilator outlet to the patient airway interface, a wye connector (in the case of dual-limb circuits), a patient interface (invasive or non-invasive), and a gas conduit to return the exhaled gas from the patient to the ventilator (in the case of dual-limb circuits). |  |  | SAD
FAB-SAD-103 | Pneumatic Architectural View | When used with an active humidification system, the breathing circuit incorporates a dry gas conduit (to convey breathing gas from the ventilator outlet to the humidifier), a humidifier chamber (to attach to the active humidifier for addition of heat and water vapor, the inspiratory gas conduit incorporating a heater wire, and a remote temperature probe port included in the inspiratory tubing no more than 18” from the patient interface. |  |  | SAD
FAB-SAD-104 | Pneumatic Architectural View | Passive humidification (HME) can be used instead of active humidification. When used with a Heat and Moisture Exchanger (HME) the HME will fit between the wye and the tracheal tube connector (ETT or tracheostomy tube). |  |  | SAD
FAB-SAD-105 | Pneumatic Architectural View | The breathing circuit incorporates a 22mm ISO 5356-1 taper connection to the patient mask. Some of the typical breathing circuit configurations for invasive and non-invasive ventilations are shown in the following sections. |  |  | SAD
FAB-SAD-106 | Electrical Architecture View | The Main Board contains the control software for the ventilation procedure and provides connectivity to the PCBAs that are dedicated to major functionalities required for the ventilation procedure. It determines the necessary controls for a programmed ventilation procedure. |  |  | SAD
FAB-SAD-107 | Electrical Architecture View | The Main Board is responsible for distributing electrical power to other system components, integrating electrical signals (e.g. transducer signal, user input, etc.) from them, and providing control signal for the ventilation procedure. |  |  | SAD
FAB-SAD-108 | Electrical Architecture View | The Main Board incorporates connectivity to NetDCU. |  |  | SAD
FAB-SAD-109 | Electrical Architecture View | The Fabian HFO Ventilator utilizes both analog and digital components and circuitry. The Main Board incorporates a watchdog circuit for monitoring the programmable electrical components. In case of a failure, the watchdog opens the pressure relief valves and stops power to the Blender module to prevent and/or relieve pressure in the Patient Circuit. The watchdog circuit monitors for failures of the following components, these are also integral parts of the main board: • Monitor PIC • Control PIC • NetDCU • Alarm PIC |  |  | SAD
FAB-SAD-110 | Electrical Architecture View | The monitoring microprocessor communicates with the User Interface Microprocessor using a SPI connection. It receives the ventilator settings from the UIM and transfers all graphic data (proximal pressure, flow, volume) values and the monitored ventilator state back to the UIM. |  |  | SAD
FAB-SAD-111 | Electrical Architecture View | Diversity and redundant monitoring of the most important parameters is achieved by using a different data path and measurements independent of the control microprocessor. Measurements:  air input pressure  oxygen input pressure  blender output pressure  patient proximal pressure  flow at the y-Piece  oxygen sensor 1  oxygen sensor 2  supply voltage |  |  | SAD
FAB-SAD-112 | Electrical Architecture View | The Monitor PIC incorporates a SPI for communication with the NetDCU. The Monitor PIC incorporates an interface for receiving tachometer signal from the chassis fan. The Monitor PIC incorporates an analog interface for connection to the O2 Sensor. The Monitor PIC incorporates an analog interface for connection to the HWFS Connector. |  |  | SAD
FAB-SAD-113 | Electrical Architecture View | The Monitor PIC incorporates an analog interface to 2 x calibration potentiometers, which is located on the Manifold Board, for the Inspiratory Pressure Transducer. Each potentiometer is used for calibration of the transducer. |  |  | SAD
FAB-SAD-114 | Electrical Architecture View | The Monitor PIC incorporates an I2C interface for connection to a FRAM. The FRAM stores calibration data, which persists through power cycles, for the following components:  Inspiratory Pressure Transducer  Proximal Pressure Transducer  Blender Pressure Regulators  Exhalation valve  HWFS |  |  | SAD
FAB-SAD-115 | Electrical Architecture View | The Monitor PIC incorporates Tacho connection with the fan module. It regulates the speed of the fan to prohibit the high heat inside the ventilator body. It also prevent ignition due to high concentration of oxygen inside the device. |  |  | SAD
FAB-SAD-116 | Electrical Architecture View | The control microprocessor PIC2 receives the ventilator settings via RS232 and controls the blender, bleed valve, HFO Module and the exhalation valve. Measurements:  air input pressure  oxygen input pressure  blender output pressure  patient proximal pressure |  |  | SAD
FAB-SAD-117 | Electrical Architecture View | The Control PIC incorporates a PWM input to the Exhalation Valve via a valve driver connection, which is integrated into the Manifold Board. |  |  | SAD
FAB-SAD-118 | Electrical Architecture View | The Control PIC incorporates an analog interface for connection to the Proximal Pressure Transducer. The Control PIC incorporates an analog interface for connection to the Inspiratory Pressure Transducer. |  |  | SAD
FAB-SAD-119 | Electrical Architecture View | The Control PIC incorporates an analog interface for connection to the Inspiratory Solenoid Valve via a valve driver connection, which is integrated into the Manifold Board. |  |  | SAD
FAB-SAD-120 | Electrical Architecture View | The Control PIC incorporates a digital IO output for control of the Electrical Safety Valve. The Control PIC has connection with HFO Module via SPI synchronous serial communication interface. |  |  | SAD
FAB-SAD-121 | Electrical Architecture View | The Control PIC incorporates an analog communication interface with a Pressure Transducer which again integrated with Blender Board. The Control PIC incorporates a hardware communication protocol UART with NetDCU to control different peripherals. |  |  | SAD
FAB-SAD-122 | Electrical Architecture View | The Net DCU is a master processor single board computer based on a RISC-Processor (CPU: Intel PXA270, Architecture ARM9) produced by F&S, Stuttgart, Germany, for industrial use in the visualization and controlling of complex systems. |  |  | SAD
FAB-SAD-123 | Electrical Architecture View | The major functional features of the NetDCU used in the design of the Fabian HFO are:  Windows CE 6.0  CPU: Intel PXA270, Architecture ARM9  Frequency: 520  Video optional SM502 or CPU, Video in optional  Flash 64MB, RAM 64MB  Ethernet 10/100M o 3 x RS232, USB, I2C, SPI, Parallel Interface, Digital IO, Matrix Keyboard  Max.Resolution : SVGA ( optional SXGA) |  |  | SAD
FAB-SAD-124 | Electrical Architecture View | NetDCU is directly mounted on the Main Board, it is an SOM. NetDCU incorporates a digital IO input connection for communication to the UDP. NetDCU incorporates a digital IO input connection for communication to the Hardware Keys. NetDCU incorporates a USB connection for communication to an external USB device (i.e. nebulizer or for SW update). NetDCU incorporates a SPI for communication with the Monitor PIC. NetDCU provides PWM signal for control of the chassis fan. NetDCU incorporates a UART for communication to the EtCO2 Sensor Connector. |  |  | SAD
FAB-SAD-125 | Electrical Architecture View | NetDCU incorporates a serial connection for communication to the SpO2 Sensor Connector via a USB-to-Serial Converter, which is integrated into the Sensor Connector Board. NetDCU incorporates a HDMI-Out connection NetDCU incorporates a RS-232 interface for the PDMS. NetDCU incorporates a UART for communication with the Control PIC. NetDCU incorporates a DIO interface for communication with the Alarm PIC. NetDCU incorporates an I2C input to the Power Supply Board. NetDCU incorporates a PWM input to the Display Backlight. |  |  | SAD
FAB-SAD-126 | Electrical Architecture View | NetDCU incorporates an integrated touch controller for analog touch signal from the touchscreen. NetDCU incorporates a RGB interface to provide drive signal to the LCD display. NetDCU incorporates connectivity to an exchangeable SD memory card, which provides storage for trend & log files and special font characters. The log and trend files may be extracted via a WinCE-compatible USB device. NetDCU incorporates a real-time clock. |  |  | SAD
FAB-SAD-127 | Electrical Architecture View | The Alarm PIC incorporates a I2C interface for connection to the Speaker Driver, which connects to a downstream Loudspeaker via an analog interface. |  |  | SAD
FAB-SAD-128 | Electrical Architecture View | The Alarm PIC incorporates a digital IO output for control of the Alarm LED. The Alarm PIC incorporates a DIO interface for communication with NetDCU. |  |  | SAD
FAB-SAD-129 | Electrical Architecture View | Monitor PIC updates every 100 ms the status of flow sensor, oxygen sensor and gas supply pressure. The Alarm is triggered by the master processor. |  |  | SAD
FAB-SAD-130 | Electrical Architecture View | In case of a Sensor defect or problem, an alarm is triggered and the calibration screen opens automatically. A recalibration can be done, if it fails the sensor could be changed or deactivated. |  |  | SAD
FAB-SAD-131 | Electrical Architecture View | With a deactivated flow sensor, only ventilation modes without trigger function can be used. |  |  | SAD
FAB-SAD-132 | Electrical Architecture View | In case of a valve or flow sensor defect in the blender unit, a wrong gas mixture will be detected by the oxygen sensor. A defect or not right calibrated oxygen sensor will also be detected. In both cases an alarm is triggered and the calibration screen opens automatically. |  |  | SAD
FAB-SAD-133 | Electrical Architecture View | The gas supply (air and oxygen) is supervised by pressure sensors. If the pressure is lower than 0.8bar, an alarm is triggered by the pressure sensors. |  |  | SAD
FAB-SAD-134 | Electrical Architecture View | In case of a defect pressure sensor the loss of one gas supply is secondary detected by the O2 Sensor because of the differ between set and monitored gas mixture. |  |  | SAD
FAB-SAD-135 | Electrical Architecture View | The loss of both gas supplies is secondary detected by the alarm settings because of a decreasing pressure and flow |  |  | SAD
FAB-SAD-136 | Electrical Architecture View | When mains power fails when the device is switched on, the battery delivers sufficient energy for an acoustical and optical power fail alarm for a minimum time of 10 minutes. |  |  | SAD
FAB-SAD-137 | Electrical Architecture View | The fan is an important component to prevent ignition because of high oxygen concentration in the device, in case of a defect in the pneumatic system. It is supervised by the Monitor PIC by a speed signal with an upper limit (speed > 190Hz) and a lower limit (speed < 80Hz). The signal is measured every 100ms, the filter for alarm triggering is 5seconds. |  |  | SAD
FAB-SAD-138 | Electrical Architecture View | The Sensor Connection Board incorporates an isolated pass-through communications interface from the below-listed accessory connections to the Monitor PIC/NetDCU. |  |  | SAD
FAB-SAD-139 | Electrical Architecture View | The Sensor Connection Boards provides electrical circuitry and isolated DC power (5 V and 3.33 V) for the following accessory connections: SpO2 module- |  |  | SAD
FAB-SAD-140 | Electrical Architecture View | To enable the SpO2 measurement, the module must be enabled. This could be done in the menu of the SpO2 module. For the SpO2 module, the Sensor Connection Board incorporates a USB-to-Serial converter which receives USB communication from NetDCU and outputs UART for interface to the SpO2 Sensor Connector. |  |  | SAD
FAB-SAD-141 | Electrical Architecture View | The Fabian HFO supports two different CO2 module types to provide CO2 monitoring: 1. Microstream, MicroPod™ External etCO2 Module from Oridion® (Covidien/Medtronic) 2. LoFlo Sidestream etCO2 Sensor from Respironics (Philips). |  |  | SAD
FAB-SAD-142 | Electrical Architecture View | The main board contains the analog circuits of the hot-wire sensor interface; the 12-bit analog-to- digital converter of the monitor PIC is used. |  |  | SAD
FAB-SAD-143 | Electrical Architecture View | The Fabian HFO flow sensor is based on a hot wire anemometer with two platinum wires to determine flow and direction. |  |  | SAD
FAB-SAD-144 | Electrical Architecture View | Each hot wire is arranged in its own Wheatstone bridge circuit. The bridge supplies a sensor heating current - that varies with the flow velocity - and maintains constant sensor resistance and, thus, constant sensor temperature. |  |  | SAD
FAB-SAD-145 | Electrical Architecture View | Flow applied to the hot wire cools the sensor and unbalances the bridge. As a result, the electronic circuits increase the bridge voltage until the nominal resistance (and temperature) is reached again. The controlling microprocessor can switch off the bridge supply when overload is detected. |  |  | SAD
FAB-SAD-146 | Electrical Architecture View | The Blender Board provides the electrical circuitry required for the control of the Air and O2 Proportional Valves. |  |  | SAD
FAB-SAD-147 | Electrical Architecture View | For air and oxygen, 1 electronically controlled valve (5) each is used in the blender (4) to allow for sufficiently exact flow values. The valves are proportional type and can be controlled from zero to the maximum flow. So, the exact oxygen concentration can be achieved. Each valve is driven by a switch mode regulator. |  |  | SAD
FAB-SAD-148 | Electrical Architecture View | The actual flow of each path is being measured by a flow sensor similar to the patient flow sensor. A local control loop regulates the actual flow to the desired level. |  |  | SAD
FAB-SAD-149 | Electrical Architecture View | There are two more solenoid valves on the blender module: For calibration of the oxygen sensor, a pneumatic bypass can be activated by opening either the air- or the oxygen calibration valve. A step up converter delivers 24 V to the calibration valves. |  |  | SAD
FAB-SAD-150 | Electrical Architecture View | The Blender Board also incorporates the Blender PIC, which is described below. The PIC microprocessor sets the two blender valves for delivery of the desired oxygen concentration using a SPI interface to the blender module with the signals MCS, MDATA, and MCLK. |  |  | SAD
FAB-SAD-151 | Electrical Architecture View | The Blender PIC incorporates a SPI interface for connection to the Control PIC within the Main Board. The Blender PIC incorporates an analog interface for connection to the Air Flow Sensor The Blender PIC incorporates a PWM input for the control of the O2 Proportional Valve. The Blender PIC incorporates a PWM input for the control of the Air Proportional Valve. The Blender PIC incorporates an analog interface for connection to the O2 Flow Sensor. The Blender PIC incorporates a PWM input for the control of the O2 Calibration Valve. The Blender PIC incorporates a PWM input for the control of the Air Calibration Valve. |  |  | SAD
FAB-SAD-152 | Electrical Architecture View | The Blender PIC incorporates a I2C interface for connection to a FRAM. The FRAM stores persistent calibration data. The Blender Board incorporates 2 x calibration potentiometers for the calibration of the Air Proportional Valve & Flow Sensor (21%) and the O2 Proportional Valve & Flow Sensor (100%). |  |  | SAD
FAB-SAD-153 | Electrical Architecture View | The Manifold Board incorporates signals from the Proximal Pressure Transducer and the Inspiratory Pressure Transducer and relays the signals to the Control PIC within the Main Board. The Manifold incorporates valve drivers for the Inspiratory Solenoid Valve and the Exhalation Valve. The Manifold Board incorporates 2 x calibration potentiometers for the calibration of the Inspiratory Pressure Transducer. The Manifold incorporates an outlet switching relay connected with HFO module. |  |  | SAD
FAB-SAD-154 | Electrical Architecture View | The device can be operated from an external power source but features an internal battery (NiMH Power pack) also. Power module is the sum of the following components: - Power PCB - Powerpack NiMH/ Inner battery - Power supply “XP” - Power inlet module The Power Board comprises of Power PIC, which is described below. |  |  | SAD
FAB-SAD-155 | Electrical Architecture View | The Power PIC incorporates circuitry for power selection, DC / DC conversion, battery charge control, and all power supply-related monitoring and control mechanisms. The Power PIC incorporates a digital IO input from the On/Off hardware key. The Power PIC defaults to AC input power when available. When it is not available, the Power PIC can switch to the battery power without user-action and without disruption to the ventilation procedure The Power PIC incorporates a digital IO output to the Main LED within the User Control Interface. The Power PIC incorporates a digital IO output to the Battery Status LED within the User Control Interface. |  |  | SAD
FAB-SAD-156 | Electrical Architecture View | The Ventilator incorporates an internal battery (NiMH Power pack) that provides DC power to the Power Board (16.8 VDC nominal). The battery connects to the Power Board via an AKKU cable. |  |  | SAD
FAB-SAD-157 | Electrical Architecture View | The AC/DC Power Supply transforms the input from an external AC power supply (100~240 VAC) to 24 VDC. The Power Supply Board provides DC power to the Power Board (24 VDC). |  |  | SAD
FAB-SAD-158 | Electrical Architecture View | The Potential Equalization Pin is for additional safety and can be connected to an equi-potential zone. Adhere to local guidelines when using this PIN. The guidelines may vary between countries, localities and power companies. Always keep the PIN for Potential equalization accessible. |  |  | SAD
FAB-SAD-159 | Electrical Architecture View | The HFO Module’s pneumatic consists of an oscillator and an HFO open valve. An electromagnetic drive together with a flexible membrane forms the oscillator. It delivers the sinus oscillation and is via the HFO open valve connected to the patient circuit. A CLASS-D amplifier is used to drive the oscillator. From the electrical point of view HFO module is connected with main board and outlet switching relay of the manifold. Controller PIC controls the HFO modules to generate required high frequency oscillation. Controller PIC Data is sent via a local SPI to the HFO-Module. Data include: HFO Frequency, I / E-ratio, Amplitude The measured amplitude is compared with the setting value and a correcting value is delivered to the HFO-PIC. In this way it makes a control loop. |  |  | SAD
FAB-SAD-160 | Electrical Architecture View | Video In (VGA Port) cable to be connected to the Mainboard. The video option is not enabled yet. The digitalization of analog VideoIn signals results from Philips plug-in SAA7114. The SAA7114 is a video capture device for applications at the image port of Video Graphics Array (VGA) controllers. |  |  | SAD
FAB-SAD-161 | Electrical Architecture View | The SAA7114 is a combination of a two-channel analog preprocessing circuit including source selection, anti-aliasing filter and Analog-to-Digital Converter (ADC), an automatic clamp and gain control, a Clock Generation Circuit (CGC), a digital multi-standard decoder containing two-dimensional chrominance/luminance separation by an adaptive comb filter and a high performance scaler, including variable horizontal and vertical up and downscaling and a brightness, contrast and saturation control circuit. |  |  | SAD
FAB-SAD-162 | Electrical Architecture View | The fan is an important component to prevent ignition because of high oxygen concentration in the device, in case of a defect in the pneumatic system. |  |  | SAD
FAB-SAD-163 | Electrical Architecture View | It is supervised by the Monitor PIC by a speed signal with an upper limit (speed > 190Hz) and a lower limit (speed < 80Hz). The signal is measured every 100ms, the filter for alarm triggering is 5seconds. |  |  | SAD
FAB-SAD-164 | Electrical Architecture View | An external supply with a voltage of 12 – 24 volts must be connected to the back of the device in this case. VDC Connection to EXT +/-, which connects with the Power PCB. This is the connection for 24V DC external power supply. It is not for battery charging, for battery charging there is a separate connection form main board. It is for the ventilator operation. |  |  | SAD
FAB-SAD-165 | Electrical Architecture View | An independent thread toggles through via I2C line every 40ms. If due to an error the pneumatic control software spends excessive time in an outside routine or collapses completely, these pulses will fail, and therefore a system failure alarm will be triggered by hardware. |  |  | SAD
FAB-SAD-166 | Electrical Architecture View | The following signals are monitored by hardware: • WD0, WD1, WD2, WD2 : square wave signals from the micro controllers a upper and lower frequency limit ensures proper function Tolerances: approx. ± 20% • WD 0: Frequency: minimum 10 Hz, maximum 450 Hz, T_abfall = 5 s • WD 1,2,3: Frequency: minimum 50 Hz, maximum 450 Hz, T_abfall = 1 |  |  | SAD
FAB-SAD-167 | Electrical Architecture View | These signals are monitored by a dedicated hardware circuit. If one or more fail, the following actions take place:  System alarm  Drop of safety relay  Opening of the bleed valve  Shut down of the blender |  |  | SAD
FAB-SAD-168 | Electrical Architecture View | The sound generator uses a magnetic speaker and a full bridge amplifier. It generates the sound frequency and sets the adjustable amplifier to the desired level. It also activates the system fail-LED in case of a system fail caused by the Hardware watchdog circuit. The different sound sequences get activated from the IU via parallel bus and. |  |  | SAD
FAB-SAD-169 | Electrical Architecture View | It is expected that a Class I rated supply will be used, thus PE is required to allow for 2 MOP on primary. The ventilator system shall be Class I – internally powered. Any PE conductive parts meet Class I PE requirements. Electric isolation circuit |  |  | SAD
FAB-SAD-170 | Electrical Architecture View | One means of patient protection between applied parts as per Clause 8.5.2.1 for a working voltage equal to the maximum mains voltage. |  |  | SAD
FAB-SAD-171 | Electrical Architecture View | Multiplication factor applied to CL distances. 10,000ft --> 3,048m --> 1.29x for MOOP, 1.14x for MOPP. |  |  | SAD
FAB-SAD-172 | Electrical Architecture View | Where a working voltage for a circuit is below 24V, 24V is used in the analysis (conservative). |  |  | SAD
FAB-SAD-173 | Electrical Architecture View | IFU contraindicates the use of SIP/SOPs while the ventilator is not protectively earthed. |  |  | SAD
FAB-SAD-174 | Electrical Architecture View | All applied parts, except SpO2, meet this requirement by means of solid insulation and physical separation of the patient from electrical circuitry inherent to the way the applied parts are used. SpO2 meets this requirement by means of separation of electrical circuitry to the chassis within the system. |  |  | SAD
FAB-SAD-175 | Mechanical Packaging Architectural View | The ventilator is packaged in a mechanical enclosure to meet the environmental requirements specified in the Fabian HFO PRS. The Ventilator provides capability for mechanical interface to the Fabian mobile cart or for shelf use. |  |  | SAD
FAB-SAD-176 | Mechanical Packaging Architectural View | The Ventilator incorporates a sheet metal chassis that provides a mechanical interface for all components of the ventilator assembly, including the PCB assemblies, electronic components, and the pneumatics components. Assemblies/components that are mounted on the metal chassis are oriented in a way that optimizes serviceability of the Ventilator. |  |  | SAD
FAB-SAD-177 | Mechanical Packaging Architectural View | Following PCBAs/components are accessible on removing the housing as follows: Top cover: Mainboard, NetDCU Right cover (looking the Fabian from behind): Proportional gas blender with O2 sensor (O2 sensor cell, Proportional valves, Calibration valves, Input pressure regulators, Blender PCB, Inner flow sensor), Flow sensor connector PCB Left cover (looking the Fabian from behind): Power PCB, Powerpack/battery pack Bottom cover: HFO module Rear cover: Loudspeaker, Fan, Power inlet module, Power Supply ‘XP’, Video connector Board |  |  | SAD
FAB-SAD-178 | Mechanical Packaging Architectural View | The housing allows access to expiratory membrane, proximal connector, user knob and main fuses, for replacement without the need to remove housing. |  |  | SAD
FAB-SAD-179 | Mechanical Packaging Architectural View | The chassis also provides mechanical interface for the integrated user control interface components. The user control interface is designed to be installed with a vertical tilt for ease of viewing. The tilt angle is not adjustable. The external housing of the Ventilator incorporates plastic panels that augment the protection of internal components from environmental exposure. An opening on the rear side of the Ventilator allows for ambient air to flow into the fan and subsequently into the Ventilator. A separate exhaust opening (on the bottom side) also allows for the relief of circulating air from the fan and for the gas passing through the exhalation valve into the atmosphere. |  |  | SAD
FAB-SAD-180 | Mechanical Packaging Architectural View | The external packaging includes hand-hold area(s) to help with lifting the ventilator. |  |  | SAD
FAB-SAD-181 | Mechanical Packaging Architectural View | The front side of the housing incorporates the following user interfaces: • Manifold connectors o Inspiratory limb port/HFO port o Expiratory limb port o Proximal pressure port o External bias flow port (Fresh Gas) and port for nCPAP system based on flow generators (single limb systems) |  |  | SAD
FAB-SAD-182 | Mechanical Packaging Architectural View | • User control interface o Graphical user Interface  Touchscreen  LED Indicators o Key/Control Panel  Hardware Keys • Home • Graphics • Manual Breath • Alarm Limits • Alarm Silence • Menu/Calibration • Nebulizer • O2 flush • Start/Stop • ON/OFF  Rotary Push Knob |  |  | SAD

## Hardware_Sub
Index | category | Inputs | Failure_Modes | Hazard | Source
FAB-MHFO-001 | Interface Connectors Requirements | The HFO pneumatic subsystem shall have HFO outlet port connected to Inspiratory tube port at manifold through HFO flexible hose, specification of HFO flexible hose mentioned below. HFO outlet port shall connect patient-connection port close to the patient for connection to an airway device. HFO flexible hose: Adult breathing hose with sleeves for connectors with ISO 22 mm on both ends. Length of breathing hose shall be determined based on placement of connecting ports. Biocompatibility requirement: ISO 18562-1: 2017, ISO 15001:2010. |  |  | Hardware_pneumatic
FAB-MHFO-002 | Interface Connectors Requirements | Specification of HFO outlet port mentioned below. HFO outlet port: Barb connector compatible with 22 mm hose HFO flexible hose. HFO outlet port shall connect patient-connection port close to the patient for connection to an airway device. Material: Stainless steel |  |  | Hardware_pneumatic
FAB-MHFO-003 | Operational requirement | The HFO pneumatic subsystem shall include a mechanism to shut off/on pneumatic output at HFO outlet port that consists of HFO valve placed between HFO generator and HFO output port. |  |  | Hardware_pneumatic
FAB-MHFO-004 | Operational requirement | The HFO pneumatic subsystem shall provide interface between HFO PCBA and HFO Valve. HFO Valve will be controlled as CLOSED (not actuated) or OPEN (actuated) by HFO PCBA and shall have the following control specifications. Interface type: PWM drive through GPIO Notes: For details please refer HFO PCBA subsystem requirement documents. |  |  | Hardware_pneumatic
FAB-MHFO-005 | Operational requirement | HFO valve shall be electrically operated by 2/2 way actuation valve, normally closed (shutoff). HFO valve shall have below specifications. (Requires detailed data sheet to confirm below specifications) • Insulation class of the winding: Class F • Reference Temperature: 0 to 50 deg C • Operating mode: S1 (100 %) • Rated Voltage: 12 V(DC), • Rated current: 0.35 amp • Holding current: 0.08 amp • Peak current: 0.39 amp • Power: 4,12 Watt Notes: For more details please refer HFO PCBA subsystem requirement documents. |  |  | Hardware_pneumatic
FAB-MHFO-006 | Operational requirement | The HFO pneumatic subsystem shall have interface between HFO PCBA and High frequency oscillation generator by electrical input parameters through Class D amplifier for generating the High Frequency (HF) oscillations. High Frequency (HF) oscillations controls different output parameters as mentioned below. • Frequency 5 to 20 Hz with resolution 1 Hz with accuracy of ± (0.5 Hz). • 1:1 to 1:3 I:E ratio with resolution 1 (i.e. range is 1:1, 1:2, 1:3) |  |  | Hardware_pneumatic
FAB-MHFO-007 | Operational requirement | The HFO module includes an electromagnetic solenoid and horns attached with diaphragm which converts the electrical signal come from the PCB, converting it to a pneumatic output mentioned below. • 0 to 120 mbar peak to peak pressure with resolution <= 1 mbar and accuracy ± (0.5 mbar+ 4%). • Displacement volume per oscillation 0 to TBD millilitres. Notes: For details of controlling parameters please refer HFO PCBA subsystem requirement documents. |  |  | Hardware_pneumatic
FAB-MHFO-008 | Operational requirement | HFO pneumatic subsystem with HFO PCBA shall provide High frequency oscillation in presence of mean airway pressure in range 0 to 50 mbar. |  |  | Hardware_pneumatic
FAB-MHFO-009 | Operational requirement | The mechanical horn coils shall vibrate the Membrane which in turn shall generate the pneumatic oscillations. Specifications of Mechanical Horn coils mentioned below: • Rated Voltage: TBD • Rated current: TBD • Holding current: TBD • Peak current: TBD • unidirectional from holding current Notes: For details of controlling parameters please refer HFO PCBA subsystem requirement documents. |  |  | Hardware_pneumatic
FAB-MHFO-010 | Operational requirement | The HFO valve shall be operated by HFO PIC controller. |  |  | Hardware_pneumatic
FAB-MHFO-011 | Operational requirement | The Mechanical Horn coils shall be operated by HFO PIC controller. |  |  | Hardware_pneumatic
FAB-MHFO-012 | Alarm functions | HFO pneumatic subsystem shall trigger alarm to main board PCBA for HFO related fault happening with HFO oscillation. • HFO module shutoff valve electrical fault HFO module oscillator electrical fault |  |  | Hardware_pneumatic
FAB-MHFO-013 | Construction requirement (Subsystem/Process) | The HFO generator diaphragm shall sustain minimum and maximum pressure up to -60 to +140 mbar at HFO Oscillation generator. |  |  | Hardware_pneumatic
FAB-MHFO-014 | Construction requirement (Subsystem/Process) | HFO generator Body shall have exhaust hole to assist pneumatic oscillation and free movement of diaphragm. |  |  | Hardware_pneumatic
FAB-MHFO-015 | Construction requirement (Subsystem/Process) | All pneumatic connections shall incorporate means for mechanical retention. |  |  | Hardware_pneumatic
FAB-MHFO-016 | Construction requirement (Subsystem/Process) | The HFO PCBA & HFO pneumatic subsystem shall have proper mounting to avoid self-dismantling or damages during the function / transportation. |  |  | Hardware_pneumatic
FAB-MHFO-017 | Construction requirement (Subsystem/Process) | The HFO PCBA & HFO pneumatic subsystem when operating in HFO mode shall comply with ISO 80601-2-87:2021 Medical electrical equipment - Part 2-87: Particular requirements for basic safety and essential performance of high frequency ventilators. Current protection: TBD |  |  | Hardware_pneumatic
FAB-MHFO-018 | Material requirement | The material in contact with medical gas shall be compatible with oxygen and AIR. Note 1: Corrosion resistance includes resistance to moisture and surrounding material. Note 2: Compatibility with oxygen involves both combustibility and ease of ignition. Material which bun in AIR will burn violently in pure oxygen. |  |  | Hardware_pneumatic
FAB-MHFO-019 | Material requirement | All pneumatic connections shall consist of following Gas contact components and other components that can be planned with aligned material as mentioned below. Gas Contact Materials: Aluminum (6082 alloy or 6061 alloy), Nitrile rubber, Silicon Rubber, Stainless steel Non-Gas Contact Materials: Aluminum (6082 alloy or 6061 alloy), Low Carbon Steel, Spring Steel, Rare-earth magnet, Off shelf item |  |  | Hardware_pneumatic
FAB-MHFO-XXX | Material requirement | Regulatory compliance applicable are REACH requirements and RoHS requirements for all intended markets. |  |  | Hardware_pneumatic
FAB-MHFO-020 | Physical requirement | The HFO pneumatic subsystem shall have the maximum physical dimensions placed inside as per below specifications: 1. Length – 190mm 2. Width – 170mm 3. Thickness – 170mm |  |  | Hardware_pneumatic
FAB-MHFO-021 | Physical requirement | The HFO pneumatic subsystem shall be mounted on plate with 2 Nos of M5 screws and same plate shall be mounted with mounting holes of TBD mm aligned to mounting threaded inserts at enclosures. |  |  | Hardware_pneumatic
FAB-MHFO-022 | Physical requirement | The HFO PCBA shall have 3 NPTH Mounting holes of 3.2±0.1mm for mounting the card to the studs in enclosure. The mounting holes shall have routing clearance of min. 1mm around their circumference. |  |  | Hardware_pneumatic
FAB-MHFO-023 | Physical requirement | The HFO PCBA & HFO pneumatic subsystem shall be weighted not more than TBDkg. |  |  | Hardware_pneumatic
FAB-MHFO-024 | Operating environmental | The HFO PCBA & HFO pneumatic subsystem shall withstand and perform as intended within the following environmental conditions. • Temperature +0 to +50 °C • Relative Humidity 10 to 90%, non-condensing • Ambient pressure 70 to 106 kPa (10.2 to 15.0 PSIA) Note: the ambient pressure range is intended to equate to approximately -1350 to +10,000 feet altitude. Note: Temperature range of HFO pneumatic subsystem should withstand and perform at higher range of temperature +/-10°C w.r.t device since its generating High Frequency pneumatic oscillation in terms of flow and pressure. |  |  | Hardware_pneumatic
FAB-MHFO-025 | Operating environmental | The HFO PCBA & HFO pneumatic subsystem shall withstand and conform to specifications following shipping & distribution (transport) for up to 72h within the following environmental conditions. • Temperature -20 to +60 °C • Relative Humidity 10 to 95%, non-condensing • Ambient pressure 50 to 106 kPa (7.2 to 15.0 PSIA) Note: the ambient pressure range is intended to equate to approximately -1350 to +10,000 feet altitude. |  |  | Hardware_pneumatic
FAB-MHFO-026 | Other requirement | HFO pneumatic subsystem shall qualify leak test with pressure of 250 mbar. Leak rate should not exceed 10 cc/min when HFO valve open and HFO generator isolated between flexible hose outlet and subsystem. |  |  | Hardware_pneumatic
FAB-MGAS-001 | Pneumatic Interface Requirements | Gas Input subsystem shall provide standard connector to connect separate pneumatic interfaces for 02 as gas input at rear panel for following purpose. 1. Connection for compressed oxygen “O2” supply 2.0 – 6.0 bar (200 - 600 kPa) compatible with flow rate not greater than 40 L/min. Notes: Gas input subsystem shall be capable to supplying sufficient range of flow and pressure to deliver various modes of ventilation/therapy requirement. |  |  | Hardware_gas
FAB-MGAS-002 | Pneumatic Interface Requirements | Gas Input subsystem shall provide standard connector to connect separate pneumatic interfaces for AIR as gas input at rear panel for following purpose. 1. Connection for compressed AIR supply 2.0 – 6.0 bar (200 - 600 kPa) compatible with flow rate not greater than 40 L/min. Notes: Gas input subsystem shall be capable to supplying sufficient range of flow and pressure to deliver various modes of ventilation/therapy requirement. |  |  | Hardware_gas
FAB-MGAS-003 | Pneumatic Interface Requirements | O2 inlet connector at rear panel shall have NIST connector (NIST body adaptor) as per ISO 18082:2014. The specifications of the standard connector as below: Inlet Port: Gas specific NIST Body adaptor with threaded connector. Dimensions of NIST body adaptor is per image and table mentioned below: |  |  | Hardware_gas
FAB-MGAS-004 | Pneumatic Interface Requirements | AIR inlet connector at rear panel shall have NIST connector (NIST body adaptor) as per ISO 18082:2014. The specifications of the standard connector as below: Inlet Port: Gas specific NIST Body adaptor with threaded connector. Dimensions of NIST body adaptor is per image and table mentioned below: |  |  | Hardware_gas
FAB-MGAS-005 | Pneumatic Interface Requirements | O2 inlet connector at rear panel shall be compatible with Gas(O2) specific NIST low pressure connectors (NIST nipple) for medical gasses (attached at gas hose along with NIST nut) only. Connector size: Gas specific NIST Nipple shall have dimensions are per image and table mentioned below: |  |  | Hardware_gas
FAB-MGAS-006 | Pneumatic Interface Requirements | AIR inlet connector at rear panel shall be compatible with Gas (AIR) specific NIST low pressure connectors (NIST nipple) for medical gasses (attached at gas hose) only. Connector size: Gas specific NIST Nipple shall have dimensions are per image and table mentioned below: |  |  | Hardware_gas
FAB-MGAS-007 | Pneumatic Interface Requirements | Gas input subsystem shall be compatible with detachable water trap (filter) at AIR input gas supply line, It provides NIST body adaptor (male) as input connector tilted 90 downwards and NIST nipple and nut (female) as outlet connector coaxial with AIR Gas input subsystem body adaptor without gauge mounting. |  |  | Hardware_gas
FAB-MGAS-008 | Pneumatic Interface Requirements | Water trap have following specification. Pressure bearing capacity shall be 10 bar. Note: The inlet filters provide protection for the valves and blender flow sensor from contamination. |  |  | Hardware_gas
FAB-MGAS-009 | Pneumatic Interface Requirements | Gas input subsystem of O2 supply line) shall have separate output port as standard push pneumatic connector. Maximum Inlet Pressure rating: 10 Bar Type of Connector: Push fit tubing connector Standard polyurethane Tubing: 6 mm OD Connector size: Straight - Straight connector Threaded/6 mm tubing Tubing Color: Gray |  |  | Hardware_gas
FAB-MGAS-010 | Pneumatic Interface Requirements | Gas input subsystem of AIR supply line shall have separate output port as standard push pneumatic connector. Maximum Inlet Pressure rating: 10 Bar Type of Connector: Push fit tubing connector Standard polyurethane tubing: 6 mm OD Connector size: Straight - Straight connector Threaded/6 mm tubing Tubing Color: Black |  |  | Hardware_gas
FAB-MGAS-011 | Operational requirement | Gas input subsystem’s each component and joint (sealed with O ring / sealant/ Loctite) shall sustain 10 bar maximum inlet pressure. |  |  | Hardware_gas
FAB-MGAS-012 | Operational requirement | Gas input subsystem shall comply following functional leak requirements. Leak test: Leakage test for subsystem shall not exceed 50ml/min at following test parameters. a. Compressed gas pressure 6.0 bar with blocked output’s b. At (23+/-2 deg C) Test Gas: Carry out test specific gas for which gas inlet is designed with dry gas with maximum moisture content of 50 microgram/gram, corresponding to dew point of -48 deg C. |  |  | Hardware_gas
FAB-MGAS-013 | Operational requirement | Gas input subsystem shall have inbuilt sintered filters having a pore size of less than or equal to 100 microns to particulate inlet filters to provide protection for the valves and blender flow sensor from contamination. Specification of sintered filer is mentioned below. Shape: Conical 100 microns perforation Maximum flow rate: 40 L/min TBD Note: Sintered filter to be provided with a particulate filter at each gas (O2 and AIR) subsystem. |  |  | Hardware_gas
FAB-MGAS-014 | Operational requirement | Gas input subsystem (with inbuilt filter) shall comply following functional flow restriction requirements. The pneumatic output shall be 40 l/min least with at 2 bar pressure applied at gas input subsystem, pressure drop of input to output to be less than 0.2 bar. Notes: Carry out test at (23+/-2 deg C). |  |  | Hardware_gas
FAB-MGAS-015 | Construction Requirements | The ventilator shall provide the gas O2/AIR input construct/connector in such way so that they are non-interchangeable to prevent wrong pneumatic connections. Note: Purpose of this requirement is to verify Probe of O2 and Air can be connected to their specific NIST connector only in order to prevent wrong connection. This design feature is implemented by NIST nipple and body adapter in accordance with ISO: 18082. |  |  | Hardware_gas
FAB-MGAS-016 | Construction Requirements | AIR/O2 Gas input subsystem shall have locking mechanism (anti rotation) at rear panel internally mounted to avoid self-dismantling or damages due to vibration during the function / transportation. |  |  | Hardware_gas
FAB-MGAS-017 | Construction Requirements | All pneumatic connections at Gas input subsystem components shall be constructed by following material. • NIST Connector: – Stainless steel • Lock Nut: Stainless steel • Lock pin: Stainless steel • O-Ring - Nitrile rubber (NBR) • Any other seal -Silicon Rubber • Sealant/Loctite: TBD |  |  | Hardware_gas
FAB-MGAS-018 | Labeling Requirement | Each of the NIST connector shall be laser marked or labelled with the gas name or chemical symbol (“AIR” and “O2” for oxygen) for proper identification while assembly and storage, since they are identical in design. |  |  | Hardware_gas
FAB-MGAS-019 | Biocompatibility Requirements | The material in contact with medical gas shall be compatible with oxygen and AIR. . Note 1: Corrosion resistance includes resistance to moisture and surrounding material. Note 2: Compatibility with oxygen involves both combustibility and ease of ignition. Material which bun in AIR will burn violently in pure oxygen. |  |  | Hardware_gas
FAB-MGAS-020 | Physical Requirements | Gas input subsystem shall be place inside the ventilator enclosure and mounted on rear panel with proper mechanical retention. |  |  | Hardware_gas
FAB-MGAS-021 | Physical Requirements | The Gas input subsystem shall have the maximum physical dimensions as per below specifications: Approximate 50% of length portion shall be projected outside of enclosure(rear panel). 1. Length – 100mm 2. Width – 40 mm 3. Thickness – 40mm |  |  | Hardware_gas
FAB-MGAS-022 | Environmental requirements | The gas input subsystem shall withstand and perform as intended within the following environmental conditions. • Temperature +10 to +40 °C • Relative Humidity 10 to 90%, non-condensing • Ambient pressure 70 to 106 kPa (10.2 to 15.0 PSIA) Note: The ambient pressure range is intended to equate to approximately -1350 to +10,000 feet altitude. |  |  | Hardware_gas
FAB-MGAS-023 | Environmental requirements | The gas input subsystem shall withstand and conform to specifications following shipping & distribution (transport) for up to 72h within the following environmental conditions. • Temperature -20 to +60 °C • Relative Humidity 10 to 95%, non-condensing • Ambient pressure 50 to 106 kPa (7.2 to 15.0 PSIA) |  |  | Hardware_gas
FAB-EBLN-001 | Interface Connectors Requirement | The Blender PCBA shall have 20 position connector (2X10) interface to connect with the Main Board PCBA. The Specification of the connector shall be as below: 1. No. Of position – 20 (2X10) 2. Pitch – 2.54 mm 3. Style – Board to cable 4. Contact mating length – 6.10±0.1mm 5. Flammability rating – UL94 V-0 6. Keying slot – Yes |  |  | Hardware_BlenderPCBA
FAB-EBLN-002 | Interface Connectors Requirement | The Blender PCBA connector interface with the Main Board PCBA shall have the below signal specifications: 1. MGND (Pin#1, 5) – Module Ground connection 2. AGND (Pin#10,11,12) – Analog Ground connection 3. TP100 (Pin#20) – Test Point 4. VM (Pin#7,8,9) – 24 Volts supply Input 5. MVCC (Pin#14,15)) – 5V supply 6. DVCC (Pin#13,19) – Digital 5V Supply 7. DRM1 & DRM2 (Pin#4&3)– Pressure Transducer signal O/P; Type Analog; Value – 0~3V max 8. MOSI, MI_STAT, MCS & MCLK (Pin# 17,6,16&18) – SPI lines to Main Board PCBA, Type – Digital I/O, 5V logic (Note: SPI communication specifications are mentioned in the Blender PIC requirement) |  |  | Hardware_BlenderPCBA
FAB-EBLN-003 | Interface Connectors Requirement | The Blender PCBA shall communicate with the Mainboard PCBA via SPI communication medium and have the below mentioned signal specifications: 1. MI_DATA (Pin#17) – SPI MOSI signal at 5V logic 2. MI_STAT (Pin#6) – SPI MISO signal at 5V Logic 3. MCS (Pin#16) – SPI Chip select signal at 5V Logic 4. MCLK (Pin#18) – SPI clock at minimum XXMHz |  |  | Hardware_BlenderPCBA
FAB-EBLN-004 | Interface Connectors Requirement | The Blender PCBA shall have 2 quantity of 4 Pole connectors for Hot wire flow sensor as per below mentioned specifications: (* The PCB connector part no. is not available in the BOM) |  |  | Hardware_BlenderPCBA
FAB-EBLN-005 | Interface Connectors Requirement | The Hot wire flow sensor connectors (quantity 2) shall have the below mentioned signal specifications: 1. Pin#1 – NC 2. Pin#2 – HA1A1 OR HA1B1 (Flow sensor signal) 3. Pin#3 – HA2A1 OR HA2B1 (Flow sensor signal) 4. Pin#4 - NC |  |  | Hardware_BlenderPCBA
FAB-EBLN-006 | Interface Connectors Requirement | The Blender PCBA shall have a 2 Pole connector for the “Air Proportional valve drive” output. The connector shall have the below specification: 1. No. Of Pos. – 2 2. Pitch – 2.54 mm 3. Style – Board to cable 4. Connector height from board – 14±0.1mm 5. Flammability rating – UL94 V-0 6. Keying slot – Yes 7. Current rating per contact – Minimum 238mA |  |  | Hardware_BlenderPCBA
FAB-EBLN-007 | Interface Connectors Requirement | The “Air Proportional valve Drive” connector shall have the below mentioned signal specifications: 1. Pin#1 – Varying voltage up to 12V Max.; Output current 0~238 mA 2. Pin#2 – VGND; Ground potential |  |  | Hardware_BlenderPCBA
FAB-EBLN-008 | Interface Connectors Requirement | The Blender PCBA shall have a 2 Pole connector for the “Oxygen Proportional valve drive” output. The connector shall have the below specification: 1. No. Of Pos. – 2 2. Pitch – 2.54 mm 3. Style – Board to cable 4. Connector height from board – 14±0.1mm 5. Flammability rating – UL94 V-0 6. Keying slot – Yes 7. Current rating per contact – Minimum 238mA |  |  | Hardware_BlenderPCBA
FAB-EBLN-009 | Interface Connectors Requirement | The “Oxygen Proportional valve Drive” connector shall have the below mentioned signal specifications: 1. Pin#1 – Varying voltage up to 24V Max.; Output current 0~238 mA 2. Pin#2 – VGND; Ground potential |  |  | Hardware_BlenderPCBA
FAB-EBLN-010 | Interface Connectors Requirement | The Blender PCBA shall have a 2 Pole connector for the “Air Calibration valve drive” output. The connector shall have the below specification: 1. No. Of Pos. – 2 2. Pitch – 2.5 mm 3. Style – Board to cable 4. Connector height from board – 7±0.25mm 5. Flammability rating – UL94 V-0 |  |  | Hardware_BlenderPCBA
FAB-EBLN-011 | Interface Connectors Requirement | The “Air Calibration valve Drive” connector shall have the below mentioned signal specifications: 1. Pin#1 – 24V supply; Output current TBD A Max 2. Pin#2 – Ground potential |  |  | Hardware_BlenderPCBA
FAB-EBLN-012 | Interface Connectors Requirement | The Blender PCBA shall have a 2 Pole connector for the “Oxygen Calibration valve drive” output. The connector shall have the below specification: 1. No. Of Pos. – 2 2. Pitch – 2.5 mm 3. Style – Board to cable 4. Connector height from board – 7±0.25mm 5. Flammability rating – UL94 V-0 |  |  | Hardware_BlenderPCBA
FAB-EBLN-013 | Interface Connectors Requirement | The “Oxygen Calibration valve Drive” connector shall have the below mentioned signal specifications: 1. Pin#1 – 24V supply; Output current TBD A Max 2. Pin#2 – Ground potential |  |  | Hardware_BlenderPCBA
FAB-EBLN-014 | Interface Connectors Requirement | The Blender PCBA shall have a 6-pos. programming connector for programming the Blender PIC controller. The Programming connector shall have the below specifications: 1. No. Of Pos. – 6 (1X6) 2. Pitch – 2.54 mm 3. Style – Board to Board or Cable 4. Connector height from board – 9.49±0.1mm 5. Flammability rating – UL94 V-0 |  |  | Hardware_BlenderPCBA
FAB-EBLN-015 | Interface Connectors Requirement | The Blender PIC Programming connector shall have the below signal specifications: 1. PCLR# (Pin#1) – Programming Pin; Digital I/O at 3V3 logic. 2. A3V3 (Pin#2) – 3V3 supply pin 3. AGND (Pin#3) – Ground Pin 4. PGD2 (Pin#4) – Programming Pin (Data); Digital I/O at 3V3 logic. 5. PGC2 (Pin#5) - Programming Pin (Clock); Digital O/P at 3V3 logic. 6. NC (Pin#6) – Not connected Note: No Capacitors, pullups or diodes are recommended on the PGC & PGD pins |  |  | Hardware_BlenderPCBA
FAB-EBLN-016 | Interface Connectors Requirement | The Blender PCBA shall have a 4-pos. “Debug” connector for debugging the Blender PIC controller. The Debug connector shall have the below specifications: 1. No. Of Pos. – 4 (1X4) 2. Pitch – 2.54 mm 3. Style – Board to Board or Cable 4. Connector height from board – 9.49±0.1mm 5. Flammability rating – UL94 V-0 |  |  | Hardware_BlenderPCBA
FAB-EBLN-017 | Interface Connectors Requirement | The Blender PIC Debugging connector shall have the below signal specifications: 1. AVCC (Pin#1) – 3V3 supply pin 2. U1TX (Pin#2) – Data Transmission pin; Digital I/O at 3V3 logic. 3. U1RX (Pin#3) – Data Receiving pin; Digital I/O at 3V3 logic. 4. AGND (Pin#4) – Ground Pin Note: A Pull down of 10K is required in the U1RX signal pin |  |  | Hardware_BlenderPCBA
FAB-EBLN-018 | Inspiratory Pressure Sensing Requirement | The Blender PCBA shall have 2 quantity of Inspiratory pressure transducer, one each for Oxygen and Air Inputs. The transducers (used for sensing inspiratory pressures) shall have the below specifications: 1. Pressure Type - Vented Gauge 2. Operating Pressure – Minimum 1.16 PSI (80mbar) 3. Output Type - Wheatstone Bridge 4. Output - 0 mV ~ 90 mV (12V) 5. Accuracy - ±2% 6. Response time – 100uS±2% (at 10-90% of FSS) 7. Voltage – Supply - 3V ~ 16V 8. Port Size - Male - 4.83mm±0.08mm 9. Port Style - Barbed 10. Mounting Type – Through Hole PCB (Note: The pressure sensors shall be classified as high integrity components subject to purchasing controls.) |  |  | Hardware_BlenderPCBA
FAB-EBLN-019 | Inspiratory Pressure Sensing Requirement | The Blender PCBA shall have 2 quantity of Instrumentation Amplifiers one each for Oxygen and Air Pressure transducer Inputs. The Amplifiers to be used for amplifying the inspiratory pressures “DRM1 & DRM2” to an analog value ranging from 0~3V that can be read by the ADC of the Blender PIC Controller |  |  | Hardware_BlenderPCBA
FAB-EBLN-020 | Hot Air Flow Sensing Requirement | The Blender PCBA shall include 2 calibration potentiometer circuits for calibration of the HWFS signals “HA1A1 / HA1B1” and “HA2A1 / HA2B1”. The calibration circuit shall be created with Rail to Rail O/P Op-Amps having the below specifications: 1. Amplifier Output Type - Rail-to-Rail 2. Slew Rate - 7V/µs 3. Voltage Supply - 5V±10% 4. Voltage Input Offset - 250 µV |  |  | Hardware_BlenderPCBA
FAB-EBLN-021 | Hot Air Flow Sensing Requirement | The Blender PCBA shall include 2 of Rail to Rail amplifier circuits for amplification of the HWFS signals “HA1A1 / HA1B1” and “HA2A1 / HA2B1”. The amplifier circuit shall be created with Rail to Rail O/P Op-Amps having the below specifications: 1. Amplifier Output Type - Rail-to-Rail 2. Slew Rate - 7V/µs 3. Voltage Supply – 5V±10% 4. Gain Bandwidth Product - 10 MHz 5. Voltage Input Offset - 500 µV (The amplified signals of flow sensors “Flow1 and Flow2” will not exceed 3V) |  |  | Hardware_BlenderPCBA
FAB-EBLN-022 | Hot Air Flow Sensing Requirement | The Blender PCBA shall have a 2 channel 12 Bit DAC that acts as a digitally controlled potentiometer for generating precise reference voltages “VREFA” and “VREFB” for the Calibration circuits. The DAC shall have the below specifications: 1. Number of Bits - 12 2. Number of D/A Converters - 2 3. Settling Time - 4.5µs (Typ) 4. Output Type - Voltage - Buffered 5. Data Interface – I2C 6. Reference Type - External 7. Voltage – Supply - 5V±10% (Note: Its reference input “VREFB” & “VREFA” is driven by the bridge supply voltage while its output voltages “VOUTB” & “VOUTA” can tune the bridge by ±22 % through a resistor. |  |  | Hardware_BlenderPCBA
FAB-EBLN-023 | Hot Air Flow Sensing Requirement | The Flow sensor signals “Flow1 & Flow2” shall be analog Type and vary from 0~TBDV (The Flow rates shall not be more than 40lpm from either inlet connection) |  |  | Hardware_BlenderPCBA
FAB-EBLN-024 | Proportional Valve Drive Requirement (Oxygen & Air) | The Blender PCBA shall include a 2 channel Rail to Rail unity gain Amplifier that can buffer the PWM signals “PWM_L” generated by Blender PIC for the proportional drive of the Air Valve and “PWM_O” generated by Blender PIC for the proportional drive of the Oxygen Valve. The unity gain amplifiers shall act as a buffers for transmitting the PWM signals from the Blender PIC to the Buck converters responsible for the proportional Valve drives. |  |  | Hardware_BlenderPCBA
FAB-EBLN-025 | Proportional Valve Drive Requirement (Oxygen & Air) | The Blender PCBA shall include 2 DC-DC Buck converter that can generate the requisite Power required for the proportional Valve drive for the Oxygen & Air. The DC-DC converter shall have the below specifications: 1. Driver Function - Step-Down 2. Output Configuration – Positive, Adjustable 3. Topology - Buck 4. Voltage Input – 24VDC±10% 5. Voltage Output – 12V±10% 6. Output Current – 0 to 211mA±3% |  |  | Hardware_BlenderPCBA
FAB-EBLN-026 | Proportional Valve Drive Requirement (Oxygen & Air) | The Proportional Valve drive outputs “PROP_LUFT” for Air and “PROP_OXI” for Oxygen shall range from 0 to 211mA±3% at 12V for a varying duty cycle from 0-100% (Note: The proportional valves shall have proof pressure at least10 bar, and be classified as high integrity components subject to purchasing controls) |  |  | Hardware_BlenderPCBA
FAB-EBLN-027 | Calibration Valve Drive Requirement (Oxygen & Air) | The Blender PCBA shall include 1 converter that can generate the Power required for the Power Switch to drive the individual Calibration Valves for Oxygen & Air. The Converter shall have the below specifications: 1. Output Configuration – Positive, Adjustable, single O/P 2. Topology - Boost, Flyback, Forward Converter 3. Voltage Input – 24V±20% 4. Voltage Output – 24V±5% 5. Output Current – Minimum 500mA |  |  | Hardware_BlenderPCBA
FAB-EBLN-028 | Calibration Valve Drive Requirement (Oxygen & Air) | The Blender PCBA shall include a 2 N-Channel Power switch that can distribute the power required for the Air & Oxygen Calibration Valves as and when their channels activated 1. Number of Outputs - 2 2. Output Type – N-Channel (On/Off interface) 3. Output Configuration – Low Side 4. Input Type – Non - Inverting 5. Voltage Load Max – 24V±5% 6. Output Current – Minimum 250mA each for both the Output drives (“VENT_LUFT” & “VENT_OXI”) |  |  | Hardware_BlenderPCBA
FAB-EBLN-029 | Calibration Valve Drive Requirement (Oxygen & Air) | The N channel Power switch shall get the below mentioned control signals from the Blender PIC for operating the calibration valves: 1. Air calibration Valve signal “LUFT_ON” – Active low signal. 2. Oxygen calibration Valve signal “VENT_OXI” - Active low signal. (Note: Both the control signal requires Pull up of 3K3 with the 5V supply) |  |  | Hardware_BlenderPCBA
FAB-EBLN-030 | Blender PIC Controller Requirement | The Blender PIC controller shall have the following functional features: 1. Operating voltage - 1.8-5.5V 2. Operating speed / Clock frequency - 8MHz±5% 3. Memory – Program Memory (64KB); RAM (3.6KB) 4. Analog Channels – Min 8 (of 12 Bit) 5. Digital I/O s – Min. 9 6. Other Interface requirements (Min. Nos.) – 2XPWM; 1 UART; 1XSPI and 2XI2C |  |  | Hardware_BlenderPCBA
FAB-EBLN-031 | Blender PIC Controller Requirement | The Blender PIC Controller shall have 1 SPI interface to communicate with the Main Board. The Signal Specifications of the SPI Interface shall be as below: 1. MCS – SPI Chip select 2. MCLK – SPI Clock 3. MI_STAT – SPI MISO 4. MOSI – SPI MOSI |  |  | Hardware_BlenderPCBA
FAB-EBLN-032 | Blender PIC Controller Requirement | The Blender PIC Controller shall have 3 Analog channels to monitor the below mentioned Voltage supplies: 1. AVCC – 5V supply (Sensing signal ranges from 0~3V) 2. VM – 24 V supply (Sensing signal ranges from 0~2.2V) 3. +24V – 24 V Supply (Sensing signal ranges from 0~2.7V) |  |  | Hardware_BlenderPCBA
FAB-EBLN-033 | Blender PIC Controller Requirement | The Blender PIC Controller shall have 2 Analog channels to monitor the Flow sensor signals: 1. FLOW1 – Analog signal 0~3V 2. FLOW2 – Analog signal 0~3V |  |  | Hardware_BlenderPCBA
FAB-EBLN-xxx | Blender PIC Controller Requirement | The Blender PIC Controller shall have 1 Analog channel to monitor the Inspiratory Pressure transducer signal of the Manifold PCBA: 1. DRM3_A – Analog signal 0~3V |  |  | Hardware_BlenderPCBA
FAB-EBLN-034 | Blender PIC Controller Requirement | The Blender PIC controller shall have 1 Analog channel to measure the VREF+ Signal of 3.3V |  |  | Hardware_BlenderPCBA
FAB-EBLN-035 | Blender PIC Controller Requirement | The Blender PIC Controller shall have 1 I2C communication channel to interface with the 12 Bit DAC responsible for generating the reference voltages for Flow sensor circuitry. The I2C signals shall be accompanied with a Chip select signal “DAC_CS#”. The signal specifications shall be as below: 1. DAC_C – I2C clock; Digital O/P at 3V3 logic 2. DAC_D – I2C Data; Digital I/O at 3V3 logic 3. DAC_CS# – DAC Chip select; Digital I/O at 3V3 logic (Note: All the above 3 signal lines shall be pulled up with “AVCC” with 4K7 resistors) |  |  | Hardware_BlenderPCBA
FAB-EBLN-036 | Blender PIC Controller Requirement | The Blender PIC Controller shall have 2 Digital output signals to enable the Proportional drive of the Air and Oxygen Valves. The signal specifications shall be as below: 1. P_L_EN – Enable signal for Air proportional valve drive; Digital O/P at 3V3 logic 2. P_O_EN – Enable signal for Oxygen proportional valve drive; Digital O/P at 3V3 logic |  |  | Hardware_BlenderPCBA
FAB-EBLN-037 | Blender PIC Controller Requirement | The Blender PIC Controller shall have 2 PWM output signals to feed the amplifiers of the Proportional drive circuits of the Air and Oxygen Valves. The signal specifications shall be as below: 1. PWM_L – PWM signal for Air proportional valve drive; Freq > 1KHz; Duty cycle 0-100%; 3V3 Logic 2. PWM_O – PWM signal for Oxygen proportional valve drive; Freq > 1KHz; Duty cycle 0-100%; 3V3 Logic |  |  | Hardware_BlenderPCBA
FAB-EBLN-038 | Blender PIC Controller Requirement | The Blender PIC Controller shall have one Digital output signal to enable the Forward converter responsible for generating power for the calibration valves. The signal specifications shall be as below: 1. 24V_EN – Enable signal for Forward converter; Digital O/P at 3V3 logic |  |  | Hardware_BlenderPCBA
FAB-EBLN-039 | FRAM Requirement | The Blender PCBA shall have a 64KBit FRAM to store the calibration data, having below specifications 1. Memory Type - Non-Volatile 2. Memory Format - FRAM 3. Memory Size - 64Kb (8K x 8) 4. Memory Interface - I²C 5. Voltage Supply – 3V3±5% |  |  | Hardware_BlenderPCBA
FAB-EBLN-040 | FRAM Requirement | The FRAM shall communicate with the Blender PIC controller via I2C communication. The signal specifications of the I2C communication shall be as below: 1. SCL – I2C clock line; Digital O/P with 3V3 logic 2. SDA – I2C data line; Digital I/O with 3V3 Logic (Note: both the SCL and SDA lines shall be pulled up with a 4K7 resistor with 3V3 supply) |  |  | Hardware_BlenderPCBA
FAB-EBLN-041 | General requirements | The Blender PCBA shall have multiple layer design (4 Layers). There shall be at least 2 Ground planes for noise immunity. |  |  | Hardware_BlenderPCBA
FAB-EBLN-042 | General requirements | The Blender PCBA shall satisfy the below mentioned manufacturing standards: 1. Flammability standard of UL94 V-0 2. ROHS Directive 2011/65/EU 3. REACH Directive (EG) Nr. 1907/2006 4. IPC-A-600 – Acceptability of Printed Boards 5. IPC-SM-840 – Qualification & Performance spec for Permanent solder mask |  |  | Hardware_BlenderPCBA
FAB-EBLN-043 | General requirements | The Blender PCBA shall be developed on FR4 material and having 35u copper finish on each of the layers |  |  | Hardware_BlenderPCBA
FAB-EBLN-044 | General requirements | The Blender PCBA shall have clear and specific marking of all the mating connectors to different modules or sub systems |  |  | Hardware_BlenderPCBA
FAB-EBLN-045 | General requirements | The Blender PCBA shall have labels to define the revision no. of hardware and the software applicable to that version. |  |  | Hardware_BlenderPCBA
FAB-EBLN-046 | Physical Requirements | The Blender PCBA shall have the maximum physical dimensions as per below specifications: 1. Length – 164mm 2. Width – 75mm 3. Thickness – 1.6mm |  |  | Hardware_BlenderPCBA
FAB-EBLN-047 | Physical Requirements | The Blender PCBA shall have 6 NPTH Mounting holes of 6.5±0.1mm (having routing clearance of minimum 0.2mm across the hole circumference) for mounting the card to the studs in enclosure |  |  | Hardware_BlenderPCBA
FAB-EETSP-001 | Interface Requirements | The EtCO2 & SpO2 PCBA shall have a 7position mating header connector to receive input signals from the EtCO2 Sensor via sensor cable assembly having below specifications: 1. Number of Positions: 7 (1X7) 2. Pin Pitch: 2 mm 3. Connector System: Wire-to-Board 4. PCB Mount Orientation: Vertical 5. Contact Current Rating (Minimum) (A): TBD 6. Flammability Rating: UL94 V-0 |  |  | Hardware_SpO2PCBA
FAB-EETSP-002 | Interface Requirements | The 7 position EtCO2 mating connector interface shall have the below signal specifications: 1. +5V_CO2 (Pin# 1, 2): +5V DC Supply to sensor 2. GND_CO2 (Pin# 3, 4): Ground 3. RXD3_CO2 (Pin# 5): Receive Data at 5V logic 4. TXD3_CO2 (Pin# 6): Transmit Data at 5V logic 5. STAT_CO2 (Pin# 7): Status signal (Pulled down at Main Board) |  |  | Hardware_SpO2PCBA
FAB-EETSP-003 | Interface Requirements | The EtCO2 & SpO2 PCBA shall have a 5-position mating connector interface to receive inputs from the SpO2 Sensor via sensor cable assembly having below specifications: 1. Number of Positions: 5 (1X5) 2. Pin Pitch: 2 mm 3. Connector System: Wire-to-Board 4. PCB Mount Orientation: Vertical 5. Contact Current Rating (Minimum) (A): TBD 6. Flammability Rating: UL94 V-0 |  |  | Hardware_SpO2PCBA
FAB-EETSP-004 | Interface Requirements | The SpO2 sensor mating connector interface shall have the below signal specifications: 1. GND_SpO2 (Pin# 1, 2): Ground connection 2. 3V3_SpO2 (Pin# 3): 3.3V supply 3. RXD_SpO2 (Pin# 4): Receive Data 4. TXD_SpO2 (Pin# 5): Transmit Data |  |  | Hardware_SpO2PCBA
FAB-EETSP-005 | Interface Requirements | The EtCO2 & SpO2 PCBA shall have a common Connector interface of 14 position to transmit the EtCO2 and SpO2 Sensor Signals to the Main Board as per below specifications: 1. Number of Positions: 14 (2X7) 2. Pin Pitch: 2.54 mm 3. Connector System: Wire-to-Board 4. PCB Mount Orientation: Vertical 5. Contact Current Rating (Max) (A): 2.2 6. Flammability Rating: UL94 V-0 |  |  | Hardware_SpO2PCBA
FAB-EETSP-006 | Interface Requirements | The 14 position Mainboard Connector Interface shall have the below signal specifications: 1. +5V (Pin# 1, 2): +5V DC Supply 2. STAT (Pin# 3): Status Signal 3. NC (Pin# 4): Not Connected 4. TXD3 (Pin# 5): Transmit Data 5. RXD3 (Pin# 6): Receive Data 6. GND (Pin# 7, 8, 9, 10, 11) Ground 7. USB+ (Pin# 12) USB + 8. USB- (Pin# 13) USB – 9. VBUS+ (Pin# 14) USB Supply |  |  | Hardware_SpO2PCBA
FAB-EETSP-007 | Interface Requirements | The EtCO2 & SpO2 PCBA shall be connected with the Mainboard via cable interface. The Flat Ribbon Cable interface terminated at the Main Board Interface Connector shall have the below Specifications: 1. Cable length: 232 mm±10mm 2. Number of cores: 14 3. Interface connector type: Receptacle (Female socket) 4. Connector Pitch: 2.54mm |  |  | Hardware_SpO2PCBA
FAB-EETSP-008 | EtCO2 & SPO2 Signal Conditioning Requirements | All the EtCO2 & SPO2 sensor signal and power lines shall be protected from Human body ESD air discharge of 15KV. |  |  | Hardware_SpO2PCBA
FAB-EETSP-009 | EtCO2 & SPO2 Signal Conditioning Requirements | The SpO2 Sensor Power supply and the output signals shall be galvanically isolated to satisfy regulatory safety requirements per IEC 60601-1. The below specifications shall be complied with respect to power lines Isolation: 1. Classification: Two MOPP 2. Isolation: 4KV 3. Creepage: 4 mm 4. Clearance: 2.3mm |  |  | Hardware_SpO2PCBA
FAB-EETSP-010 | EtCO2 & SPO2 Signal Conditioning Requirements | The EtCO2 sensor shall transmit its sensor data to the mainboard via UART communication at 5V Logic |  |  | Hardware_SpO2PCBA
FAB-EETSP-011 | EtCO2 & SPO2 Signal Conditioning Requirements | The SPO2 sensor shall transmit via serial communication using RX and TX signal lines. These RX and TX signals shall be converted to a USB data format for the communication with the Mainboard |  |  | Hardware_SpO2PCBA
FAB-EETSP-012 | General requirements | The EtCO2 & SpO2 PCBA shall have multiple layer design (4 Layers). There shall be at least 1 Ground plane and 1 power plane for noise immunity. |  |  | Hardware_SpO2PCBA
FAB-EETSP-013 | General requirements | The EtCO2 & SpO2 PCBA shall satisfy the below mentioned manufacturing standards: 1. Flammability standard of UL94 V-0 2. ROHS Directive 2011/65/EU 3. REACH Directive (EG) Nr. 1907/2006 4. IPC-A-600 Class 2 – Acceptability of Printed Boards 5. IPC-SM-840 – Qualification & Performance spec for Permanent solder mask |  |  | Hardware_SpO2PCBA
FAB-EETSP-014 | General requirements | The EtCO2 & SpO2 PCB shall be developed on FR4 material (complying to UL 94V-0) and having min 35u copper finish on external layers and 17.5u copper finish on inner Layers |  |  | Hardware_SpO2PCBA
FAB-EETSP-015 | General requirements | The EtCO2 & SpO2 PCBA shall have clear and specific marking of all the mating connectors to different modules or sub systems. |  |  | Hardware_SpO2PCBA
FAB-EETSP-016 | General requirements | The EtCO2 & SpO2 PCBA shall have labels to define the revision no. of hardware/software applicable to that version. |  |  | Hardware_SpO2PCBA
FAB-EETSP-017 | Physical Requirements | The EtCO2 & SpO2 PCBA shall have max. physical dimensions as per below specifications: 1. Length: 75 mm±5% 2. Width: 55.5 mm±5% 3. Thickness: 1.6mm±10% |  |  | Hardware_SpO2PCBA
FAB-EETSP-018 | Physical Requirements | The EtCO2 & SpO2 PCBA shall have 4 Mounting holes of 3.3mm dia. for mounting the card to the studs in enclosure The Mounting holes shall have a routing clearance of at least 0.2mm across the mounting hole |  |  | Hardware_SpO2PCBA
FAB-EHFB-001 | Connectors Interface Requirements | The HFO PCBA shall have 26 position connector (2X13) of 2.54mm pitch for Mainboard PCBA connectivity with the below specifications: 1. No. Of Pos. – 26 (2X13poles) 2. Pitch – 2.54 mm 3. Contact mating length – 6.10±0.1mm 4. Flammability rating – UL94 V-0 5. Keying slot - Yes |  |  | Hardware_HFOPCBA
FAB-EHFB-002 | Connectors Interface Requirements | The HFO PCBA connector shall have the below signal specifications: 1. PROX (Pin#1) – Proximal Pressure sensor Input; Analog Voltage I/P; 0~4.096V max. 2. MCLK (Pin#2) – SPI Clock Pin; Digital I/O pin at 5V logic 3. VAVS_A(Pin#3) – 5V Supply switched by the Safety relay that supplies power mainly to Manifold Board. 4. MDATA (Pin#4) – SPI MISO signal for the HFO PIC controller; Digital I/O at 5V logic 5. HF_CS (Pin#5) – SPI chip select signal; 5V active low signal. 6. +5V (Pin#6) – 5V supply for the HFO card 7. HF_STAT (Pin#7) – SPI MOSI signal for the HFO PIC controller; Digital I/O at 5V logic 8. DGND (Pin#8) – Digital Ground 9. PGND (Pin#9-15) – Power Ground for the HFO Board 10. +VHF (Pin#16-22) – 24V supply for the HFO Board 11. UPDRX_2 (Pin#25) – Receiver pin of the UART line; Digital I/O at 5V Logic. 12. UPDTX_2 (Pin#26) – Transmitter pin of the UART line; Digital I/O at 5V Logic. 13. TP1 (Pin#24) – Test Point (Not used) 14. Pin#23 – Not connected. |  |  | Hardware_HFOPCBA
FAB-EHFB-003 | Connectors Interface Requirements | The HFO PCBA shall have 6 pin RJ11 connector for programming and debugging purpose. |  |  | Hardware_HFOPCBA
FAB-EHFB-004 | Connectors Interface Requirements | The HFO PCBA programming connector shall have following signal specifications connected to HFO PIC Controller: 1. MCLR - Reset Signal for the controller 2. +5V - Supply voltage for the Programming interface 3. DGND - Supply Ground for the Programming interface 4. PGD - Programming Data for the Controller 5. PGC - Programming Clock for the Controller 6. NC - Not Connected |  |  | Hardware_HFOPCBA
FAB-EHFB-005 | Power supply Requirements | The HFO PCBA shall receive “+5V” and “+VHF” (24V) from the Main board. The current drawn on “+5V” shall not exceed TBD A and +VHF TBD A |  |  | Hardware_HFOPCBA
FAB-EHFB-006 | Power supply Requirements | The “+VHF” supply shall be filtered using an LC low pass filter of TBD Hz cut off frequency. |  |  | Hardware_HFOPCBA
FAB-EHFB-007 | Power supply Requirements | The HFO PCBA shall generate a stable “+24V” on board from Switch Mode Regulator for operating the HFO Valve. This regulator shall have the following specifications: 1. Supported Topology – Boost /Flyback / Forward Converter 2. Voltage Input Max – 24V±10% 3. Voltage Output – 24V±5% 4. Output Configuration – Positive The current delivered by the Switch Mode regulator will be minimum TBD A (Note: The HFO PIC controller shall provide an active high signal “24_EN” of 5V to enable and disable the +24V Power Supply) |  |  | Hardware_HFOPCBA
FAB-EHFB-008 | Power supply Requirements | The HFO PCBA shall generate a 12 V DC supply to power the Class D Amplifier having below specification: 1. Output Type – Fixed, 12V±5% 2. Quiescent (Iq) Current - < 5mA 3. Voltage – Maximum 24V±10% For biasing the Class D amplifier feedback inputs, a signal of 12V “V_MITTE” shall be generated using transistor circuits. |  |  | Hardware_HFOPCBA
FAB-EHFB-009 | HFO PIC Controller requirement | The HFO PIC controller shall have the following functional features for the HFO PCBA: 1. Operating voltage – 5V±10% 2. Operating speed (Max) – 14.7456 MHz +/-50 PPM 3. Memory – Program Memory (16KB); Data Memory (768B) 4. Analog Channels – Min 1 (of 10 bit) 5. Digital I/O s – Min 13 6. Other Interface requirements (Minimum Nos.) – 1 SPI and 1 UART |  |  | Hardware_HFOPCBA
FAB-EHFB-010 | HFO PIC Controller requirement | The HFO PIC controller shall generate 2 control signals “SIGNAL” and “OFFS” at 5V logic. The Signal specifications of these 2 control signals shall be as below: 1. “SIGNAL” – PWM signal of variable frequency from “XXHz-YYHz” 2. “OFFS” - PWM signal of frequency “XXHz-YYHz” and duty cycle “XX-YY”. |  |  | Hardware_HFOPCBA
FAB-EHFB-011 | HFO PIC Controller requirement | The PWM signals “SIGNAL” and “OFFS” generated by HFO PIC controller shall be used to generate sinusoidal signal to drive the horn and generate high frequency oscillations. The frequency of the signal shall be varied from 5 Hz to 20 Hz. The Amplitude of the signal shall be varied from XX V to YY resulting in pressure change at the membrane level between 5-100 cmH2O. |  |  | Hardware_HFOPCBA
FAB-EHFB-012 | HFO PIC Controller requirement | The HFO PIC controller shall have 1 analog pin dedicated for measuring the Proximal pressure received from the Manifold PCBA via the Main Board connector having below specifications: 1. PPROX – Analog Input – 0~4.096V max. |  |  | Hardware_HFOPCBA
FAB-EHFB-013 | HFO PIC Controller requirement | The HFO PIC Controller shall be capable of communicating with the NET DCU of the Main board using UART lines “UPDTX_2” and “UPDRX_2” having below specifications: 1. “UPDTX_2” – Transmitter signal at 5V logic 2. “UPDRX_2” – Receiver signal at 5V logic These lines are used for remote software upgrade. |  |  | Hardware_HFOPCBA
FAB-EHFB-014 | HFO PIC Controller requirement | The HFO PIC controller shall communicate with the Controller PIC of the Main board via SPI communication to receive the configuration of HFO frequency, I/E ratio and Amplitude. The SPI communication line shall have the below signal specifications: 1. “SCK” - SPI clock; Digital I/P at 5V logic 2. “HFSTAT” - SPI Data Out; Digital O/P at 5V logic 3. “MDATA” - SPI Data In; Digital I/P at 5V logic 4. “HF_CS” - SPI Slave select; Digital I/P at 5V logic The HFO Controller PIC will operate in configurable slave mode. |  |  | Hardware_HFOPCBA
FAB-EHFB-015 | HFO PIC Controller requirement | The HFO PIC controller shall receive “OVLD” status (active high) from Class D amplifier to determine the status of the Amplifier having the below signal specifications: 1. OVLD – Digital I/O at 5V logic |  |  | Hardware_HFOPCBA
FAB-EHFB-016 | HFO PIC Controller requirement | The HFO PIC controller shall have a Digital I/P at 5V logic to continuously monitor the Manifold 5V supply “VAVS_A”. |  |  | Hardware_HFOPCBA
FAB-EHFB-017 | HFO PIC Controller requirement | The HFO PIC controller shall enable / disable Class D amplifier output using “MUTE” signal of +5V logic level. |  |  | Hardware_HFOPCBA
FAB-EHFB-018 | HFO Horn Oscillation Requirements (Class D Amplifier) | The HFO PCBA shall have a Class D Amplifier for generating the High Frequency (HF) oscillations for the Mechanical Horn coil of TBD ohms. The coil will then be responsible for vibrating the Membrane which in turn generates the pneumatic oscillations. The Class D amplifier will have the below specifications: 1. Output Type – 2 channel Stereo 2. Bandwidth – (5Hz~20 Hz) ±0.5Hz 3. High Power Output – Minimum 30VA 4. Voltage supply – 10V ~ 16V 5. Signal Input – Differential |  |  | Hardware_HFOPCBA
FAB-EHFB-019 | HFO Horn Oscillation Requirements (Class D Amplifier) | The Class D bridge output supply shall have an inbuilt current limit of TBD A to limit any high current surge on the 24V line |  |  | Hardware_HFOPCBA
FAB-EHFB-020 | HFO Horn Oscillation Requirements (Class D Amplifier) | The Class D Amplifier output shall be connected in Parallel Bridge Tied Load (PBTL) configuration to drive load of TBD ohms (Horn coils) with a power of TBD watts The PBTL connections are made using 4 N- Channel Power MOSFETs. The MOSFETS characteristics are as below- 1. Rds ON minimum 60 m ohms 2. Vds = 30 V±10% 3. Id = 500mA minimum 4. Turn-on Delay time < 15 ns |  |  | Hardware_HFOPCBA
FAB-EHFB-021 | HFO Valve Control Requirements | The HFO PCBA shall include a HFO valve to transfer the High frequency oscillations to the pneumatic Output with the below specifications: 1. HFO Valve resistance - 35 ohm +/- 8% 2. Drive current – Minimum 0.35 A 3. Hold current - 0.08 A typical. |  |  | Hardware_HFOPCBA
FAB-EHFB-022 | HFO Valve Control Requirements | The HFO Valve operating power shall not exceed 4.5W@24 V operation |  |  | Hardware_HFOPCBA
FAB-EHFB-023 | HFO Valve Control Requirements | The HFO PIC controller shall generate Reset signals for 555 timer ICs that control HFO Valve. Reset signal keeps the HFO valve OFF during Power ON and when not in Use. |  |  | Hardware_HFOPCBA
FAB-EHFB-024 | HFO Valve Control Requirements | The HFO PIC controller shall generate Trigger signal “RB4”” for 555 timer ICs that control HFO Valve. Trigger Signal shall be used to Turn ON HFO Valve when required. |  |  | Hardware_HFOPCBA
FAB-EHFB-025 | HFO Valve Control Requirements | 555 timer shall also generate PWM timing signals based on R and C values to supply holding current with minimum power loss for HFO Valve during ON state. The frequency of the PWM signal shall be XXHz and the duty cycle shall vary from XX-YY |  |  | Hardware_HFOPCBA
FAB-EHFB-026 | HFO Valve Control Requirements | The Valve drive electronics shall include an LED for indicating the operation status of the Valve (for diagnostic purpose only) |  |  | Hardware_HFOPCBA
FAB-EHFB-027 | General Requirements | The HFO PCB shall have multiple layer design (4). There shall be at least 1 Ground plane and 1 power planes for noise immunity. |  |  | Hardware_HFOPCBA
FAB-EHFB-028 | General Requirements | The HFO PCBA shall satisfy the below mentioned manufacturing standards: 1. Flammability standard of UL94 V-0 2. ROHS Directive 2011/65/EU 3. REACH Directive (EG) Nr. 1907/2006 4. IPC-A-600 Class 2 – Acceptability of Printed Boards 5. IPC-SM-840 – Qualification & Performance spec for Permanent solder mask |  |  | Hardware_HFOPCBA
FAB-EHFB-029 | General Requirements | The HFO PCB shall be developed on FR4 material and have minimum 35u copper finish on each of the layers |  |  | Hardware_HFOPCBA
FAB-EHFB-030 | General Requirements | The HFO PCBA shall have clear and specific marking of all the mating connectors to different modules or sub systems |  |  | Hardware_HFOPCBA
FAB-EHFB-031 | General Requirements | The Mainboard shall have labels to define the revision no. of hardware and the software applicable to that version. |  |  | Hardware_HFOPCBA
FAB-EHFB-032 | Physical Requirements | The HFO PCBA shall have the maximum physical dimensions as per below specifications: 1. Length – 122mm 2. Width – 133mm 3. Thickness – 1.6mm |  |  | Hardware_HFOPCBA
FAB-EHFB-033 | Physical Requirements | The HFO PCBA shall have 3 NPTH Mounting holes of 3.2±0.1mm for mounting the card to the studs in enclosure. The mounting holes shall have routing clearance of minimum 0.2 mm around their circumference. |  |  | Hardware_HFOPCBA
FAB-EHFB-034 | Physical Requirements | The HFO PCBA shall be mounted with vibration-proof fasteners (for example, screws with spring washers or thread-lock). |  |  | Hardware_HFOPCBA
FAB-EMB-001 | SOM / NET DCU Interface | The mating header connector shall mate with the SOM / Net DCU module and have the below connector specifications: 1. Connector J1 & J3c – 8 Position (2X4) 2.54mm Pitch 2. Connector J3a – 6 Position (2X3) 2.54mm Pitch 3. Connector J2 – 14 Position (2X7) 2.54mm Pitch 4. Connector J4 – 16 Position (2X8) 2.54mm Pitch 5. Connector J5 & J7 – 26 Position (2X13) 2.54mm Pitch 6. Connector J3b – 34 Position (2X17) 2.54mm Pitch |  |  | Hardware_MainPCBA
FAB-EMB-002 | SOM / NET DCU Interface | All the mating header connectors for the SOM / Net DCU shall be of vertical mount to hold the module in stacked position with a mating of 9.49mm ±0.1mm (inclusive of mating length & insulation height) between the Mainboard PCBA and the bottom surface of the SOM / NETDCU module. |  |  | Hardware_MainPCBA
FAB-EMB-003 | SOM / NET DCU Interface | The SOM / Net DCU Module shall be having the below functional features for the Fabian HFO: 1. Operating system: Windows CE 6.0 or later 2. CPU: Intel PXA270, Architecture ARM9 3. Frequency: 520 MHz (minimum) 4. Flash 64MB, RAM 64MB (minimum) 5. Ethernet 10/100M 6. 3 x RS232, USB, I2C, SPI, Parallel Interface, Matrix Keyboard, HDMI 7. Resolution: SVGA (optional SXGA) 8. Digital I/O s – Minimum 14 (excluding the display, Keyboard & communication signal I/O s) |  |  | Hardware_MainPCBA
FAB-EMB-004 | SOM / NET DCU Interface | The SOM / Net DCU module shall be capable of operating at +10 to +40 °C |  |  | Hardware_MainPCBA
FAB-EMB-005 | SOM / NET DCU Interface | The current version of the SOM/ Net DCU to be used for Fabian HFO shall be of version 11, or version A9. Future versions may be considered upon verification. |  |  | Hardware_MainPCBA
FAB-EMB-006 | SOM / NET DCU Interface | The Main Board shall be capable of delivering 5V±5%; and minimum TBD mA of current for the Net DCU / SOM Module. |  |  | Hardware_MainPCBA
FAB-EMB-007 | Controller PIC Requirement | The Controller PIC shall have the following functional features: 1. Operating voltage - 5V±10% 2. Operating speed / Clock frequency (Min.) – 14.7456 MHz±5%. 3. Memory – Program Memory (64KB minimum); RAM (3.6KB minimum) 4. Analog Channels – Minimum 6 (of 12 Bit resolution) 5. Digital I/O s – Minimum 10 6. Other Interface requirements (Minimum quantity) – 1XPWM; 1XCCP; 1 UART & 1XSPI |  |  | Hardware_MainPCBA
FAB-EMB-008 | Controller PIC Requirement | The Controller PIC shall be capable of communicating with the SOM / Net DCU Module via UART communication (UPDRX_0/UPDTX_0). (This is used for exchange of sensor data and respective valve drives for both Manifold & Blender modules.) The UART signal specifications shall be as below: 1. UPDRX_0 – Receive signal at 5V Logic 2. UPDTX_0 – Transmit signal at 5V Logic |  |  | Hardware_MainPCBA
FAB-EMB-009 | Controller PIC Requirement | The Controller PIC section shall have a 12Bit DAC IC communicating through SPI, to convert the digital signal to PWM for PID regulation of the “Exhalation valve”. The DAC controlled by the SPI communication shall have a min. clock rate of “TBD Hz” |  |  | Hardware_MainPCBA
FAB-EMB-010 | Controller PIC Requirement | The Controller PIC shall have a PWM output “M_PWM” to connect with the “Monitoring PIC” (used for data synchronization). |  |  | Hardware_MainPCBA
FAB-EMB-011 | Controller PIC Requirement | The Controller PIC shall have an active low trigger pin/signal “TRIG” input from the “Monitoring PIC” (This is used to send the status of the Exhalation Valve through a Digital I/O pin “J/E”). |  |  | Hardware_MainPCBA
FAB-EMB-012 | Controller PIC Requirement | The Controller PIC shall have an active high pin/signal “J/E” output to the “Monitoring PIC as a status signal for the Exhalation Valve drive status |  |  | Hardware_MainPCBA
FAB-EMB-013 | Controller PIC Requirement | The Controller PIC shall act as a master for SPI communication to the HFO module with a minimum clock rate of TBD. The SPI signal specifications as below: 1. MCLK (Pin#2) – Clock signal, Type – Square wave clock min. XXMHz, Logic – 5V 2. MDATA (Pin#4) – MOSI signal; Type – Digital I/O; Logic 5V 3. HF_CS (Pin#5) – Chip Select output, Type, Digital O/P, Logic 5V 4. HF_SYNC (Pin#7) - MISO signal; Type – Digital I/O; Logic 5V (This is for the exchange of below mentioned data:) 1. HFO Frequency 2. I/E Ratio and 3. Amplitude |  |  | Hardware_MainPCBA
FAB-EMB-014 | Blender PCBA Interface Requirement | The Controller PIC shall have a 20-position connector (2X10) interface to connect with the Blender PCBA through a cable assembly. The on-board connector shall have the below specifications to mate with the cable assembly connector: 1. No. of Position – 20 (2X10) 2. Pitch – 2.54 mm 3. Style – Board to cable 4. Contact mating length – 6.10 ± 0.1mm 5. Flammability rating – UL94 V-0 6. Keying slot – Yes |  |  | Hardware_MainPCBA
FAB-EMB-015 | Blender PCBA Interface Requirement | The Blender PCBA connector interface with the Controller PIC shall have the below signal specifications: 1. AGND (Pin#1&5) – Analog Ground connection 2. VGND (Pin#10,11&12) – Virtual Ground connection 3. DGND (Pin#20)– Digital Ground connection 4. BL24V (Pin#7,8&9) – 24 Volts supply (For Valve drives) 5. A5V (Pin#14&15) – Analog 5V supply 6. D5V (Pin#13&19) – Digital 5V Supply 7. DRM1, DRM2 & DRM 3 (Pin#4,3&2) – Pressure Transducer signals; Type Analog; Value – 0~4.1V max 8. MI_DATA, MI_STAT, MCS & MCLK (Pin#17,6,16&18) – SPI lines to Blender PCBA, Type – Digital I/O, 5V logic |  |  | Hardware_MainPCBA
FAB-EMB-016 | Blender PCBA Interface Requirement | The Controller PIC shall communicate with the Blender PCBA via SPI communication medium and have the below mentioned signal specifications: 1. MI_DATA (Pin#17) – SPI MOSI signal at 5V logic 2. MI_STAT (Pin#6) – SPI MISO signal at 5V Logic 3. MCS (Pin#16) – SPI Chip select signal at 5V Logic 4. MCLK (Pin#18) – SPI clock at minimum XXMHz |  |  | Hardware_MainPCBA
FAB-EMB-017 | Blender PCBA Interface Requirement | The Controller PIC shall have 3 analog inputs (DRM1. DRM2 & DRM3) ranging from 0~3 V for the pressure transducer input connected via Blender PCBA connector interface |  |  | Hardware_MainPCBA
FAB-EMB-018 | Blender PCBA Interface Requirement | The Blender Connector interface having 24VDC supply shall be capable of handling min. TBD A of current. |  |  | Hardware_MainPCBA
FAB-EMB-019 | Patient Gas Supply / Manifold Board PCBA Interface | The Controller PIC shall have a 16-position connector (2X8) interface to connect with the Manifold PCBA through a cable assembly. The on-board connector shall have the below specifications to mate with the cable assembly connector: 1. No. of Position – 16 (2X8) 2. Pitch – 2.54 mm 3. Style – Board to cable 4. Contact mating length – 6.10 ±0.1mm 5. Flammability rating – UL94 V-0 6. Keying slot - Yes |  |  | Hardware_MainPCBA
FAB-EMB-020 | Patient Gas Supply / Manifold Board PCBA Interface | The Manifold PCBA connector interface with the Controller PIC shall have the below signal specifications: 1. AGND (Pin#1) – Analog Ground connection 2. VGND (Pin#14,15&16)) – Virtual Ground connection 3. BL24V (Pin#8) – 24 Volts supply 4. A5V (Pin#2)– Analog 5V supply 5. DRM 3 (Pin#4)– Pressure Transducer signals; Type Analog; Value – 0~4.1V max 6. BIAS & BIAS_H (Pin#5&6)– Electrical Safety Valve Signal, Type – Digital I/O, 5V logic 7. I_NEG & ENTL (Pin#7&9)– Inspiratory Solenoid Valve drive signal, Type – Digital I/O, 5V logic 8. PWM1 (Pin#10)– Exhalation valve drive signal; Type – Digital I/O, 5V logic |  |  | Hardware_MainPCBA
FAB-EMB-021 | Patient Gas Supply / Manifold Board PCBA Interface | The Manifold connector Interface shall have a varying voltage signal output “PWM1” (generated through DAC for the drive and control of the Exhalation Valve), ranging between “XXV-YYV” |  |  | Hardware_MainPCBA
FAB-EMB-022 | Patient Gas Supply / Manifold Board PCBA Interface | The Manifold connector interface shall have 2 digital outputs “BIAS & BIAS_H” at 5V logic to control the “Outlet Switching Relay” |  |  | Hardware_MainPCBA
FAB-EMB-023 | Patient Gas Supply / Manifold Board PCBA Interface | The Manifold connector interface shall have 1 analog input “DRM3” (of range 0~4.1V) for the “Inspiratory Pressure Transducer” |  |  | Hardware_MainPCBA
FAB-EMB-024 | Patient Gas Supply / Manifold Board PCBA Interface | The Manifold connector interface shall have 1 analog input “PROX” (of range 0~4.1V) for the “Proximal Pressure sensing” |  |  | Hardware_MainPCBA
FAB-EMB-025 | Patient Gas Supply / Manifold Board PCBA Interface | The Manifold Connector interface having the 24VDC supply shall be capable of handling min. TBD A of current. |  |  | Hardware_MainPCBA
FAB-EMB-026 | HFO Module | The Controller PIC circuit shall have 26 position connector (2X13) for HFO PCBA connectivity through a cable assembly. The on-board connector shall have the below specifications to mate with the cable assembly connector: 1. No. Of Position – 26 (2X13) 2. Pitch – 2.54 mm 3. Style – Board to cable 4. Contact mating length – 6.10 ±0.1mm 5. Flammability rating – UL94 V-0 6. Keying slot - Yes |  |  | Hardware_MainPCBA
FAB-EMB-027 | HFO Module | The HFO PCBA interface connector shall have the below signal specifications: 1. PROX_A (Pin#1) – Analog input Pin ranging from 0~4.1V 2. MCLK, MDATA, HF_CS & HF_SYNC (Pin#2,4,5&7)– SPI Communication Lines for the HFO PCBA 3. AVS5V/VAVS_A (Pin#3) – 5V Supply of Manifold that is monitored by the HFO PCBA 4. +5VHF (Pin#6)– 5V±10% supply for the HFO PCBA 5. DGND (Pin#8)– Ground Pin 6. HGND (Pin#9-15) – Power Ground for the HFO PCBA 7. 24V (Pin#16-22) – 24V±10% supply to HFO PCBA 8. Test Points (Pin#23&24) – Test Points 9. UPDRX_2 (Pin#25) – UART Receive signal at 5V Logic 10. UPDTX_2 (Pin#26) – UART Transmit Signal at 5V Logic (The details of SPI Signals are mentioned in the below req. in HFO section) |  |  | Hardware_MainPCBA
FAB-EMB-028 | HFO Module | The current consumption of the HFO module 24VDC supply (while operating in the HFO Mode) shall not exceed 500mA. |  |  | Hardware_MainPCBA
FAB-EMB-029 | Miscellaneous Requirements of PIC Controller | The Controller PIC shall have a dedicated “Watchdog” pin “WD1” (Active High of 5V Logic) with a timeout period between 2.2mS to 20mS, output to the Net DCU (for detecting the system failure of the Controller PIC) |  |  | Hardware_MainPCBA
FAB-EMB-030 | Miscellaneous Requirements of PIC Controller | The Controller PIC shall have a programming header of 6 pins that gets connected to the VCC, GND and 3 Programming pins of the Controller. The signal pinouts for programming are: 1. MCLR (Pin#1) – Programming pin; RESET signal with 10K pull-up to VDD 2. A5V (Pin#2) – Supply voltage 3. AGND (Pin#3) – Ground pin 4. PGD (Pin#4) – Programming pin; Data signal at 5V logic 5. PGC (Pin#5) - Programming pin; Clock signal at 5V logic 6. NC (Pin#6) – Not Connected. Note: No Capacitors, pullups or diodes are recommended on the PGC & PGD pins. |  |  | Hardware_MainPCBA
FAB-EMB-031 | Miscellaneous Requirements of PIC Controller | The Controller PIC Programming header connector shall have the below specifications: 1. No. Of Position – 6 (1X6) 2. Pitch – 2.54 mm 3. Style – Board to board or cable 4. Contact mating length – 6.70 ±0.1mm 5. Flammability rating – UL94 V-0 |  |  | Hardware_MainPCBA
FAB-EMB-032 | Miscellaneous Requirements of PIC Controller | The Controller PIC shall have at least 1 dedicated pin to source (active high) 20mA at 5V for an LED used for diagnostics. |  |  | Hardware_MainPCBA
FAB-EMB-033 | Miscellaneous Requirements of PIC Controller | The Controller PIC circuit shall have provision of a 10 Pin “Test Connector” interface to physically test all the below mentioned vital signal manually or through a test jig: 1. AGND (Pin#1) - Analog Ground 2. DGND (Pin#2) - Digital ground 3. PROX (Pin#3) - Proximal Pressure sensor signal. Analog signal (0~4.1V) 4. ENTL (Pin#4) - Inspiratory solenoid drive signal at 5V Logic) 5. DRM1 (Pin#5) - Inspiratory pressure transducer signal; Analog signal (0~4.1V) 6. PWM1 (Pin#6) - Exhalation valve drive signal; (signal of frequency TBD and Duty Cycle TBD) 7. DRM2 (Pin#7) - Inspiratory pressure transducer signal; Analog signal (0~4.1V) 8. LED (Pin#8) – Logic High 5V for LED supply check 9. DRM3 (Pin#9) - Inspiratory pressure transducer signal; Analog signal (0~4.1V) 10. TSTSR (Pin#10) - 24VDC power switch bias Signal at 5V Logic |  |  | Hardware_MainPCBA
FAB-EMB-034 | Miscellaneous Requirements of PIC Controller | The Test connector shall be a 10-position connector (2X5) and have the below specifications: 1. No. Of Position – 10 (2X5) 2. Pitch – 2.54 mm 3. Style – Board to Board or cable 4. Type – Header (Unshrouded) 5. Contact mating length – 6.7 ±0.1mm 6. Flammability rating – UL94 V-0 |  |  | Hardware_MainPCBA
FAB-EMB-035 | Monitoring PIC Requirement | The Monitoring PIC shall have the following functional features for the Fabian HFO: 1. Operating voltage – 3.3VDC ±5% 2. Operating speed– 40 MIPS 3. Memory – Program Memory (128KB); RAM (16KB) 4. Analog Channels – Min 10 (of 12 Bit) 5. Digital I/O s – Min. 15 6. Other Interface requirements (Min. Nos.) – 1XPWM; 1XCCP; 2 UART; 2XSPI & 2XI2C |  |  | Hardware_MainPCBA
FAB-EMB-036 | Monitoring PIC Requirement | The Monitoring PIC shall have at least 1 dedicated pin to source (active high) 20mA at 3V3 for an LED used for diagnostics. |  |  | Hardware_MainPCBA
FAB-EMB-036 | Monitoring PIC Requirement | The Monitoring PIC shall be capable of communicating with the SOM / Net DCU Module via SPI protocol (e.g. for receiving the ventilator settings and transfer back the Proximal Pressure value and ventilator state to the SOM / Net DCU Module). |  |  | Hardware_MainPCBA
FAB-EMB-037 | Monitoring PIC Requirement | The Monitoring PIC shall act as a slave and shall have the below SPI signal specifications: 1. SCK – SPI clock signal having min. clock rate of XXHz at 3V3 logic 2. MOSI – SPI MOSI Signal at 3V3 logic 3. MISO - SPI MISO Signal at 3V3 logic & 4. FRM - SPI Chip select Signal at 3V3 logic |  |  | Hardware_MainPCBA
FAB-EMB-038 | Monitoring PIC Requirement | The Monitoring PIC shall have a dedicated 4 pin Debug interface Header with below signal specifications: 1. AGND (Pin#4) – Ground Pin 2. A5V (Pin#1) – 5V Supply 3. U1RX (Pin#3) – UART Receive pin at 3V3 logic. 4. U1TX (Pin#4) - UART Transmit pin at 3V3 logic. |  |  | Hardware_MainPCBA
FAB-EMB-039 | Monitoring PIC Requirement | The 4 pin Debug Interface connector shall have the below specifications to mate with the Debug cable connector: 1. No. Of Position – 4 (1X4poles) 2. Pitch – 2.54 mm 3. Style – Board to Board or cable 4. Contact mating length – 6.7 ±0.1mmFlammability rating – UL94 V-0 |  |  | Hardware_MainPCBA
FAB-EMB-040 | Monitoring PIC Requirement | The Monitoring PIC shall have a programming header of 6 pins with the signals as below: 1. PCLR# (Pin#1) – Programming pin; RESET signal with 10K pull-up to VDD 2. A3V3 (Pin#2) – Supply voltage 3V3 3. AGND (Pin#3) – Ground pin 4. PGED2 (Pin#4) – Programming pin; Data signal at 3V3 logic 5. PGEC2 (Pin#5) - Programming pin; Clock signal at 3V3 logic 6. NC (Pin#6) – Not Connected. Note: Pull downs of 10K are recommended on the PGEC2 & PGED2 Pins |  |  | Hardware_MainPCBA
FAB-EMB-041 | Monitoring PIC Requirement | The Monitoring PIC Programming header connector shall have the below specifications to mate with the Programming cable connector: 1. No. Of Position – 6 (1X6poles) 2. Pitch – 2.54 mm 3. Style – Board to board or cable 4. Contact mating length – 6.70 ±0.1mm 5. Flammability rating – UL94 V-0 |  |  | Hardware_MainPCBA
FAB-EMB-042 | Monitoring PIC Requirement | The Monitoring PIC shall have an active high (3.3V) output pin/signal “TRIG” to the “Controller PIC” (for Synchronization). |  |  | Hardware_MainPCBA
FAB-EMB-043 | Monitoring PIC Requirement | The Monitoring PIC shall have a dedicated input pin for FAN operation check that is capable to detect a frequency between 60Hz~210Hz. (Note: This signal is the Tachometer O/P, square wave pulses, of the fan and shall be limited via 5V TVS.) The Fan operation shall be supervised by the Monitor PIC by a speed signal with an upper limit (speed > 190Hz OR 5700 RPM) and a lower limit (speed < 80Hz OR 2400 RPM). This signal shall be measured every 100ms |  |  | Hardware_MainPCBA
FAB-EMB-044 | Monitoring PIC Requirement | The Monitoring PIC shall have 12 Bit analog Inputs for sensing the below mentioned voltage levels by high precision Voltage dividers: 1. 24V of Blender Module – “BL24” (sensing range 0~2.2V) 2. 5V of Manifold Module – “AVS5V” (Sensing Range (0~3V) 3. 5V of VCC – “A5V” (Sensing Range (0~1.2V) |  |  | Hardware_MainPCBA
FAB-EMB-045 | Monitoring PIC Requirement | The Monitoring PIC shall have 12 Bit analog Inputs for sensing the below mentioned analog signals: 1. Proximal Pressure – “PROX_A” (Sensing Range (0~3V) 2. Inspiratory pressure from Manifold module – “DRM3_A” (Sensing Range (0~3V) |  |  | Hardware_MainPCBA
FAB-EMB-046 | Monitoring PIC Requirement | The Monitoring PIC shall have a PWM output “M_PWM” to connect with the “Controller PIC” (for data synchronization). |  |  | Hardware_MainPCBA
FAB-EMB-047 | Monitoring PIC Requirement | The Monitoring PIC shall have two Digital outputs at 3V3 Logic, “SHUTDN & P/N” to feed an Analog SPDT switch for protection of the flow sensor. (e.g. To protect the sensor wire in case of excessive contamination the bridge can be shutdown at terminal “SHUTDN” by generating an offset, which sets the bridge supply voltage to Zero.) |  |  | Hardware_MainPCBA
FAB-EMB-048 | Monitoring PIC Requirement | The Monitoring PIC shall have a dedicated Watchdog pin “WD2” (Active High of 3V3 Logic) with a timeout period between 2.2mS to 20mS, output to the Net DCU (for detecting the system failure of the Monitoring PIC). |  |  | Hardware_MainPCBA
FAB-EMB-049 | Monitoring PIC Requirement | The Monitoring PIC shall have one Digital input “SPKRSTAT” (active low signal) to check the fault status of the audio amplifier for the Speaker. |  |  | Hardware_MainPCBA
FAB-EMB-050 | Monitoring PIC Requirement | The Monitoring PIC shall have one I2C Channel “SDA_M & SCL_M” at 3V3 logic to connect with the Oxygen sensor as an optional I2C Interface. |  |  | Hardware_MainPCBA
FAB-EMB-051 | Monitoring PIC Requirement | The Monitoring PIC shall have a dedicated I2C Channel “SDA & SCL” at 3V3 logic for FRAM. (The FRAM will store calibration data, which persists through power cycles, for the following components: - Inspiratory Pressure Transducer - Proximal Pressure Transducer - Blender Pressure Regulators - Exhalation valve - HWFS) |  |  | Hardware_MainPCBA
FAB-EMB-052 | Monitoring PIC Requirement | The calibration FRAM shall have the below specifications: 1. Memory Size – 4Kb (512X8) 2. Interface – I2C 3. Supply Voltage – 3.3V ±5% |  |  | Hardware_MainPCBA
FAB-EMB-053 | Flow Sensor Requirement Under Monitoring PIC | The Monitoring PIC section shave have one 10 Pole connector (2x5) for the Hot wire Flow sensor (HWFS) Input. The HWFS connector interface shall have the below specifications: 1. No. of Position – 10 (2X5) 2. Pitch – 2.54 mm 3. Style – Board to cable 4. Contact mating length – 6.10 ±0.1mm 5. Flammability rating – UL94 V-0 |  |  | Hardware_MainPCBA
FAB-EMB-054 | Flow Sensor Requirement Under Monitoring PIC | The Hot Wire Flow sensor shall have the below mentioned signal specifications: 1. HW1A (Pin#1&2) – HW flow sensor 1 I/P sig. 2. HW1B (Pin#3&4) – HW flow sensor 1 I/P sig. 3. HW2A (Pin#5&6) – HW flow sensor 2 I/P sig. 4. HW2B (Pin#7&8) – HW flow sensor 2 I/P sig. 5. AGND (Pin#9&10) - GND |  |  | Hardware_MainPCBA
FAB-EMB-055 | Flow Sensor Requirement Under Monitoring PIC | The Bridge circuit shall be designed for raw HWFS signals ranging from XXmV to yymV. (Note: These flow sensor signals are amplified and fed to the Monitoring PIC (as “FLOW1” & “FLOW2”) and should not exceed the range of 3V3.) |  |  | Hardware_MainPCBA
FAB-EMB-056 | Flow Sensor Requirement Under Monitoring PIC | The Monitoring PIC shall have one SPI interface having signals “DAC_C, DAC_D & DAC_CS#” that connect with a 12Bit DAC used as a digitally controlled potentiometer. |  |  | Hardware_MainPCBA
FAB-EMB-057 | Flow Sensor Requirement Under Monitoring PIC | A 12-bit DAC shall function as a digitally controlled potentiometer with an output voltage “VOUTB” range of xxV~yyV. Its reference input “VREFB” is driven by the bridge supply voltage while its output voltage “VOUTB” can tune the bridge by ±22 % through a resistor. |  |  | Hardware_MainPCBA
FAB-EMB-058 | Oxygen Sensor Requirement Under Monitoring PIC | The Monitoring PIC Section shall have 3 pin connector interface for the Oxygen sensor Input. The interface connector specifications are as below: 1. No. of poles – 3 (1X3) 2. Pitch – 2.49mm 3. With mechanical Locking |  |  | Hardware_MainPCBA
FAB-EMB-059 | Oxygen Sensor Requirement Under Monitoring PIC | The Oxygen Sensor connector interface shall have the below signal specifications: 1. OS2 (Pin#1) – Analog signal (9mv~13.5mV for an oxygen concentration of 21% in Ambient air and 42.85mV~64.28mV for an Oxygen concentration of 100%) 2. AGND – GND pin 3. OS2 (Pin#1) - Analog signal (9mv~13.5mV for an oxygen concentration of 21% in Ambient air and 42.85mV~64.28mV for an Oxygen concentration of 100%) |  |  | Hardware_MainPCBA
FAB-EMB-060 | Oxygen Sensor Requirement Under Monitoring PIC | The oxygen sensor shall have accuracy ± 3 % after proper calibration |  |  | Hardware_MainPCBA
FAB-EMB-061 | Watchdog & Alarms PIC Requirement | The Alarms PIC shall have the following functional features for the Fabian HFO: 1. Operating voltage – 5VDC ±10% 2. Operating speed (Max) – 32 MHz 3. Memory – Program Memory (3.5KB); RAM (8KB) 4. Digital I/O s – Min. 10 5. Other Interface requirements (Min. Nos.) –1 UART & 1XI2C |  |  | Hardware_MainPCBA
FAB-EMB-062 | Watchdog & Alarms PIC Requirement | The Alarms PIC shall be capable of communicating with the SOM / Net DCU Module via UART communication UPDRX_3 and UPDTX_3. |  |  | Hardware_MainPCBA
FAB-EMB-063 | Watchdog & Alarms PIC Requirement | The Alarms PIC shall be receiving the alarms signals from the SOM / Net DCU module via 6 digital I/Os to be labelled from AL0-AL5 (via a D-Type transparent Latch). (Note: The 6 digital I/Os for Alarm signals should be capable of handling maximum 32 (2^6) different types of alarm signals.) |  |  | Hardware_MainPCBA
FAB-EMB-064 | Watchdog & Alarms PIC Requirement | The Alarms PIC shall have a programming header of 6 pins that gets connected to the VCC, GND and 3 Programming pins of the Controller. The Signal specifications for programming are as below: 1. MCLR (Pin#1) – Programming pin; RESET signal with 10K pull-up to VDD 2. AL5V (Pin#2) – Supply voltage 3V3 3. DGND (Pin#3) – Ground pin 4. ICSPDAT (Pin#4) – Programming pin; Data signal at 5V logic 5. ICSPCLK (Pin#5) - Programming pin; Clock signal at 5V logic 6. NC (Pin#6) – Not Connected. (Note: No Pull ups or Pull downs are recommended on the ICSPDAT & ICSPCLK Pins) |  |  | Hardware_MainPCBA
FAB-EMB-065 | Watchdog & Alarms PIC Requirement | The Alarms PIC shall have a dedicated Watchdog pin “WD3” (Active High of 5V Logic) with a timeout period between 2.2mS to 20mS, output to the Net DCU, (for detecting the system failure of the Alarms PIC). |  |  | Hardware_MainPCBA
FAB-EMB-066 | Speaker Drive Requirements | The Alarms PIC shall have one dedicated I2C channel to communicate and control the Audio power amplifier for alarms generation. |  |  | Hardware_MainPCBA
FAB-EMB-067 | Speaker Drive Requirements | The Audio Power Amplifier for speaker drive shall have the below functional features: 1. Volume control > 3 steps, within range 49 to 65 dB(A). 2. Drive capacity – Minimum 4W (@4 Ohm Load) 3. Shall have low power shutdown mode and low quiescent current (max. 8mA) 4. Shall have fault condition feedback to Microcontroller 5. Shall have function to “Mute” Note: The Audio amplifier will be connected with an external speaker having the below specifications: 1. Rated Power – 4W 2. Nominal Impedance – 4 Ohm 3. Mean sound pressure level – Max. 52dB(A) in normal operation 4. Resonance frequency – 250 Hz |  |  | Hardware_MainPCBA
FAB-EMB-068 | Speaker Drive Requirements | The Alarms PIC shall have 2 pin connector (Male header) of 2.54mm Pitch for speaker connection. |  |  | Hardware_MainPCBA
FAB-EMB-069 | Watchdog Requirement | All the Watchdog status signals (WDST0-WDST3) shall have LED indications as per below specifications: 1. LED color – Green (diffused) The Watchdog LEDs shall be driven by Digital I/O pins of Alarms PIC at 5V Logic. |  |  | Hardware_MainPCBA
FAB-EMB-070 | Watchdog Requirement | There Alarms PIC shall generate one signal of Watchdog alarm “WDAL” to drive the Safety Relay that generates the 5V Supply for the Manifold. (The relay will switch the 5V supply “AV5V” to “AVS5V” to Manifold module and send a sensing signal to the HFO Module.) |  |  | Hardware_MainPCBA
FAB-EMB-071 | Watchdog Requirement | The Specification of the safety Relay shall be as below: 1. Coil Voltage – 5V 2. Contact Current – minimum TBD A 3. Contact form – DPDT (2 form C) |  |  | Hardware_MainPCBA
FAB-EMB-072 | Miscellaneous Requirements for Mainboard | The Mainboard PCBA shall have interface connectors for the “ETCo2 and SPO2” PCBA inputs. The ETCO2 & SPO2 connector interface on the Mainboard PCBA shall have the below specifications: 1. No. Of Position – 14 (2X7) 2. Pitch – 2.54 mm 3. Style – Board to cable 4. Contact mating length – 6.10 ±0.1mm 5. Keying Slot - Yes |  |  | Hardware_MainPCBA
FAB-EMB-073 | Miscellaneous Requirements for Mainboard | The ETCO2 Interface shall communicate via UART and have the below mentioned signal specifications: 1. +5V (Pin#1&2) – 5V Supply for the ETCO2 sensor 2. DGND (Pin#3,7,8,9,10&11)) – GND pin 3. Test Point/STAT(Pin#4) – Test Point or Status signal 4. TXD3 (Pin#5) – UART Transmitter 5. RXD3 (Pin#6) – UART receiver signal |  |  | Hardware_MainPCBA
FAB-EMB-074 | Miscellaneous Requirements for Mainboard | The SPO2 Interface shall communicate via USB and have the below mentioned signal specifications: 1. VBUS (Pin#14) – 5V Supply for the USB Bus 2. DGND (Pin#7,8,9,10&11) – GND pin 3. USBDOWN3-2P (Pin#12) – USB+ Signal 4. USBDOWN3-2N (Pin#13) – USB- Signal |  |  | Hardware_MainPCBA
FAB-EMB-075 | Miscellaneous Requirements for Mainboard | The Mainboard PCBA shall have a 6 pin UDP connector for front panel rotary encoder (including push switch) interface with the Net DCU. The specification of the UDP connector is as below: 1. Connector system – Board to board 2. No. of positions – 6 6. Connector pitch – 1.26mm |  |  | Hardware_MainPCBA
FAB-EMB-076 | Miscellaneous Requirements for Mainboard | The 6 pin UDP connector shall have the below signal specifications: 1. DGND (Pin#1&3) – Ground potential 2. NO (Pin#2) – Push Button Switch signal 3. Output B (Pin#4) – Open collector O/P pulled up with VDD (3V3) 4. Output A (Pin#5) – Open collector O/P pulled up with VDD (3V3) 5. POWER (Pin#6) – Encoder supply pin |  |  | Hardware_MainPCBA
FAB-EMB-077 | Miscellaneous Requirements for Mainboard | The Mainboard shall contain RTC battery holder to hold the Coin cell. (This is meant for providing continuous supply of 3V to the RTC section which maintains the time stamp of all events.) The specification of the holder shall be as below: 1. Battery size accepted – 2032 2. Battery holder mount – Through hole 3. Flammability rating UL 94 V-0 1. Safety against inverse polarity (polarized) |  |  | Hardware_MainPCBA
FAB-EMB-078 | Miscellaneous Requirements for Mainboard | The Mainboard shall contain a 40 position (2X20) connector interface from Power card (Akku Board) (for all the required power supplies) with the following specifications: 1. Connector type – Male header 2. Mating pitch – 2.54mm 3. Connector style – board to cable 4. Contact mating length – 6.10 ±0.1mm 4. Flammability standard – UL94 V-0 5. Current rating per contact – minimum TBD A |  |  | Hardware_MainPCBA
FAB-EMB-079 | Miscellaneous Requirements for Mainboard | The Power Board interface to the Mainboard PCBA shall have the below mentioned signal specifications: 1. AL5V (Pin#7,8,9&10) – 5V supply for Analog circuits 2. D5V (Pin#11~14) – 5V Supply for Digital circuits 3. DGND (Pin#17~20) – Digital Ground 4. VGND (Pin#21&22) – Virtual Ground 5. HGND (Pin#23~30) – HFO Module Power Ground 6. +12V (Pin#15&16) – 12 V supply 7. 24V (Pin#31~36) – 24V Supply 8. P100.SCL_A (Pin#40) & P100.SDA_A (Pin#39) – I2C signals from Net DCU (after 5V logic level shifting) 9. P100.AK_3 (Pin#37) – 5V bias signal for N-Channel high side switch that switches 24 volts 10. TP101 (Pin#3) & TP121 (Pin#38) – Test Points (TP121 is pulled down) 11. P100.KON (Pin#5) & P100.KON1 (Pin#6) – 5V Digital Logic signal for the matrix keyboard 12. P100.LED_EXT (Pin#4) – 5V LED Supply for the Matrix (membrane) keyboard after getting filtered through Ferrite bead 13. P100.LED_Netz (Pin#2)– 5V LED Supply for the Matrix (membrane) keyboard after getting filtered through TVS diode 5. P100.LP5_2 (Pin#1)– 5V Digital I/O for the Matrix keyboard |  |  | Hardware_MainPCBA
FAB-EMB-080 | Miscellaneous Requirements for Mainboard | The Power board (Akku board) shall have a dedicated pin for an external LED indication for the “System Fail” alarms. This LED Signal shall be active Low and routed through the “Membrane keypad Interface”. (e.g. The System Fail LED shall get activated in case of a system fail caused by the Hardware watchdog circuit, Blender failure, Fan failure or exhalation valve calibration failure.) |  |  | Hardware_MainPCBA
FAB-EMB-081 | Miscellaneous Requirements for Mainboard | The Mainboard PCBA shall have an interface to a FRAM (for storing the configuration data of the ventilator) with the below specifications: 1. Memory Type – FRAM (Non-Volatile) 2. Memory size – min. 1Mb 3. Interface – I2C 4. Supply voltage – 3.3V ±5% |  |  | Hardware_MainPCBA
FAB-EMB-082 | Miscellaneous Requirements for Mainboard | The Mainboard PCBA shall have a connector interface for the Membrane keypad at the front panel. The Membrane connector interface shall have the below specifications: 1. No. of positions – 13 (1X13) 2. Pitch – 2.54mm 3. Fastening type – Push pull 4. Flammability rating – UL94 V-0 |  |  | Hardware_MainPCBA
FAB-EMB-083 | Miscellaneous Requirements for Mainboard | The Membrane Keyboard shall have the below mentioned Signal Specifications: 1. MX2 (Pin#9) & MX3(Pin#4) – Digital I/O of 5V logic for Row Inputs 2. MY0~MY3 (Pin#5~8) - Digital I/O of 5V logic for Column Inputs 3. LED_E2 (Pin#12) & P100.LED_Netz (Pin#13) – 5V LED supply for the Membrane Keyboard 4. LP4 (Pin#10) & P100.LP5_2 (Pin#11) – Digital I/O of 5V logic to the Membrane keyboard 5. KAS (Pin#2) & P100.KON (Pin#3) – Digital I/O of 5V Logic from the Membrane Keyboard |  |  | Hardware_MainPCBA
FAB-EMB-084 | Miscellaneous Requirements for Mainboard | The SOM / Net DCU module of the Mainboard PCBA shall interface with the front panel buttons via matrix key positioning using 2 rows (MX2-MX3) & 4 columns (MY0-MY3) architecture. (This is to address the 8 user input buttons and have 2 additional buttons for power and alarms silencing.) |  |  | Hardware_MainPCBA
FAB-EMB-085 | Miscellaneous Requirements for Mainboard | The Mainboard PCBA shall have an altitude sensor for measurement of operating altitudes of the ventilator with the below mentioned specifications: 1. Operating pressure – 70~106kPa 2. Interface – I2C 3. Accuracy - ±2kPa (Note: The accuracy is defined as per altitude variation of minimum 500 meters between 2 levels) |  |  | Hardware_MainPCBA
FAB-EMB-086 | Miscellaneous Requirements for Mainboard | The Mainboard PCBA shall have 3 position connector for the nurse call cable assembly interface. The nurse call interface must be isolated to satisfy regulatory safety requirements per IEC 60601-1. The Nurse call header interface has the below specifications: 1. No. of positions – 3 (1X3) 2. Pin Pitch – 2.54mm 3. Flammability rating – UL94 V-0 4. Per Contact current – Minimum 1A |  |  | Hardware_MainPCBA
FAB-EMB-087 | Miscellaneous Requirements for Mainboard |  |  |  | Hardware_MainPCBA
FAB-EMB-088 | Miscellaneous Requirements for Mainboard
FAB-EMB-089 | Miscellaneous Requirements for Mainboard
FAB-EMB-090
FAB-EMB-091
FAB-EMB-092

## Software_Sub
Index | Requirement Title | Inputs | Failure_Modes | Hazard | Source
BLAPIC_SRS_001 | Commands from UIM | The software shall receive following commands from UIM via UART. • Program • Erase • Hello • Reset |  |  | Alarm PIC Bootloader
BLAPIC_SRS_002 | Response to UIM | The software shall send the following response to UIM via UART. • Acknowledgement, If program command successful • Checksum error, if checksum error occurred • Read/write error, if read/write unsuccessful • Address error, If address error • Size error, If size error. |  |  | Alarm PIC Bootloader
BLAPIC_SRS_003 | Erase memory | The software shall erase the EEPROM data, when the erase command received via UART. |  |  | Alarm PIC Bootloader
BLAPIC_SRS_004 | Alarm PIC Bootloader | The software shall load Alarm PIC software into EEPROM memory, when program command received via UART. |  |  | Alarm PIC Bootloader
BLAPIC_SRS_005 | Check existing application-Reset | The software shall perform soft reset, when reset command received via UART. |  |  | Alarm PIC Bootloader
BLAPIC_SRS_006 | Bootloader version | The software shall return bootloader version, when hello command received via UART. |  |  | Alarm PIC Bootloader
BLAPIC_SRS_007 | Checksum check | The software shall perform checksum operation to ensure integrity of data received via UART. |  |  | Alarm PIC Bootloader
APIC_SRS_001 | Low priority pattern | The software shall have low priority alarm according to IEC 60601-1-8. Tone pattern consists of one single pulse. Attribute Requested value by regulation Rise/Fall time 10-20% of effective pulse length effective pulse length 125-250ms pause after pulse groups >15s |  |  | Alarm PIC
APIC_SRS_002 | Medium priority pattern | The software shall have medium priority alarm according to IEC 60601-1-8. Tone pattern consists of three pulses. Attribute Requested value by regulation Rise/Fall time 10-20% of effective pulse length effective pulse length 125-250ms pause Y (after 1st/2nd/3rd pulse) 125-250ms pause after pulse groups 2.5-30s |  |  | Alarm PIC
APIC_SRS_003 | High priority pattern | The software shall have high priority alarm pattern according to IEC 60601-1-8. Tone pattern consists of ten pulses. Attribute Requested value by regulation Rise/Fall time 10-20% of effective pulse length effective pulse length 75-200ms pause x (after 1st/2nd/4th/6th/9th 50-125ms pulse) pause after 3rd & 8th pulse 2*(pause x)+(effective pulse length) (±5%) pause after 5th pulse 350ms-1.3s pause after pulse groups 4s -5s |  |  | Alarm PIC
APIC_SRS_004 | Tone pulse rise- and fall time | The software shall increase/decrease the tone volume at the start/end of each pulse to meet slope requirements specified in IEC 60601-1-8. Rise and fall time of each tone must be between 10% and 20% of the effective tone pulse length. |  |  | Alarm PIC
APIC_SRS_005 | Speaker output | The software shall set volume and operation mode of the connected audio amplifier via l2C. The output tone is generated via a PWM-signal that is hardware filtered, so that the tone frequencies required by IEC60601-1-8 are output. |  |  | Alarm PIC
APIC_SRS_006 | Low volume output | When both LSB and MSB volume inputs are low, the software shall set all alarms with low volume (Amplifier volume level 7) except system alarm. |  |  | Alarm PIC
APIC_SRS_007 | Medium volume output | When only the LSB volume input is high, the software shall set all alarms with medium volume (Amplifier volume level 12) except system alarm. |  |  | Alarm PIC
APIC_SRS_008 | High volume output | When MSB volume input is high (LSB volume input is don't care), the software shall set all alarms with high volume (Amplifier volume level 17). |  |  | Alarm PIC
APIC_SRS_009 | Audible alarm off | The software shall audibly silence the alarm for a period not exceeding 120 seconds when the audio is paused (pressing alarm silence button). |  |  | Alarm PIC
APIC_SRS_010 | Device shutdown detection | The software shall set shutdown alarm as an output when shutdown input gets low. The shutdown alarm must remain active no matter of the parallel input stages. It is only reset by user or a valid reboot of the NetDCU. |  |  | Alarm PIC
APIC_SRS_011 | Shutdown alarm | The software shall set shutdown alarm to be active until the user performs a reset by pressing "Alarm reset" and holding it for two seconds after device is fully powered down. |  |  | Alarm PIC
APIC_SRS_012 | Reboot detection while shutdown alarm is active | The software shall detect the device reboot when the shutdown alarm is not reset by the user. After a valid boot is detected the shutdown alarm must be reset and alarms are regularly output. |  |  | Alarm PIC
APIC_SRS_013 | System alarm volume | The software shall provide high volume to the system alarms irrespective of what volume input levels are set. (Amplifier volume level 17). |  |  | Alarm PIC
APIC_SRS_014 | No priority | The software shall not activate any alarm if both priority inputs (LSB and MSB) are low, except when the hardware watchdog input is low. |  |  | Alarm PIC
APIC_SRS_015 | Low priority | The software shall provide a low priority alarm if only the LSB priority input is high, except when the hardware watchdog input is low. |  |  | Alarm PIC
APIC_SRS_016 | Medium priority | The software shall provide a medium priority alarm if only the MSB priority input is high, except when the hardware watchdog input is low. |  |  | Alarm PIC
APIC_SRS_017 | High priority | The software shall provide a high priority alarm when both priority inputs (LSB and MSB) are high. |  |  | Alarm PIC
APIC_SRS_018 | Hardware watchdog priority | The software shall provide a high priority alarm with maximum volume when the hardware watchdog input is low, except the Alarm PIC is waiting for the NetDCU to boot. No matter what alarm priority is requested via the priority input. |  |  | Alarm PIC
APIC_SRS_019 | System alarms | When the alarms triggered by the hardware watchdog input, the software shall enable the alarm LED, during this period a tone is output via the speaker. |  |  | Alarm PIC
APIC_SRS_020 | Test clicks | After boot detection the device, the software shall output one or two test clicks so that the user can recognize a speaker defect. During this test click output, the alarm LED must be illuminated. |  |  | Alarm PIC
APIC_SRS_021 | Volume input | The software shall set output alarm with user selected volume with an exception of system-alarm and after delay of the power-off alarm, which are always high volume. |  |  | Alarm PIC
APIC_SRS_022 | Alarm LED output | The software shall handle an alarm LED via regular IO output (Port B6). |  |  | Alarm PIC
APIC_SRS_023 | Shutdown alarm volume | The software shall initiate the shutdown alarm with the user selected volume and then increase the volume to maximum in less than one minute. |  |  | Alarm PIC
APIC_SRS_024 | Parallel interface to NetDCU | The software shall read the desired alarm request communicated by NetDCU via a parallel IO connection. Signal Processor port Description AL_0 RA0 Priority LSB AL_1 RA1 Priority MSB AL_2 RA2 Volume LSB AL_3 RA3 Volume MSB AL_4 RA4 Click on toggle AL_5 RA6 Device shutdown Additionally, the Alarm PIC receives the common hardware watchdog signal (that turns low on failure of at least one processor) |  |  | Alarm PIC
APIC_SRS_025 | Boot detection | The software shall hold back the alarm during NetDCU boot. When system alarm goes low and click toggle is detected, regular operation shall be established. |  |  | Alarm PIC
APIC_SRS_026 | Boot Failure detection | The software shall set system alarm as an output when no valid boot was detected after 30 seconds. |  |  | Alarm PIC
APIC_SRS_027 | Click output | The software shall output a single pulse tone (click) with currently set volume for every click input toggles state. If an alarm tone is output just at the time of the click- toggle, the alarm has priority and the click is ignored. Clicks must be output when the alarm is active and alarm tone is paused. |  |  | Alarm PIC
APIC_SRS_028 | Amplifier reset | The software shall reset the amplifier after each tone pulse in order to prevent amplifier freeze. |  |  | Alarm PIC
APIC_SRS_029 | PWM waveform generation | The software shall produce 50% PWM waveform to the amplifier to output tone pulses. After hardware filtering, the speaker must output a tone pulse of 150Hz-1kHz base frequency with at least 4 harmonics between 300Hz and 4kHz. All frequencies' amplitudes must be within a ±15dB band of the base frequency. |  |  | Alarm PIC
APIC_SRS_030 | Boot - 2 clicks | During boot up, the software shall allow the alarm LED to blink twice together with 2 audible clicks simultaneously. |  |  | Alarm PIC
APIC_SRS_031 | Shutdown alarm - Restart | The software shall reset the shutdown alarm once the device restarted with active shutdown alarm. |  |  | Alarm PIC
APIC_SRS_032 | Shutdown alarm - Disabling | When all watchdogs (controller, Monitor, Alarm, NetDCU) are present, the software shall disable the shutdown alarm |  |  | Alarm PIC
APIC_SRS_033 | Alarm PIC | The software shall annunciate audible alarm when the rerouted NetDCU GUI WD0 line fails to toggle. |  |  | Alarm PIC
BLCPIC_SRS_001 | Start-up | The bootloader software shall be the first program to be executed on power up or reset of device. |  |  | Controller PIC Bootloader
BLCPIC_SRS_002 | Wait for ‘#’ | The software shall wait until ‘#’ character is received via UART. |  |  | Controller PIC Bootloader
BLCPIC_SRS_003 | EEPROM or FLASH erase | The software shall erase the EEPROM or FLASH memory area after ‘#’ character received from UART |  |  | Controller PIC Bootloader
BLCPIC_SRS_004 | Reading Data via UART | The software shall read the Controller PIC software data via UART and load into EEPROM or FLASH memory based on the record type (Data:0, End:1, Extend Address:4) received over UART. |  |  | Controller PIC Bootloader
BLCPIC_SRS_005 | Call- Reset | The software shall perform the soft reset when there is a buffer over run or framing error occurred before or during the software load operation. |  |  | Controller PIC Bootloader
BLCPIC_SRS_006 | Direct to Controller PIC software | The software shall transfer control to Controller PIC software when the controller PIC software is loaded into EEPROM or FLASH memory. |  |  | Controller PIC Bootloader
BLCPIC_SRS_007 | Checksum check | The software shall perform checksum operation to ensure integrity of data received via UART. |  |  | Controller PIC Bootloader
BLCPIC_SRS_008 | Bootloader - Configuration bits | The bootloader software shall not write the configuration bits. |  |  | Controller PIC Bootloader
BLCPIC_SRS_009 | Bootloader reset - UART error | UART error shall not cause bootloader reset, therefore UART framing/overflow should not be interpreted. |  |  | Controller PIC Bootloader
BLCPIC_SRS_010 | Bootloader - ROM address | The bootloader should not be activated at the beginning of the startup, therefore the location of the bootloader address should be adjusted to the end of the memory |  |  | Controller PIC Bootloader
BLCPIC_SRS_011 | Bootloader – Timeout | The software shall allow the bootloader timeout to be shortened to allow the ventilation to restart faster following a reset during runtime. |  |  | Controller PIC Bootloader
CPIC_SRS_001 | Exhalation valve control in single-limb ventilation | The software shall set the exhalation valve control to zero (valve fully open), when single limb ventilation modes (nCPAP, DuoPAP, O2 Therapy) are selected. |  |  | Controller PIC
CPIC_SRS_002 | Adjustment of amplitude in HFO Mode for Volume Guarantee | When volume guarantee is enabled in HFO Mode, the software shall adjust the oscillatory pressure amplitude to maintain the tidal volume for each oscillation at the set target. |  |  | Controller PIC
CPIC_SRS_003 | Volume Guarantee operation (conventional modes) | When Volume Guarantee is enabled in conventional dual limb ventilation modes, the software shall adjust the inspiratory pressure for mandatory breaths to maintain the expired tidal volume Vte at the set target. |  |  | Controller PIC
CPIC_SRS_004 | Volume Guarantee with mandatory breaths/PSV breaths | The software shall apply the Volume Guarantee as below: • Apply the VG to mandatory breaths in IPPV, SIPPV, SIMV and SIMV+PSV modes. • Apply the VG to mandatory breaths in PSV mode. • Do not apply the VG to PSV breaths in SIMV+PSV mode. |  |  | Controller PIC
CPIC_SRS_005 | Inspiratory Pressure constraint in vG | The software shall constrain the inspiratory Pressure in VG between PEEP+2mbar to set Pmax or 80 mbar, whichever is lesser. |  |  | Controller PIC
CPIC_SRS_006 | Volume Guarantee maximum delivered pressure | While VG is activated, the software shall prevent delivered pressure from exceeding "Pmax" for a period greater than 50 ms. |  |  | Controller PIC
CPIC_SRS_007 | Volume Guarantee - First breath | The software shall have a mechanism to detect the first breath after the activation of Volume Guarantee |  |  | Controller PIC
CPIC_SRS_008 | Volume Guarantee test breath pressure | The VG test breath (first breath) pressure shall be delivered between Pinsp setting [+0.5 mbar +3% of Pinsp] and (PEEP setting + 2 mbar) – (0.5 mbar + 3% of the pressure). Note: The VG algorithm for the test breath pressure target set by software is PEEP setting + 75% of (Pinsp – PEEP), or PEEP + 2.0 mbar if greater. The actual pressure is subject to the tolerance for pressure delivery. |  |  | Controller PIC
CPIC_SRS_009 | Volume Guarantee for patient-triggered and time-triggered breaths | When VG is enabled, the software shall maintain independent inspiratory pressure control for patient-triggered and time-triggered mandatory breaths. |  |  | Controller PIC
CPIC_SRS_010 | Volume Guarantee (VG) maximum breath-to-breath pressure change (non-HFO modes) | While VG is activated, the change in the delivered pressure between two consecutive breaths of the same trigger category (i.e. time-triggered and patient-triggered) shall not exceed ± [3.0 mbar + 0.5 mbar +3% of 3 mbar]. Note: The change in VG algorithm targeted pressure between two consecutive breaths of the same trigger category (i.e. time-triggered and patient-triggered) must not exceed ± 3.0 mbar. |  |  | Controller PIC
CPIC_SRS_011 | Volume Guarantee auto set volume limit | If the inspired volume exceeds 130% of the set target volume during Volume Guarantee ventilation, the software shall terminate the mandatory breaths. Note: this volume is leak-compensated to avoid premature termination in the presence of airway leak. |  |  | Controller PIC
CPIC_SRS_012 | Volume Guarantee deactivation | If the Volume guarantee is switched off, the software shall use the last set Pinsp value for further ventilation. |  |  | Controller PIC
CPIC_SRS_013 | Risetime - Flow control | The software shall control the flow in a way that the set inspiratory pressure is always reached. |  |  | Controller PIC
CPIC_SRS_014 | Rise Time breaths – flow control | When Rise Time control is enabled, the flow provided to the outlet shall be determined by the software to ensure that the specified pressure rise time can be achieved. Note: In this use case, the I-Flow control is inactive. |  |  | Controller PIC
CPIC_SRS_015 | Rise Time breaths – exhalation valve control | When Rise-Time is enabled, the software shall adjust/control the exhalation valve to enable the pressure to reach the preset target within the time set by the Rise-Time control. Note: in presence of spontaneous inspiratory effort, or with a high-pressure difference from PEEP to target pressure, the measured rise time may be greater than set |  |  | Controller PIC
CPIC_SRS_016 | Apnea ventilation in PSV mode | In PSV mode, the software shall provide the apnea ventilation at the set “Backup Rate” after the preset “Apnea Time”. |  |  | Controller PIC
CPIC_SRS_017 | Apnea ventilation in PSV mode with apnea alarm disabled | In PSV mode, with the “Apnea Time” is set Off, the software shall trigger the backup breaths if the patient breath rate falls below the set “Backup Rate”. |  |  | Controller PIC
CPIC_SRS_018 | Setting Apnea time | In CPAP mode, the software shall have a feature to set the Apnea time, the time during which patient does not breath at all. |  |  | Controller PIC
CPIC_SRS_019 | Backup Control | In CPAP mode, the software shall have a "backup control" that sets the backup breaths from 1 to 5 breaths. |  |  | Controller PIC
CPIC_SRS_020 | Settings for backup breaths in CPAP mode | In CPAP mode, when an interruption in spontaneous breathing exceeds the “Apnea Period”, from 1 to 5 backup breaths (based on the Backup Control) the software shall trigger using pre-determined settings. |  |  | Controller PIC
CPIC_SRS_021 | Apnea ventilation reset | The software shall reset the apnea timer at the start of inspiration of each detected patient breath. |  |  | Controller PIC
CPIC_SRS_022 | Pressure Trigger | When Trigger Mode is set to Pressure Trigger, the software trigger control shall determine the threshold for inspiratory pressure (as a decrement below baseline pressure) required to initiate a patient-triggered breath. |  |  | Controller PIC
CPIC_SRS_023 | NIV flow trigger minimum exhalation | The software shall inhibit the NIV Trigger until: • Exhalation time has exceeded 260 ms, and • Pprox < [CPAP + 0.125 ∙ ΔP + 0.9 mbar] Note: ΔP represents the difference between the inspiratory pressure and expiratory pressure, this ensures that pressure has fallen to within 12.5% of the expiratory level. |  |  | Controller PIC
CPIC_SRS_024 | TCPL pressure control | When Volume Guarantee is disabled, the software shall modulate the exhalation valve force to set the opening pressure for the exhalation valve to a set level (“P-Insp”). |  |  | Controller PIC
CPIC_SRS_025 | TCPL pressure with volume guarantee active | While Volume Guarantee enabled, the software shall modulate the exhalation valve force to set the opening pressure for the exhalation valve to an algorithmically determined pressure target as per set parameters (Volume, Pressure Limit). |  |  | Controller PIC
CPIC_SRS_026 | Rise Time controlled breaths | When the Rise Time control is enabled, the software shall set the mandatory breath type as a modified form of TCPL breath, with auto-set inspiratory flow, and ramped control of the exhalation valve pressure target waveform. |  |  | Controller PIC
CPIC_SRS_027 | Rise time control breath - pressure control | When Rise Time control is enabled, the software shall modulate the exhalation valve force to a target pressure waveform with a ramped increase to the set pressure level over the set Rise Time, followed by a plateau for the remaining I-Time. |  |  | Controller PIC
CPIC_SRS_028 | PSV breath pressure control | The software shall modulate the exhalation valve force to prevent the airway pressure from exceeding a set level: • P-PSV (when configured as absolute value) • PEEP + P-PSV (when configured as delta) |  |  | Controller PIC
CPIC_SRS_029 | DuoPAP breath flow control | The software shall provide a fixed flow of oxygenated air to the breathing circuit throughout the duration of the inspiratory phase of a DuoPAP breath, pre-determined to achieve the target inspiratory pressure (P-High). Note: in DuoPAP mode, the inspiratory pressure level is relative to ambient and not additive to PEEP. Note 2: the flow required to achieve a given pressure target is different for the three supported nCPAP interfaces (Infant Flow, Infant Flow LP, Medijet) |  |  | Controller PIC
CPIC_SRS_030 | Manual breath profile | The software shall make manual breaths are mandatory breaths in all modes, with inspiratory pressure determined by the Pinsp or Pmanual control; and with inspiratory phase continuing as long as the manual breath control is held, or until the maximum set time is reached. |  |  | Controller PIC
CPIC_SRS_031 | Disabling Manual breath | Because manual breath control is edge-triggered, the software shall disable it until 200 milliseconds after the preceding inspiratory period ends. |  |  | Controller PIC
CPIC_SRS_032 | Manual breaths in IPPV, SIPPV, SIMV and SIMV+PSV modes | In IPPV, SIPPV, SIMV and SIMV+PSV modes, manual breaths shall be a mandatory breath with pressure limit determined by the “Pinsp” control, and with inspiratory phase continuing as long as the manual breath control is held, or until the maximum set time is reached. |  |  | Controller PIC
CPIC_SRS_033 | Manual breaths in PSV and CPAP mode | In CPAP mode, manual breaths shall be TCPL breaths with the pressure limit determined by the “P-Manual” control, and with inspiratory phase continuing as long as the manual breath control is held, or until the maximum set time is reached. |  |  | Controller PIC
CPIC_SRS_034 | Manual breaths in nCPAP and DuoPAP modes | In nCPAP and DuoPAP modes, the software shall provide a fixed flow of oxygenated air to the breathing circuit throughout the duration of the inspiratory phase of a manual breath, pre-determined to achieve the target pressure (“P-Manual” in nCPAP mode, “P-High” in DuoPAP mode) |  |  | Controller PIC
CPIC_SRS_035 | Manual breaths in HFO mode | In HFO mode, the manual breaths shall be TCPL breaths with the pressure limit determined by the P-Manual control, with oscillation paused and with inspiratory phase continuing as long as the manual breath control is held, or until the maximum set time is reached. Note: The manual breath function can be enabled/disabled in HFO mode. |  |  | Controller PIC
CPIC_SRS_036 | Manual breath termination on high pressure alarm | The software shall terminate the manual breaths within 200 ms on activation of the High PIP alarm. |  |  | Controller PIC
CPIC_SRS_037 | Standby mode pressure control | When ventilator in Standby mode, the software shall use PEEP as constant 1.0 mbar in exhalation valve. |  |  | Controller PIC
CPIC_SRS_038 | Tube Occlusion Alarm Active | When the "Tube Occlusion" alert is triggered, the software will shut off gas flow to the outlet and open the exhalation valve to relieve pressure in the breathing circuit. |  |  | Controller PIC
CPIC_SRS_039 | Exhalation valve control during occlusion | Once the tube occlusion alarm state is activated, software shall set the exhalation valve control to zero (exhalation valve fully opened). |  |  | Controller PIC
CPIC_SRS_040 | Airway pressure measurement Minimum value | The Software shall measure the minimum Airway pressure value as -10 mbar |  |  | Controller PIC
CPIC_SRS_041 | Airway pressure measurement Maximum value | The Software shall measure the maximum Airway pressure value as +125 mbar |  |  | Controller PIC
CPIC_SRS_042 | Low Inlet Pressure Alarm | The software shall switch to alternative gas when pressure of one supply falls below the low inlet pressure alarm limit. Note: if the oxygen supply fails, the ventilator will operate on air inlet alone, and vice versa. |  |  | Controller PIC
CPIC_SRS_043 | land E flow- Trigger | The software shall, correctly trigger the inspiration/expiration flow at every breath and correctly measure the delays of inspiration/expiration peak flow. |  |  | Controller PIC
CPIC_SRS_044 | Pressure - Measurement range | The software shall adjust the gain and offset values for the gas supply board V4 to reach up to 150mbar. |  |  | Controller PIC
CPIC_SRS_045 | Ventilation - Restart time | The software shall recover faster and continue the ventilation after processor reset during runtime to prevent any patient risk (max. 7 seconds of no ventilation period is allowed). |  |  | Controller PIC
CPIC_SRS_046 | Ventilation - Start after boot | With regular boot the software shall not start the ventilation before blender calibration is ready. |  |  | Controller PIC
CPIC_SRS_047 | Triggered breath - Pressure | In triggered modes with backup capability support, the software shall set the pressure of a triggered breath after backup ventilation as high as the previous backup pressure. |  |  | Controller PIC
CPIC_SRS_048 | Second or third Triggered breath - Pressure | When triggering a breath after backup ventilation is active, the software shall support the second or third triggered breath with the last active backup pressure. |  |  | Controller PIC
CPIC_SRS_049 | Manual breath - Overpressure | When overpressure occurs during manual breath, the software shall continue the ventilation at the end of the manual breath. |  |  | Controller PIC
CPIC_SRS_050 | Disconnection alarm - Lowest settings | The software shall trigger the disconnection alarm at the lowest settings and the disconnection must be recognized even if the minimal P-insp and VG settings are active. |  |  | Controller PIC
CPIC_SRS_051 | VG - Pmax parameter | The software shall provide a new parameter called P-Max in the user interface so that it can be treated separately from P-insp during volume guarantee, and the P-Max value is considered as a new maximum pressure allowed for ventilation when VG is active. |  |  | Controller PIC
CPIC_SRS_052 | VG - Pinsp | The software shall able to reduce Pinsp down to 0.9 mbar on top of PEEP during volume guarantee. |  |  | Controller PIC
CPIC_SRS_053 | VG - Switch after boot | The software shall adjust all settings before starting the controller, when volume guarantee is turned on directly after device boot. |  |  | Controller PIC
CPIC_SRS_054 | VG + Risetime - I-flow calculation | When VG is active and risetime is chosen, the software shall calculate the I-flow based on the actual pressure applied. |  |  | Controller PIC
CPIC_SRS_055 | VG - SIMV+PSV mode breath difference | In SIMV+PSV mode when VG is active, the software shall calculate the steepness of SIMV and PSV breath independently and Ppsv must be reached independently from the VG. |  |  | Controller PIC
CPIC_SRS_056 | VG+HFO - Minimum amplitude | The software shall set the possible minimum amplitude between 1 and 2 mbar, when VG is active in HFO mode. |  |  | Controller PIC
CPIC_SRS_057 | HFO - Distortion in pressure | In HFO mode, due to exhalation valve support, the software shall not create any pressure curve distortion. |  |  | Controller PIC
CPIC_SRS_058 | HFO - Manual breath overshoot | In HFO mode when manual breath is applied, the software shall not create overpressure. |  |  | Controller PIC
CPIC_SRS_059 | HFO - Pressure while stepping to another mode | When switching from HFO to another mode, the software shall release the pressure for 1.5 seconds with the HFO valve open, then close it and begin the new ventilation. |  |  | Controller PIC
CPIC_SRS_060 | Brownout reset - Voltage | The software shall set the brownout reset to lowest possible voltage, to prevent unnecessary resets caused by voltage spikes. |  |  | Controller PIC
CPIC_SRS_061 | Command 'k' - Bit 5 | If bit 5 in the command 'k' is set, the software shall deliver a value of 128 from the offset calculation which means the offset current to the HFO-coil is zero. If the bit is reset, the calculation must be as previously. |  |  | Controller PIC
CPIC_SRS_062 | Low P, high flow-Spikes in pressure | Using low pressure values with high flow settings the software shall not create spikes at the beginning and end of the inspiration (Kirchendach-Effect) |  |  | Controller PIC
CPIC_SRS_063 | CPAP - Backup | When backup is active in CPAP mode, the software shall control the Pinsp correctly to reach its preset value. |  |  | Controller PIC
CPIC_SRS_064 | Risetime - Insp flow | In Risetime mode, the software shall calculate the inspiratory flow correctly and ensure Pinsp is reached always. |  |  | Controller PIC
CPIC_SRS_065 | HFO mode - Expiration time | When HFO is turned off, the software shall maintain the exp time for at least 200 milliseconds, and 2 seconds when recruitment is ongoing. |  |  | Controller PIC
CPIC_SRS_066 | Pressure trigger | The software shall provide the Pressure trigger in conventional trigger modes and it works identical to Flow and Volume triggers. |  |  | Controller PIC
CPIC_SRS_067 | FOT -closed loop amplitude control | The software shall create high frequency by using expiratory valve and the created amplitude is controlled by a closed loop. |  |  | Controller PIC
CPIC_SRS_068 | Variable leak compensation | Controller PIC software shall support the following NIV leak compensation options: -15% (Low) -30% (Middle) -40% (High level) |  |  | Controller PIC
CPIC_SRS_069 | SIPPV - Trigger time window | In SIPPV mode, the software shall change the trigger time window from 1/2 l-time to 180 ms fixed value. |  |  | Controller PIC
CPIC_SRS_070 | Startup Delay | In DUOPAP mode, if the 1st breath interfered with the blender start-up operation, the software shall increase the START DELAY from 8.5 to 11.5 seconds. |  |  | Controller PIC
CPIC_SRS_071 | Internal disconnection | To avoid overpressure when the patient is reconnected, the software shall identify internal disconnection more precisely and limit the flow to default in the event of disconnection. |  |  | Controller PIC
CPIC_SRS_072 | VG Optimization | The software shall optimize the VG, to smoothly control the volume at very low volumes (2.5 …0.8 ml). |  |  | Controller PIC
CPIC_SRS_073 | Tube occlusion recovery | In DUOPAP mode after tube occlusion situation, the software shall recover the pressure levels within a few breaths so that it does not create tube occlusion again if the circumstances allow it (the tubes are not occluded anymore). |  |  | Controller PIC
CPIC_SRS_074 | HFO - filter | In HFO mode, the software shall apply the filter on pressure and flow signals to compensate the underdamping effect of the pressure line. |  |  | Controller PIC
CPIC_SRS_075 | HFO - startup | When entering HFO mode, the software shall slow down the Pmean rise to avoid overshoot. |  |  | Controller PIC
CPIC_SRS_076 | Startup - timing | At startup, the software shall make the first blender call happen only after the settings are arrived/read. |  |  | Controller PIC
CPIC_SRS_077 | DUOPAP – overpressure | To avoid the hanging situation, the software shall start the recovery from overpressure situation at 4 lpm. |  |  | Controller PIC
CPIC_SRS_078 | Trigger modes – Backup frequency | The software shall correctly apply the backup frequencies /mandatory breaths. |  |  | Controller PIC
CPIC_SRS_079 | NIV – flow value | In NCPAP and DUOPAP modes, the software shall not make any corrections to the applied flow values while transmitting them to the GUI |  |  | Controller PIC
CPIC_SRS_080 | NIV – pressure measurement | The software shall compensate the effects of the purge flow and correctly measure the patient side pressure in NCPAP and DUOPAP modes. |  |  | Controller PIC
CPIC_SRS_081 | NCPAP incorrect pressure level | In Nasal CPAP mode, the software shall create the flow compensation independent of inspiratory values. |  |  | Controller PIC
CPIC_SRS_082 | Pressure not reached | The software shall not trigger the "Pressure not reached" situation anymore. |  |  | Controller PIC
CPIC_SRS_083 | CPAP/PEEP levels | The software shall achieve the correct PEEP and CPAP levels independent from the used flow setting. (The exhalation valve must be able to handle higher flow levels) |  |  | Controller PIC
CPIC_SRS_084 | CPAP – Sync mode step | When switching from CPAP to Synchronized CMV mode, the software shall ensure that no inspiratory hold occurs. |  |  | Controller PIC
CPIC_SRS_085 | CPAP - overpressure | After closing a big leak in CPAP mode, the software shall control the CPAP level correctly. |  |  | Controller PIC
CPIC_SRS_086 | False low PEEP alarm | The software shall not trigger the false "Low PEEP" alarm at startup. |  |  | Controller PIC
CPIC_SRS_087 | Pduo leak compensation | In DuoPAP mode, the software shall allow the maximum flow to reach the desired Pduo level of 25 l/min independent of the set leak compensation level setting. |  |  | Controller PIC
CPIC_SRS_088 | False “check Blender code #64” alarm | The software shall check the blender upon start-up and give an alarm, if leakage is detected at the proportional valve or at the blender flow sensor (note this alarm occurs when switching the vent off in HFO mode and restarting the vent) |  |  | Controller PIC
CPIC_SRS_089 | No Patient Disconnect alarm in Volume Guarantee mode | The software shall trigger Patient Disconnect alarm, when the proximal pressure line is disconnected while Volume Guarantee is active. |  |  | Controller PIC
CPIC_SRS_090 | False tube occlusion situation | The software shall be able to regulate the ventilation without false tube occlusion alarms even at extreme settings. (e.g.: high I-flow, short I-time, such as: · Mode SIMV+PSV · E Flow = 6 lpm · I Flow = 32 lpm · Trigger Vol = 0.7 L · I-time = 0.1 s · Freq = 76 bpm · PEEP = 5 cmH2O · Pinsp = 25 cmH2O · Pmax = 25 cmH2O · Vguarant = 24 mL · O2 = 21 %) |  |  | Controller PIC
CPIC_SRS_091 | Overpressure condition | The software shall perform below during overpressure conditions: · Reduce the pressure to PEEP when the proximal pressure reaches the High PIP Alarm Limit (i.e. high-pressure condition) and activate the High PIP Alarm. · Reduce the pressure to atmospheric pressure when the high-pressure condition occurs, and the pressure fails to drop below the High PIP Alarm Limit within 100 milliseconds. |  |  | Controller PIC
CPIC_SRS_092 | Inspiratory pressure jumped during manual breath | The software shall maintain its inspiratory pressure level during the entire manual breath period. |  |  | Controller PIC
CPIC_SRS_093 | HFO amplitude re-ramping | The software shall maintain consistent pressure amplitude and mean airway pressure (Pmean) in HFO mode. |  |  | Controller PIC
CPIC_SRS_094 | Non-terminated alarm | The software shall terminate disconnection alarm after the reconnection of the patient circuit. |  |  | Controller PIC
CPIC_SRS_095 | NIV flow - limitation | In NCPAP and DUOPAP modes, the software shall ensure the applied flow values are not exceeding 20 lpm and the leak compensation setting related maximum values, whichever is smaller. |  |  | Controller PIC
CPIC_SRS_096 | DUOPAP – leak compensation | In DUOPAP mode, the software shall make both CPAP and DUO flows dependent on leak compensation setting. |  |  | Controller PIC
CPIC_SRS_097 | NCPAP/DUOPAP flow compensation | The software shall allow the flow compensation up to the following values corresponding to user selected levels of compensation in NCPAP and DUOPAP modes: · 0% (OFF) · 10% (LOW) · 20% (MED) · 30% (HIGH) |  |  | Controller PIC
CPIC_SRS_098 | Missing patient disconnection alarm | The software shall detect patient disconnect and alarm for all applicable generator types in NCPAP and DUOPAP modes when leak compensation is set to OFF. |  |  | Controller PIC
CPIC_SRS_099 | Exhalation Valve Control | The Controller PIC software shall control the exhalation valve through a digital-to-analog converter (DAC). |  |  | Controller PIC
CPIC_SRS_100 | Controller PIC 12-bit ADC converter | The Controller PIC software shall read below sensors data through 12-bit ADC converter. • Air inlet pressure • Oxygen inlet pressure • Inspiration port pressure • Proximal pressure |  |  | Controller PIC
CPIC_SRS_101 | Volume Guarantee (VG) minimum delivered breath pressure (non-HFO modes) | While VG is activated, the software shall ensure the delivered pressure for each breath is not less than (PEEP setting + 2 mbar) - [0.5 mbar +3% of the pressure]. Note: 1. The VG algorithm targeted pressure for each breath must not be less than the PEEP setting + 2 mbar. 2. Pressure error tolerance off ± (0.5 mbar + 3% of actual pressure) |  |  | Controller PIC
CPIC_SRS_102 | Volume Guarantee (VG) maximum delivered breath pressure (non-HFO modes) | While VG is activated, the software shall ensure the delivered pressure for the test breath or breath following a setting change or compliance/resistance change is not exceeding the Pmax setting by more than 3.0 mbar + [0.5 mbar +3% of 3.0 mbar]. Note: The VG algorithm targeted pressure for each breath must not exceed Pmax setting. |  |  | Controller PIC
CPIC_SRS_103 | Manual breath in HFO mode | The software shall deliver a breath with pressure set by Pmanual setting when Manual breath is activated in HFO mode. |  |  | Controller PIC
CPIC_SRS_104 | Commands to the Blender PIC | The software shall send below commands to the Blender PIC via SPI bus: • Requested air flow • Requested O2 flow • O2 21% calibration • O2 100% calibration • Inspiration/expiration • Non-invasive mode |  |  | Controller PIC
CPIC_SRS_105 | Error status from Blender PIC | The software shall receive the below operational error status from the Blender PIC via SPI bus: • Air PWM Limit • O2 PWM Limit • Communication • Air Flow sensor • O2 Flow sensor • Program Memory Checks • Zero Calibration • FRAM |  |  | Controller PIC
CPIC_SRS_106 | Commands to HFO PIC | The software shall send below commands to the HFO PIC via SPI bus: • HF_FREQ: Frequency • HF_AMPL: Amplitude • HF_STATUS: Status • HF_FORM: Waveform • HF_OFFS: Offset • SPI_CHECKS: Take over setting |  |  | Controller PIC
CPIC_SRS_107 | Communication with Monitor PIC | The software shall read the breath trigger signal from the Monitor PIC and send the inspiration/exhalation status to the Monitor PIC via digital IO. |  |  | Controller PIC
CPIC_SRS_108 | Communication with Alarm PIC | The software shall send the watch dog status to Alarm PIC via digital Output. |  |  | Controller PIC
CPIC_SRS_109 | Commands from NetDCU | The software shall receive below commands from NetDCU via UART: 's': Volume correction 'S': Autoflow correction: 128=Mitte 'r': Option to the Mode 'w': I-Time 'e': S_EXP_TIME 'i': Oxi_ist 't': I_Pressure 'z': S_PEEP 'c': Pressure_psv_target 'd': Oxi_target 'u': I_Flow 'a': E_Flow 'o': Mode to adjust 'g': Prox_Cal 'f': HF-Frequenz 'm': HF-Amplitude mmWs 'j': MAP_Pressure 'h': HFO:restart Amplitude 'v': T_apnea_s 'k': HW_Configuration '(': Mode_option_2 '[': Plateau 0... 80% '!': P_recruit '<': P_max_VG ':': Hi PIP Alarm [0,1 mb]] '#': Reset f.Download |  |  | Controller PIC
CPIC_SRS_110 | Response to NetDCU | The software shall send below response to NetDCU via UART: 'F': HF-Freq. 'M': HF-Ampl. 'J': HF-MAP 'x': N-Mode E_Flow 'X': N-Modes I_Flow 'R': Vol. Garantie 'W': I-Time 'E': E-Time 'T': I-Pressure 'Z': E-Pressure 'C': Pressure_psv_Target 'D': Oxi_Target 'U': I_flow 'A': E_flow 'O': Mode 'B': Prox_scale 'b': Demand Flow 'G': Prox_offs 'L': Status all 2 sec 'N': Checksumme 'Y': Vers-Nr 'V': T_apnea 'H': Mixer-Vers-Nr 'I': Misxer-Status 'K': Hardware_Config ')': ModeOption_2 ']': Plateau % '?': P_recr '>': P_max_VG '*': Hi PIP Alarm |  |  | Controller PIC
CPIC_SRS_111 | ATG module - Support | The controller PIC shall support ATG module 4.1. |  |  | Controller PIC
CPIC_SRS_112 | VG - Tube occlusion reference pressure | The software shall set the Tube occlusion pressure threshold value to higher value between Pinsp and Pmax, when Volume Guarantee is active. |  |  | Controller PIC
CPIC_SRS_113 | Correction values - Resend | When flow detects no breath, the correction values shall be resent. If the flow breath detection does not trigger, use the Inspiration flag to start data transfer to the controller PIC. set the correction value to cause a pressure increase on the next breath in this situation. |  |  | Controller PIC
CPIC_SRS_114 | Inspiration - Volume excess Termination | The Software shall not terminate the Inspiration when volume exceeds 130% of settings |  |  | Controller PIC
CPIC_SRS_115 | Volume Value | The software shall not have zero volume value during stable ventilation conditions at minimum settings. |  |  | Controller PIC
CPIC_SRS_116 | Flow calibration - Bar placement | The calibration bars shall be displayed at the rightmost edge of the calibration symbol when flow sensor is turned off. |  |  | Controller PIC
CPIC_SRS_117 | Exhalation Valve Control | The Controller PIC software shall control the exhalation valve through a digital-to-analog converter (DAC). |  |  | Controller PIC
CPIC_SRS_118 | Controller PIC 12-bit ADC converter | The Controller PIC software shall read below sensors data through 12-bit ADC converter. • Air inlet pressure • Oxygen inlet pressure • Inspiration port pressure • Proximal pressure |  |  | Controller PIC
CPIC_SRS_119 | Therapy – lookup table | At the device Fabian Therapy evolution, the lookup table shall be adjusted to make compatible with the situation that there is no purge flow. |  |  | Controller PIC
BLHPIC_SRS_001 | Start-up | The bootloader software shall be the first program to be executed on power up or reset of device. |  |  | HFO PIC Bootloader
BLHPIC_SRS_002 | BOOT TIMEOUT duration | The software shall wait either BOOT TIMEOUT duration or until ‘#’ character is received (via UART), whichever is earlier. |  |  | HFO PIC Bootloader
BLHPIC_SRS_003 | Soft Reset | The software shall perform soft reset to launch the HFO PIC application software when there is a buffer over run or framing error while waiting for BOOT TIMEOUT duration to expire. |  |  | HFO PIC Bootloader
BLHPIC_SRS_004 | Direct to HFO PIC software | The software shall transfer control to HFO PIC software address to execute the already loaded software when ‘#’ is not received from UART while waiting for BOOT TIMEOUT duration to expire. |  |  | HFO PIC Bootloader
BLHPIC_SRS_005 | EEPROM or FLASH erase | The software shall erase the EEPROM or FLASH memory area when ‘#’ is received from UART during BOOT TIMEOUT duration. |  |  | HFO PIC Bootloader
BLHPIC_SRS_006 | Reading Data via UART | The software shall read the HFO PIC software data via UART and load into EEPROM or FLASH memory based on the record type (Data:0, End:1, Extend Address:4) received over UART. |  |  | HFO PIC Bootloader
BLHPIC_SRS_007 | Call- Reset | The software shall perform the soft reset after HFO PIC software is loaded into EEPROM or FLASH memory. |  |  | HFO PIC Bootloader
BLHPIC_SRS_008 | Checksum check | The software shall perform checksum operation to ensure integrity of data received via UART. |  |  | HFO PIC Bootloader
HPIC-SW_001 | Voice coil activation | The software shall drive a voice coil through Pulse Width Modulation (PWM). |  |  | HFO PIC
HPIC-SW_002 | Oscillations to maintain airway pressure value | The software shall deliver fresh gas oscillation to maintain the airway pressure value. |  |  | HFO PIC
HPIC-SW_003 | Proximal Pressure Sensor Reading | The software shall read the airway pressure (Proximal pressure) through the Proximal Pressure Sensor (analog input 0 to 4.096V max). |  |  | HFO PIC
HPIC-SW_004 | Voice Coil Sinusoidal signal generation | The software shall calculate the Voice Coil sinusoidal oscillation based on below parameters received from controller PIC via SPI bus. • Frequency • Amplitude • Status (HFO mode, non-HFO mode) • Waveform • Offset Note: The Controller PIC plays the master role in SPI communication scheme. |  |  | HFO PIC
HPIC-SW_005 | High frequency Valve ON | The software shall turn ON the HF Valve to direct mixed fresh gas to the HFO module when HFO mode is active. |  |  | HFO PIC
HPIC-SW_006 | High frequency Valve OFF | The software turns the HF Valve OFF when non-HFO modes are active. |  |  | HFO PIC
HPIC-SW_007 | 3-Way Valve ON | The software shall turn ON the 3-Way Valve, when HFO mode is active. |  |  | HFO PIC
HPIC-SW_008 | 3-Way Valve OFF | The software shall turn OFF the 3-Way Valve, when HFO mode is active. |  |  | HFO PIC
HPIC-SW_009 | Communication with NetDCU | For debugging purpose, the software shall read the below commands communicated by NetDCU via UART. • '#': Reset for Download • x': Output ON/Off • 'f': Frequency (enable frequency value change received from Controller PIC) • 'a': Amplitude (enable amplitude value change received from Controller PIC) • 'o': Offset (enable Offset value change received from Controller PIC) • 's': Waveform (enable waveform change received from Controller PIC) • 'r': HF ON (Run) • 'h': HF ON/OFF • 'Y': read Versions Number |  |  | HFO PIC
HPIC-SW_010 | Communication with NetDCU | For debugging purpose, the software shall send the below data to NetDCU via UART. • 'F'; Frequency • 'A'; Amplitude • 'O'; Offset • 'S'; Waveform • 'R'; HF ON (Run) • 'x'; Debug1 • 'y'; Debug2 |  |  | HFO PIC
HPIC-SW_011 | Communication with Controller PIC | The software shall receive the below parameters communicated by Controller PIC via SPI bus. • HF_FREQ: Frequency • HF_AMPL: Amplitude • HF_STATUS: Status • HF_FORM: Waveform • HF_OFFS: Offset • SPI_CHECKS: Take over setting |  |  | HFO PIC
HPIC-SW_012 | Sinusoidal Waveform | The software shall use a base sine wave that has 164 samples. For different amplitude, the waveform is scaled in y direction, which is PWM duty cycle. For different frequencies, the waveform is scaled in x direction, which is the time-period per sample. |  |  | HFO PIC
BLMPIC_SRS_001 | Commands from UIM | The software shall receive following commands from UIM via UART. • Bootloader software ID • Bootloader software version • Read program • Write program • Write configuration • Reset • No Acknowledgement |  |  | Monitor PIC Bootloader
BLMPIC_SRS_002 | Response to UIM | The software shall send the following response to UIM via UART. • Acknowledgement, if program command successful • No Acknowledgement, if program command unsuccessful |  |  | Monitor PIC Bootloader
BLMPIC_SRS_003 | Disable interrupt | The software shall disable interrupt at start-up the boot loading operation. |  |  | Monitor PIC Bootloader
BLMPIC_SRS_004 | Bootloader software id command | The software shall send the bootloader software id, when bootloader software id command received via UART. |  |  | Monitor PIC Bootloader
BLMPIC_SRS_005 | Bootloader software version command | The software shall send the bootloader software version, when bootloader software version command received via UART. |  |  | Monitor PIC Bootloader
BLMPIC_SRS_006 | Read program command | The software shall read the Monitor PIC software data from FLASH memory, when read program command received via UART. |  |  | Monitor PIC Bootloader
BLMPIC_SRS_007 | Erase memory | The software shall erase the FLASH data before loading software when write program command received via UART. |  |  | Monitor PIC Bootloader
BLMPIC_SRS_008 | Write program command | The software shall load the Monitor PIC software into FLASH memory, when write program command received via UART. |  |  | Monitor PIC Bootloader
BLMPIC_SRS_009 | Write configuration command | The software shall send 'No Acknowledgement’, when write configuration command received via UART. |  |  | Monitor PIC Bootloader
BLMPIC_SRS_010 | Reset command | The software shall perform soft reset when reset command received via UART. |  |  | Monitor PIC Bootloader
BLMPIC_SRS_011 | No Acknowledgement command | The software shall perform soft reset when 'No Acknowledgement' command received via UART. |  |  | Monitor PIC Bootloader
BLMPIC_SRS_012 | Bootloader - ROM address | The bootloader software shall not be activated at the beginning of the start-up. Note: The location of the Bootloader address should be adjusted to the end of the memory. |  |  | Monitor PIC Bootloader
MPIC SRS_001 | Volume - Value display | The software shall not have zero volume value during stable ventilation conditions at minimum settings. |  |  | Monitor PIC
MPIC SRS_002 | Pressure - Measurement range | The software shall adjust the gain and offset for the gas supply board to reach up to 150 mbar. |  |  | Monitor PIC
MPIC SRS_003 | Inspiration -Volume excess Termination | The software shall not terminate the Inspiration when volume exceeds 130% of settings. |  |  | Monitor PIC
MPIC SRS_004 | Triggered breath -Flag | The software shall reset triggered breath flag when VG controller switching from a synchronized to a non-synchronized ventilation mode. |  |  | Monitor PIC
MPIC SRS_005 | Correction values – Resend | The software shall perform the following actions, when the flow breath detection did not trigger: a) Resent the correction values to controller PIC. b) Make sure to use the inspiration flag to initiate data transfer to the controller PIC c) Set the correction values to force a pressure increase with the next breath. |  |  | Monitor PIC
MPIC SRS_006 | Tube Blocked - Low settings | The software shall detect the blocked tube and the tube blocked alarm to be triggered even if the minimum Pins and VG Settings are active. |  |  | Monitor PIC
MPIC SRS_007 | Volume - Possible lowest Value | The software shall measure volume minimum value as 0.2 ml. |  |  | Monitor PIC
MPIC SRS_008 | MPIC - Restart relay test | The software shall open the relay for test when monitor PIC is restarted. |  |  | Monitor PIC
MPIC SRS_009 | Apnea – Calculation | The software shall calculate Apnea time since the end of the last breath. |  |  | Monitor PIC
MPIC SRS_010 | Leak - Negative Value | The software shall deactivate the negative leak Clean Flow Sensor Alarm when volume measures lower than 2 ml. |  |  | Monitor PIC
MPIC SRS_011 | Flow range - Alarm at change | The software shall set flow sensor not calibrated alarm upon changing the patient range. |  |  | Monitor PIC
MPIC SRS_012 | Startup - Oxygen Measurement | The software shall start the oxygen measurement after the 35 seconds of the Ventilator system startup. |  |  | Monitor PIC
MPIC SRS_013 | VG - Pmax parameter | The software shall update and apply the new value of Pmax immediately if it is changed when Volume Guarantee is active. |  |  | Monitor PIC
MPIC SRS_014 | VG - Small and stiff lung | The software shall have the Volume Guarantee operation even when the Lung Compliance (C) value is at least 0.1 and very small Stiff Test Lung. |  |  | Monitor PIC
MPIC SRS_015 | VG - l-Flow, Pmax alarms | The software shall trigger the Volume not reached alarm in below mentioned any of the two conditions: a) If Measure Flow is too low for the last 2 breaths. b) If measured Pmax is too low on the last one breath. |  |  | Monitor PIC
MPIC SRS_016 | VG - VG not reached Alarm I | The software shall set the tolerance of the alarm caused by low Pmax to 0.6 mbar. |  |  | Monitor PIC
MPIC SRS_017 | VG -VG not reached Alarm II | The software shall set the tolerance of the alarm caused by low I-flow as per below. a) Tolerance is 12.5% of set volume when set volume is above 10 ml. b) Tolerance is 50% of set volume when set volume is between 5ml - 10ml range c) Tolerance is 25% of set volume when set volume is between 2.5ml to 5ml (excluding 5ml) range. d) Tolerance is 0 when set volume is below 2.5ml. |  |  | Monitor PIC
MPIC SRS_018 | VG+Tube occlusion - Reference Pressure | The software shall set the Tube occlusion pressure threshold value to higher value between Pinsp and Pmax, when Volume Guarantee is active. |  |  | Monitor PIC
MPIC SRS_019 | VG+HFO - Minimum amplitude | When VG is activated in HFO mode the software shall have possible minimum amplitude value between 1 mbar and 2 mbar. |  |  | Monitor PIC
MPIC SRS_020 | I and E Flow – Trigger | The software shall correctly trigger the flow for every breath inspiration and expiration, and delays of inspiration/expiration peak flow measurement to be correct. |  |  | Monitor PIC
MPIC SRS_021 | 02 compensation - change | When there is a change in the state of O2 compensation (on/off), it should have an immediate effect. |  |  | Monitor PIC
MPIC SRS_022 | Bits – Nebulizer | The SW should be updated to the lack of nebulizer in MB version 4. |  |  | Monitor PIC
MPIC SRS_023 | Tube blocked – Alarm | During Volume Limit operation, the software shall handle Tube Blocked Alarm like as in Volume Guarantee operation by keeping the alarm off if it persists. |  |  | Monitor PIC
MPIC SRS_024 | Flow sensor – Calibration | The software shall terminate the Flow Sensor calibration and set the old values back if any kind of flow is detected. |  |  | Monitor PIC
MPIC SRS_025 | The percentage of respiratory part of volume (volume delivered by the respirator) | The software shall limit the value of respiratory part of volume (volume delivered by the respirator) to 200%. |  |  | Monitor PIC
MPIC SRS_026 | The average and the breath-by-breath expiratory volume of the patient should be… | The software shall limit the Vte as average and the breath-by-breath expiratory volume of the patient to max 600 ml. |  |  | Monitor PIC
MPIC SRS_027 | Flow sensor - clean alarm | The software shall set the negative leak delay to 29 breath (299 breaths in HFO) and the limit of the negative leak to 49%. |  |  | Monitor PIC
MPIC SRS_028 | Inspiration bit – Correction | The software shall set the inspiration bit to active for every ventilation mode except O2 Therapy. |  |  | Monitor PIC
MPIC SRS_029 | Measurements – Incorrect value avoidance | The software shall set Pmean and frequency values to hold zero until there is a valid measurement available in case of a setting changed. |  |  | Monitor PIC
MPIC SRS_030 | Flow sensor – Connection correction | The software shall not set "Flow Sensor not Connected" false alarm when flow sensor becomes activated manually. |  |  | Monitor PIC
MPIC SRS_031 | Pressure trigger | The software shall have the pressure trigger available in conventional trigger modes and it should work identical to flow and Volume triggers. |  |  | Monitor PIC
MPIC SRS_032 | Pressure trigger -Frequency measurement | The software shall perform the frequency measurement based on pressure measurement. |  |  | Monitor PIC
MPIC SRS_033 | Pressure trigger - Apnea alarm | The software shall have Apnea alarm even when Flow Sensor is off. |  |  | Monitor PIC
MPIC SRS_034 | Compliance – resolution | The software shall have resolution of compliance measurement as below. • Minimum value to 0.01 • Maximum value to 50.000 |  |  | Monitor PIC
MPIC SRS_035 | Alarm - Volume not reached | At VG, the software shall set the "Volume not reached" criteria to 6% below set volume. |  |  | Monitor PIC
MPIC SRS_036 | PSV - Rise time mode | The software shall terminate the inspiration when a minimum pressure of 88% (P_Psv-PEEP) above PEEP is achieved. |  |  | Monitor PIC
MPIC SRS_037 | Volume curve – reset | The software shall not reset volume curve to zero when during a long inspiratory plateau, the flow changes to expiratory direction without creating a volume higher than 25% of the average measured volume. |  |  | Monitor PIC
MPIC SRS_038 | Pmean – average | The software shall set ATZ_p_mean / AVG_p_mean averaging time to 10 sec in all modes and the manual breath to be taken out of averaging in CPAP and NCPAP modes. |  |  | Monitor PIC
MPIC SRS_039 | T_apnea set to UI-Setting when CPAP + Backup OFF | The software shall have Apnea alarm only after the user set the T_apnea. |  |  | Monitor PIC
MPIC SRS_040 | Rise time mode + PSV + VG | The software shall maintain a minimum ventilation of 2mbar over PEEP in Rise time mode + PSV + VG to not to terminate PSV breath before a minimum pressure of PEEP+2mbar is reached and when the inspired (leak corrected) volume exceeds the set volume or the flow-criteria is reached. |  |  | Monitor PIC
MPIC SRS_041 | N - Mode - Detection of overpressure | The software shall trigger no alarm when large leak is suddenly removed in NIV mode as overpressure detection is faster in NIV mode. |  |  | Monitor PIC
MPIC SRS_042 | PSV + Rise Time | The software shall extend the breath beyond T_rise if flow criteria is not reached. |  |  | Monitor PIC
MPIC SRS_043 | FOT - signal filtering | The software shall remove the flow path when FOT is active in conventional-FOT mode and HFO-FOT mode in case of any delay and SW-filtering in pressure. |  |  | Monitor PIC
MPIC SRS_044 | HFO - BTB Mvexp | The software shall calculate the Mvexp correctly in HFO mode. |  |  | Monitor PIC
MPIC SRS_045 | NIV - Breath Detection | In single limb ventilation, the software shall detect the breath based on flow and pressure signals with all the available leak compensation. |  |  | Monitor PIC
MPIC SRS_046 | Pressure trigger - PED mode | The software shall have the Pressure trigger available in Pediatric (PED) mode as well. |  |  | Monitor PIC
MPIC SRS_047 | HFO – filter | In HFO mode, the software shall apply a filter on pressure and flow sensors in order to compensate the underdamping effect of the pressure line. |  |  | Monitor PIC
MPIC SRS_048 | NIV - pressure measurement | The software shall allow the effect of the purge flow to be compensated and the patient side pressure to be measured correctly in NCPAP and DUOPAP. |  |  | Monitor PIC
MPIC SRS_049 | Tube occlusion alarm | The software shall trigger the tube occlusion alarm in case of expiratory limp is closed. |  |  | Monitor PIC
MPIC SRS_050 | SIMV skips mandatory breaths | The software shall deliver mandatory breaths matching Frequency (Freq) setting when there is no patient effort. |  |  | Monitor PIC
MPIC SRS_051 | Wrong volume measurement above 3000 meters | The software shall apply the altitude compensation factors correctly for altitudes above 3000 meters. |  |  | Monitor PIC
MPIC SRS_052 | VTe not reached alarm activation (non-HFO modes) | The software shall activate the medium priority “VTe not reached / check settings” alarm when the Monitored Tidal Volume (Vte) is less than 94% of the Volume Guarantee (VG) setting for 5 consecutive breaths. |  |  | Monitor PIC
MPIC SRS_053 | VTe not reached alarm deactivation (non-HFO modes) | Once activated, the software shall deactivate the medium priority “VTe not reached / check settings” alarm when the Monitored VTe is greater than or equal 94% of the VG setting for 1 breath. |  |  | Monitor PIC
MPIC SRS_054 | VTe not reached alarm in HFO mode activation (HFO mode) | The software shall activate the medium priority “VTe not reached / check settings” alarm when the Monitored VTe is less than 94% of the VG setting for 10 seconds. |  |  | Monitor PIC
MPIC SRS_055 | VTe not reached alarm in HFO mode deactivation (HFO mode) | Once activated, the software shall deactivate the medium priority “VTe not reached / check settings” alarm when the Monitored VTe is greater than or equal 94% of the VG setting for 3 seconds. |  |  | Monitor PIC
MPIC SRS_056 | Proximal pressure measurement | The software shall read continuously the pressure from the pressure sensor to determine the proximal pressure at patient airway interface. |  |  | Monitor PIC
MPIC SRS_057 | O2 Measurement inlet | The software shall measure the O2 which is read from inlet pressure transducers and two Oxygen sensors. |  |  | Monitor PIC
MPIC SRS_058 | Air Measurement inlet | The software shall measure the Air which is read from inlet pressure transducers. |  |  | Monitor PIC
MPIC SRS_059 | Flow sensor reading | The software shall read the flow value from the Proximal Flow Sensor (hot-wire) at patient circuit and calibrate its slope and offset if required via DAC. |  |  | Monitor PIC
MPIC SRS_060 | Inspiration pressure reading | The software shall read the inspiration from inspiration pressure transducer. |  |  | Monitor PIC
MPIC SRS_061 | Power status | The software shall monitor the voltage status of A5V, BL24V and AVS5V. |  |  | Monitor PIC
MPIC SRS_062 | Fan status | The software shall monitor the cooling fan status. |  |  | Monitor PIC
MPIC SRS_063 | Patient trigger to Controller PIC | The software shall detect and send patient triggering information to Controller PIC. |  |  | Monitor PIC
MPIC SRS_064 | Exhalation status from Controller PIC | The software shall receive the exhalation valve status from Controller PIC. |  |  | Monitor PIC
MPIC SRS_065 | Monitor PIC Watchdog (WD) status. | The software shall send hardware watch dog status to Alarm PIC. |  |  | Monitor PIC
MPIC SRS_066 | Read and store Flow Sensor - Calibrate data | The software shall read and store Proximal Flow Sensor (hot-wire) calibrated data from/to in persistent memory (FRAM). |  |  | Monitor PIC
MPIC SRS_067 | Monitor PIC Checksum | The software shall calculate the checksum for monitor PIC and report checksum error, if any. |  |  | Monitor PIC
MPIC SRS_068 | Communication interface - Trigger bits | The communication interface should provide a possibility to use more than 2 trigger mode. |  |  | Monitor PIC
MPIC SRS_069 | Noise suppression - reset I | For 15 sec no breaths, the noise suppression should reset automatically. |  |  | Monitor PIC
MPIC SRS_070 | Noise suppression - reset IL | When there is a disconnection alarm the noise suppression should reset. |  |  | Monitor PIC
MPIC SRS_071 | Disconnection alarm | The disconnection alarm should be triggered after 3 seconds not 5 as before and the values should reset. |  |  | Monitor PIC
MPIC SRS_072 | Noise suppression – Reset time change | Noise suppression should be reset after 6 seconds instead of 15. |  |  | Monitor PIC
MPIC SRS_073 | Flow measurement – High frequency-noise | Noise with direction of flow less than 40 ms in one direction must be ignored and may not have impact on the noise suppression reference volume. |  |  | Monitor PIC
MPIC SRS_074 | FOT - Dead band | The software shall disable the dead band during FOT (both in CPAP and HFO) and enabled during normal ventilation. |  |  | Monitor PIC
MPIC SRS_075 | Flow graph in PED mode | Flow graph shall display the expected waveforms in PED mode above 1600 meter setting. |  |  | Monitor PIC
MPIC SRS_076 | Flow calibration - Bar placement | The calibration bars should be displayed at the rightmost edge of the calibration symbol when flow sensor is turned off. |  |  | Monitor PIC
PPIC_SRS_001 | Power on/off button stuck | When the power on/off button is pressed for 1 to 2 seconds, then released for 1 to 5 seconds, then pressed again for at least 5 seconds, then software shall force shutdown the ventilator. Note: this is applicable only when power on/off button stuck. |  |  | Power PIC
PPIC_SRS_002 | Power PIC over-voltage protection latch | The software shall test the over-voltage protection latch at power-up. |  |  | Power PIC
PPIC_SRS_003 | Battery power lamp operation | The software shall illuminate battery LED in yellow, when the ventilator is operating on battery power. |  |  | Power PIC
PPIC_SRS_004 | Battery charging lamp operation | When ventilator is on AC main supply, the software shall illuminate the battery LED in blinking green to indicate battery is charging, irrespective of ventilator is ON or OFF. |  |  | Power PIC
PPIC_SRS_005 | Battery fully charged lamp operation | When ventilator is on AC main supply, the software shall illuminate the battery LED as steady green, when the battery is fully charged, irrespective of ventilator is ON or OFF. |  |  | Power PIC
PPIC_SRS_006 | Mains power lamp operation | The software shall illuminate main power LED in steady green, when the ventilator is connected to wall AC power. |  |  | Power PIC
PPIC_SRS_007 | Battery not charging lamp operation | The software shall disable the battery LED illumination, when the battery is not charging state. |  |  | Power PIC
PPIC_SRS_008 | Battery recharging status | The software shall continuously monitor the battery charging status and compute the remaining time to charge battery 100%. |  |  | Power PIC
PPIC_SRS_009 | System power on/off control. | The software shall control the system power on/off, when power button is turned on or off. |  |  | Power PIC
PPIC_SRS_010 | Enables and monitors DC power | The software shall enable and monitor DC power supplies for +24V, +12V, and +5V when the ventilator is powered ON. |  |  | Power PIC
PPIC_SRS_011 | Power source selection | The software shall monitor the power sources and select the battery source only if wall AC supply is not available. |  |  | Power PIC
PPIC_SRS_012 | Battery parameters | The software shall monitor and report the below battery parameters to NetDCU. • Voltage of the battery in millivolts • Discharge current of the battery in milliamps averaged for 8.2 second • Charge current of the battery in milliamps averaged for 8.2 second • Temperature of the battery in 1/10 degree Celsius • Remaining runtime with current discharge current in minutes • Remaining runtime with current discharge current in percentage • Battery capacity in milliamp/hour as determined by either last calibration cycle or default value • The number of discharge cycles of the battery in 1/10 cycle. |  |  | Power PIC
PPIC_SRS_013 | NetDCU command-shutdown | The software shall shutdown the ventilator upon receiving the power off command (99) from NetDCU via I2C. |  |  | Power PIC

## Au
Category | Sub_category | No. | Desc
Market | Voice of customer | voc-1 | Pre-op planning is a critical step to ensure safety, efficacy and efficiency of resection procedures.
Market | Voice of customer | voc-2 | Pre-op planning is intended to inform 3 clinical decisions as framed by surgeons: “where to enter”, “how to move” and the “amount of resection”.
Market | Voice of customer | voc-3 | Specifically, when reviewing pre-operative images, surgeons are looking for: 1. Size, location and vascularity of the tumor(s)
Market | Voice of customer | voc-4 | Larger tumors are more concerning for spread. Anything above 6 or 7 cm might mean a full removal and how it sits on the kidney in relationship to other organs.
Market | Voice of customer | voc-5 | Where is the cancer? Has it invaded the outside wall? Where has it invaded? What structures have been invaded?
Market | Voice of customer | voc-6 | My focus is to see the proper anatomy of the mass and its vascular anatomy
Market | Voice of customer | voc-7 | Specifically, when reviewing pre-operative images, surgeons are looking for: 2. Tumor relationship to surrounding critical structures (i.e. vasculature, ureter, chest wall, etc.)
Market | Voice of customer | voc-8 | I need to avoid critical structures like the ureter. Knowing where it is can speed things up for us.
Market | Voice of customer | voc-9 | I’m looking for the relationship of tumor to surrounding structures, particularly the vasculature. The hepatic veins and vena cava are the bigger ones for liver
Market | Voice of customer | voc-10 | Is it close to the vessel? Or is it infiltrated?
Market | Voice of customer | voc-11 | Specifically, when reviewing pre-operative images, surgeons are looking for: 3. Patient-specific anatomical variations or aberrant anatomy.
Market | Voice of customer | voc-12 | There is a great deal of variety within the anatomy. This is what we're looking for in the CT
Market | Voice of customer | voc-13 | Because of the anatomical variability, we are often operating blindly. We don't know how many bronchi we will come into.”
Market | Voice of customer | voc-14 | By integrating 3D models and patient simulation from Visible Patient into J&J ecosystem, we’ll provide the following value to the clinicians:
Market | Voice of customer | voc-15 | we’ll provide the following value to the clinicians: Highly detailed, virtual 3D model, using patients’ pre-op images (CT, MRI)
Market | Voice of customer | voc-16 | we’ll provide the following value to the clinicians: Color-coded, interactive experience that allows you to explore and visualize spatial relationships between tumors, critical structures and surrounding anatomy of your patient
Market | Voice of customer | voc-17 | we’ll provide the following value to the clinicians: Easy-to-use interface that is accessible on your team’s PC/Mac, tablet, and mobile devices for reference in the OR, at home, and on-the-go
Market | Voice of customer | voc-18 | we’ll provide the following value to the clinicians: Allows you to explore anatomy on your schedule, without reliance on interpretation by your radiology team
Market | Voice of customer | voc-19 | we’ll provide the following value to the clinicians: A virtual clip-applying tool that pre-operatively identifies where tumors are located within the anatomy and provides a real-time assessment of post-operative organ volume
Market | Voice of customer | voc-20 | Surgeons use 2D images to “envision a 3D model” in their mind when planning their approach. More experienced surgeons are more likely to be confident in this ability. However, even experienced surgeons acknowledge inefficiency and communication challenges with the surgical team when developing a surgical plan, e.g. reliance on the radiologist to develop a surgical plan.
Market | Voice of customer | voc-21 | In addition to 3D models, 2D visualization, with color, of a 3D model (Anatomical Atlas) is also a valuable tool for clinicians as it is “what I’m used to and know how to interpret, but can see things better”.
Market | Voice of customer | voc-22 | The patient simulation (Clip Applying) tool provides significant value when planning painstaking, tissue-sparing procedures. For many, the patient simulation tool is considered a “breakthrough” because of high-quality image and instant volume calculations that allow for simulation of different surgical plans pre-operatively (e.g. segmentectomies vs a lobectomy).
Market | Voice of customer | voc-23 | In addition to improving efficiency of pre-op planning with other clinicians, 3D modeling and patient simulation are seen as valuable tools for communicating the procedure to a patient or a trainee.
Market | Voice of customer | voc-24 | Surgeons see the flexibility in access across devices and locations (at home, at the office, on the go, in the OR) as a critical attribute of the solution
Market | Voice of customer | voc-25 | All surgeons for the target specialties (thoracic, HBP, colorectal and urology) are procedurally required to review pre-operative patient images before the procedure. However, only a subset of surgeons refers back to pre-operative patient images intra-operatively, mostly due to an unexpected finding or uncertainty. Hence, the primary use case for pre-operative planning is before the procedure begins; intraoperative reference of the preoperative planning is secondary.
Market | program goals |  | The primary clinical goals of this project are to design, develop and deploy a solution that will enable Digital Solution to ● launch surgical decision support, especially medical imaging software application products
Market | program goals |  | enable surgeons to order, manage, and review patient-specific 2D and 3D anatomical models
Market | program goals |  | enable surgeons to view pre-operative images and 2D and 3D anatomical models online
Market | program goals |  | enable surgeons to perform a simulated resection by applying virtual clips to approximated segments in order to simulate post-operative organ volume
Market | program goals |  | enable surgeons to save a viewing and simulation session and continue at a later time
Market | program goals |  | enable surgeons to continue their viewing and simulation session in the OR using the Digital Hub hardware device after the commercial release of the Digital Hub (viewing of data sets on the Digital Hub will be outlined in a future Charter),
Market | program goals |  | enable surgeons to continue viewing their images and simulation in the OR on the web portal,
Market | program goals |  | release an integrated solution to the EU and/or US markets late 2021 or early 2022
Market | program goals |  | release an integrated solution in APAC later in 2022 as outlined in a future Charter,
Market | program goals |  | expand and engage the overall user community post-launch.
Market | Product Description |  | Bandit Phase 2 should be comprised of the following major components: surgeon portal, digital solutions viewer, digital surgery enablement platform.
Market | Product Description |  | Surgeon Portal - a web portal for surgeons to log into, select the appropriate language, order the modeling service, review the status of the modeling service, upload DICOM images, start the Digital Solutions Viewer.
Market | Product Description |  | Digital Solutions Viewer - a DICOM viewer that allows for viewing a DICOM image and/or the 3D anatomical model selected by the surgeon through the surgeon portal. After the surgeon confirms the selected dataset, the image viewer will show the pre-operative image, allow for placement of annotations and measurements, show the anatomical model, allow for interactions with the anatomical model, allow for simulation of resection by placing a virtual clip.
Market | Product Description |  | Digital Surgery Enablement Platform (DSEP) - a cloud infrastructure that enables user management, data storage, and provides DICOM studies to the external 3D model creation service as well as to the Digital Solutions Viewer.
Market | Product Description |  | Bandit Phase 2 will deliver an integrated experiences for the customer, which can be categorized in three groups: ● 3D Model Submission and Delivery ● Digital Solution Viewer ● Patient Modeling and Outcome Simulation
Market | Product Description |  | AI-enhanced visualization and personalized insights to help surgeons get a clearer roadmap for surgery
Market | Product Description |  | 3D model submission and delivery: 3D model order submission by clinical user; 3D model order management by order admin; 3D model order billing; PACS integration
Market | Product Description |  | Image viewer: image viewer; meaurement tools
Market | Product Description |  | patient modeling and outcome simulation: integrated 3D models; patient specific simulation of resection
Market | Regulatory assessment |  | Target for the Digital Solutions Viewer is the EU. The Digital Solutions Viewer will be the first CE marked product and first product under the new EU Medical Device Regulation (MDR) for Auris.
Market | Regulatory assessment |  | The Digital Solutions Viewer is considered software as a medical device (SaMD) and will need to conform to the requirements of the Medical Device Regulation (EU) 2017/745. The Viewer will need to provide Technical Documentation (TD) covering the design, development, and state of the art with respect to safety and performance and demonstrate compliance to the General safety and Performance requirements in Annex I.
Market | Regulatory assessment |  | The proposed classification of the Digital Solutions Viewer is a Class IIa Rule 11 (Software intended to provide information which is used to make decisions with diagnosis or therapeutic purposes).
Market | Regulatory assessment |  | Type of conformity assessment route: Quality Management System, Annex IX excluding chapter II.
Market | Regulatory assessment |  | As the Digital Solutions viewer will be the first MDR product for Auris, the Auris QMS must be MDR ready to support certification. The proposed certificate scope will be for image visualization and planning software.
Market | Regulatory assessment |  | An EU Authorised Representative agreement between Auris and J&J Medical GmbH must be in place.
Market | Regulatory assessment |  | The MDR stipulates that the manufacturer must provide sufficient clinical evidence (Article 61) to justify the level of clinical evidence necessary to demonstrate conformity with the relevant GSPR.
Market | Regulatory assessment |  | The MDR provides a possibility to use clinical data related to a marketed equivalent device (found in the clinical evaluation required for a device) under conformity assessment.
Market | Regulatory assessment |  | Visible Patient Planner will be used as a comparator for the Digital Solutions Viewer as the Visible Patient Planner is CE marked and the technical characteristics, functional principle of software algorithms and clinical performance and intended purposes of the software are intended to be similar for the Auris Viewer.
Market | Regulatory assessment |  | The equivalence requirement to demonstrate sufficient levels of access to comparator data should be met as Visible Patient is collaborating with J&J and the Auris Viewer is intended to interface with the Visible Patient Lab 3D recreated models.
Market | Regulatory assessment |  | Visible Patient Planner will be used as a comparator for the Digital Solutions Viewer as the Visible Patient Planner is CE marked and the technical characteristics, functional principle of software algorithms and clinical performance and intended purposes
Market | Market overview |  | Market Target Segments: ● Cutting Edge Pioneers: Large, urban, teaching hospitals leading the way through technological innovation with a large budget.
Market | Market overview |  | Competitive Strivers: Moderately sized, urban, teaching hospitals that intentionally spend their resources to differentiate from the competition.
Market | Market overview |  | Source of Volume (SOV): Attract potential category users
Market | Market overview |  | Hospital Targets: High procedure volumes in 2+ of the following surgical specialties: Thoracic, Urology, Colorectal, and HPB; typically, urban and / or academic hospitals
Market | Market overview |  | Primary User Targets: Surgeons in Thoracic, Urology, Colorectal, and Hepato-biliary
Market | Market overview |  | Secondary User Targets (secondary expansion within primary user hospitals): Lower volume surgeons in Primary User Target specialties, Surgeons in GYN
Market | Market overview |  | Bandit opportunity immediately establishes an Ethicon position in Advanced Imaging
Market | Market overview |  | Ethicon Franchise: Aligns to Ethicon franchise objectives of developing connected products & solutions and leveraging data-driven insights to improve outcomes.
Market | Market overview |  | Advanced Imaging: Establishes an immediate Ethicon position in Advanced Imaging category, the “North Star” for the Ethicon Digital Solutions portfolio.
Market | Market overview |  | Blunts Competitive Messaging prior to JNJ Robot launch: Bandit blunts ISRG IRIS™ launch into Advanced Imaging, the #1 driver of robot brand choice
Market | Market overview |  | When the unified global Digital Solutions user portal launches in H2 2021, Bandit expands the value of Digital Solutions to Ethicon:
Market | Market overview |  | Builds Loyal Digital Surgery Customer Base: Increases the value of the online Ethicon Digital Solutions video, image & case library, creating network effects & switching costs
Market | Market overview |  | Provides Critical Data Pipeline: Adds pre-op image and 3D model data to feed the Ethicon AI Insights platform
Market | Market overview |  | Pre-operative 3D planning software is in-market today, but most focus on orthopaedics and neurosurgery. Bandit has three main competitors in the US (Fujifilm Synapse, TeraRecon iNtuition, and Intuitive Surgical IRIS™), but only Fujifilm and TeraRecon have a significant presence in EMEA, Bandit’s initial regional foothold market. Only Fujifilm is used today in APAC.
Market | Market overview |  | Fujifilm and TeraRecon require radiologist creation of 3D models and involvement with interpretation and dictation, whereas Bandit eliminates this step, giving surgeons a streamlined pre-op planning sequence that puts the control in their hands on their schedule.
Market | Market overview |  | Intuitive IRIS™ is currently cleared for use with CT only, whereas the others, including Bandit, are applicable for MRI as well.
Market | revenue and cost |  | revenue and cost: Bandit is projected to create ~€200MM in cumulative end user sales through 2025 with peak sales of €75+MM with launches in EMEA, APAC and NA.
Market | Project assumptions |  | An initial project team will be deployed to lead activities needed to complete the deliverables of Phase 1: Planning and Definition Phase per document 700-000049-00, Design Control. The Design and Development Plan (DDP) will ultimately define the organizational and team structure of the project and will be completed prior to exiting the Planning and Definition Phase.
Market | Project assumptions |  | The major project milestones and estimated timing are listed below. This schedule assumes a fully staffed development team and approved program budget by the end of Phase 1. An updated schedule and list of major milestones along with program risks will be provided in the Design and Development Plan (DDP) prior to exiting the Planning and Definition Phase.
Market | Project assumptions |  | The following assumptions were made as basis of this Project Authorization Charter: ● CE Mark is prioritized (EMEA launch) and 510K Approval is performed afterwards
Market | Project assumptions |  | Changes to UI of Digital Solutions Viewer are limited to meet timeline
Market | Project assumptions |  | Contract technical writer (w/SaMD exp) is on-board by mid-January 2021 for DHF support
Market | Project assumptions |  | CIR, DIR (and lower) are written to minimize V&V and released no later than February 2021
Market | Project assumptions |  | No issues are found in the 1st formative; and a 2nd formative is not required
Market | Project assumptions |  | Expedited Review (MDR) has minimal questions and no major revisions needed
Market | Project assumptions |  | FDA acceptance of substantial equivalence with minimal questions
Market | Project assumptions |  | Code the complete the Digital Solutions Viewer is solid with few bugs/issues by mid-March 2021
Market | regulatory |  | Web Portal & DSEP are not part of MDR submission
Market | regulatory |  | QMS audit can completed by regulatory body by end of May (earlier than planned)
Market | regulatory |  | QMS documents can be reviewed and released in Agile within 2 weeks (documents complete to released)
Market | regulatory |  | No issues found during phase reviews or phase exit reviews
Market | regulatory |  | Notified body is available for an expedited review and we meet the date provided to them
Clinical_Workflow | Clinical Hazard Analysis | CHA-01 | The intended use of the device is an aid in pre-surgical planning for procedures requiring spatial understanding of complex patient anatomy.
Clinical_Workflow | Clinical Hazard Analysis | CHA-02 | The medical device can be used to help the surgeon understand size and location of tumors, abnormalities, and/or defects, and create the surgical plan.
Clinical_Workflow | Clinical Hazard Analysis | CHA-03 | Patients undergoing procedures in the general, bariatric, HPN, gynecology, thoracic, and urology specialties, who require medical images to be taken to plan for their surgical procedure.
Clinical_Workflow | Clinical Hazard Analysis | CHA-04 | No, the medical device does not sustain or support life.
Clinical_Workflow | Clinical Hazard Analysis | CHA-05 | If the medical device fails while being used to reference medical images or 3D models during a procedure, the user may need to access the patient images on another PACS viewer to proceed.
Clinical_Workflow | Clinical Hazard Analysis | CHA-06 | Measurements are taken by the medical device. The medical device can take linear measurements between two points displayed on the DICOM or 3D model
Clinical_Workflow | Clinical Hazard Analysis | CHA-07 | Measurements are taken by the medical device. The medical device can calculate a designated area shown on a 2D DICOM slice
Clinical_Workflow | Clinical Hazard Analysis | CHA-08 | Measurements are taken by the medical device. The medical device can through the clip apply feature, calculate volume of a designated portion of an organ to be resected
Clinical_Workflow | Clinical Hazard Analysis | CHA-09 | Measurements taken on a 2D device should be accurate for images acquired with high resolution.
Clinical_Workflow | Clinical Hazard Analysis | CHA-10 | The 3D measurements and/or volumes should be considered estimates, as the anatomic models are reconstructions based on the segmentation algorithm and spacing of the 2D slices.
Clinical_Workflow | Clinical Hazard Analysis | CHA-11 | The medical device may be used with a radiology monitor
Clinical_Workflow | Clinical Hazard Analysis | CHA-12 | The medical device contains software. The software is intended to be installed by the client rep, user, or IT specialist. Periodic updates may be required to install latest version of the software.
Clinical_Workflow | Clinical Hazard Analysis | CHA-13 | The lifetime of the medical device may be determined by the software lifecycle and ability to maintain compatibility with common operating systems. The medical device must run on computer hardware, so it inherits that risk
Clinical_Workflow | Clinical Hazard Analysis | CHA-14 | When using a tablet to interface with the device during a procedure, there is a risk that the procedure would coincide with battery depletion.
Clinical_Workflow | Clinical Hazard Analysis | CHA-15 | successful application of the medical device is critically dependent on human factors such as the user interface. The user interface is fundamental to interacting with and displaying measurements of the patient images and anatomic models.
Clinical_Workflow | Clinical Hazard Analysis | CHA-16 | use errors can arise from not understanding or properly navigating the user interface. Symbols for the user interface have been based on those common in other PACS viewer products, but may be arranged differently or applied slightly differently for this medical device.
Clinical_Workflow | Clinical Hazard Analysis | CHA-17 | In general, the medical device is used in the office environment for planning prior to a case. When using the medical device in the operating room, the user may be more distracted and could make use errors such as loading the incorrect patient’s data, or ignoring the orientation indicators.
Clinical_Workflow | Clinical Hazard Analysis | CHA-18 | The medical device displays information. The display relies on text, colors, sizing, etc to relay spatial information and distinguish tissue types. These can be adjusted for accessibility but must take colorblind users into account
Clinical_Workflow | Clinical Hazard Analysis | CHA-19 | the medical device must be used with a display large enough and with high enough resolution so that the user can clearly see details of the patient images and anatomic models.
Clinical_Workflow | Clinical Hazard Analysis | CHA-20 | The device must render the patient models taking all these factors into account.
Clinical_Workflow | Clinical Hazard Analysis | CHA-21 | the medical device is controlled by a menu. The device requires simple menu navigation for loading the correct patient information. Otherwise, there is not a specific sequence required and the user will access the different functions as desired. Most settings are accessible at all times on a sidebar.
Clinical_Workflow | Clinical_procedure | CPR_01 | User shall launch program by openning Aurora viewer during the selecting an order phase
Clinical_Workflow | Clinical_procedure | CPR_02 | User shall launch program by navigation to "OVERVIEW" tab which is shown at the top of the screen during the selecting an order phase
Clinical_Workflow | Clinical_procedure | CPR_03 | User shall find patient(s) of interest by searching for patient by name, MRN, surgeon during the selecting an order phase
Clinical_Workflow | Clinical_procedure | CPR_04 | User shall find patient(s) of interest through sorting lists by specialty, patient, surgeon, delivery date, or status to find patient of interest during the selecting an order phase
Clinical_Workflow | Clinical_procedure | CPR_05 | User shall find patient(s) of interest through filters list using "FILTER BY" tool in order to find patient during the selecting an order phase
Clinical_Workflow | Clinical_procedure | CPR_06 | User shall select patient and launch image viewer by touching/clicking on row and row will highlight when selected during the selecting an order phase
Clinical_Workflow | Clinical_procedure | CPR_07 | User shall select patient and launch image viewer by clicking "NEXT" to load patient images and 3D model during the selecting an order phase
Clinical_Workflow | Clinical_procedure | CPR_08 | User shall verify images are correct by double checking patient informaiton during the selecting an order phase
Clinical_Workflow | Clinical_procedure | CPR_09 | User shall view 2D image stacks by monitoring loading of the 2D images on the "Viewing" panel during the load and view desired image stacks phase
Clinical_Workflow | Clinical_procedure | CPR_10 | After the images are loaded, the option to toggle "2D Source Image" appears. User toggles "2D Source Image" to view the 2D image stacks.
Clinical_Workflow | Clinical_procedure | CPR_11 | User shall select which of the following views they would like to see, by using the checkboxes that now appear under "2D Source Image" menu: Axial, Sagittal, Coronal, Anatomical Atlas during setting screen layout phase
Clinical_Workflow | Clinical_procedure | CPR_12 | User shall select whether they would like the segmented anatomic regions to be displayed by infill or contour when turning on and off segmentation layers
Clinical_Workflow | Clinical_procedure | CPR_13 | User shall adjust view of desired image stacks by operating pan, rotate, zoom, and scroll to select stack they would like to manipulate during the adjust view of desired image stacks phase
Clinical_Workflow | Clinical_procedure | CPR_14 | User shall use touchscreen, mouse, or side menu to select "Pan, Rotate, Zoom or Scroll"
Clinical_Workflow | Clinical_procedure | CPR_15 | User shall undo a single adjustment with the "Undo" button during the adjust view of desired image stacks phase
Clinical_Workflow | Clinical_procedure | CPR_16 | Usershall reset the view altogether with the "Re-Center" button during the adjust view of desired image stacks phase
Clinical_Workflow | Clinical_procedure | CPR_17 | User shall navigate to the Windowing menu on the sidebar to asjust the contrast or brightness of the image during the adjust view of desired image stacks phase
Clinical_Workflow | Clinical_procedure | CPR_18 | User shall select from presets to auto-adjust window width and level based on target anatomy the image was acquired to show during the adjust view of desired image stacks phase
Clinical_Workflow | Clinical_procedure | CPR_19 | User shall choose to select "Width/Level" button to make customize adjustments to the brightness and contrast during the adjust view of desired image stacks phase
Clinical_Workflow | Clinical_procedure | CPR_20 | User shall drag slider(s) to desire width and level percentages when adjust windowing during the adjust view of desired image stacks phase
Clinical_Workflow | Clinical_procedure | CPR_21 | User shall reset the view using the "Reset" button when adjust windowing during the adjust view of desired image stacks phase
Clinical_Workflow | Clinical_procedure | CPR_22 | User shall select the annotation menu by navigate to tools from the sidebar during the annotation tool phase
Clinical_Workflow | Clinical_procedure | CPR_23 | User shall select "distance" tool to measure the straight-line distance between two points on the image during the annotation tool phase
Clinical_Workflow | Clinical_procedure | CPR_24 | User shall click on the image where measured distance should start, and then clicks again where measured distance should end. Linear distance will appear onscreen during the annotation tool phase
Clinical_Workflow | Clinical_procedure | CPR_25 | User shall drag the endpoints to resize the line, or drag the line itself to reposition the line on the image by using distance tool during the annotation tool phase
Clinical_Workflow | Clinical_procedure | CPR_26 | User shall delete an added line by clicking the trash icon next to the line name in the Annotation sidebar menu by using distance tool during the annotation tool phase
Clinical_Workflow | Clinical_procedure | CPR_27 | User shall add and delete arrows in a similar fashion after clicking the "Arrow" menu option. Arrows added to the screen will also appear in the Annotation sidebar menu list by arrow tool during the annotation tool phase
Clinical_Workflow | Clinical_procedure | CPR_28 | User shall add a note by selecting the "note" annotation option and clicking anywhere on the image during the annotation tool phase
Clinical_Workflow | Clinical_procedure | CPR_29 | when a textbox appears, user may type in any notes about the case or image. The note will also appear in the Annotation sidebar menu during the annotation tool phase
Clinical_Workflow | Clinical_procedure | CPR_30 | User shall return to a previously added note by selecting it from the Annotation menu, and clicking the "Edit" button on the textbox during the annotation tool phase
Clinical_Workflow | Clinical_procedure | CPR_31 | User shall remove any added note using the trash icon in the Annotation menu or on the note itself during the annotation tool phase
Clinical_Workflow | Clinical_procedure | CPR_32 | User shall select any of the annotation distances or arrows from the Annotation menu list and rename them by rename annotation during the annotation tool phase
Clinical_Workflow | Clinical_procedure | CPR_33 | User shall reset all annotations using the "Reset" button during the annotation tool phase
Clinical_Workflow | Clinical_procedure | CPR_34 | User shall toggle off "2D Source Image" and toggle on "3D Model" in the Viewing menu when openning 3D Model View during the visualizing and manipulating 3D model phase
Clinical_Workflow | Clinical_procedure | CPR_35 | User shall confirm they ar viewing the correct 3D model by checking the information shown onscreen. They togge on "Patient Info" if not already on when openning 3D Model View during the visualizing and manipulating 3D model phase
Clinical_Workflow | Clinical_procedure | CPR_36 | User shall select "3D Model" menu to see which anatomic structures are available when selecting 3D models to display during the visualizing and manipulating 3D model phase
Clinical_Workflow | Clinical_procedure | CPR_37 | User shall select/deselect the checkboxes next to the structure they want to see, or "ALL" if they would like to see all that are available when selecting 3D models to display during the visualizing and manipulating 3D model phase
Clinical_Workflow | Clinical_procedure | CPR_38 | User shall scan the volumes of the different models to check against expectations when selecting 3D models to display during the visualizing and manipulating 3D model phase
Clinical_Workflow | Clinical_procedure | CPR_39 | User shall select "Pan, Rotate, Zoom" and/or uses the touchscreen and mouse controls to achieve the view they desire that lets them see the relevant anatomical spatial relationships when setting desired orientation during the visualizing and manipulating 3D model phase
Clinical_Workflow | Clinical_procedure | CPR_40 | Alternatively, user shall select from anatomical plane icons or cube at the bottom right of the viewing window to auto-orient the model to the desired view when setting desired orientation during the visualizing and manipulating 3D model phase
Clinical_Workflow | Clinical_procedure | CPR_41 | User shall enter fullscreen at times for a better view of the model when setting desired orientation during the visualizing and manipulating 3D model phase
Clinical_Workflow | Clinical_procedure | CPR_42 | User shall select an individual structure when adjusting visibility of individual structures during the visualizing and manipulating 3D model phase
Clinical_Workflow | Clinical_procedure | CPR_43 | User shall adjust the opacity of that structure using the slider (unknown if color can be adjusted in current iteration) when adjusting visibility of individual structures during the visualizing and manipulating 3D model phase
Clinical_Workflow | Clinical_procedure | CPR_44 | User shall select "Done" button to complete the adjustment when adjusting visibility of individual structures during the visualizing and manipulating 3D model phase
Clinical_Workflow | Clinical_procedure | CPR_45 | User shall select "Clip Apply Model" menu and allows module to load when navigating to Clip Apply feature during applying clips to simulate resection volume phase
Clinical_Workflow | Clinical_procedure | CPR_46 | User shall manipulate the model as necessary to establish the correct orientation to apply a clip when adding and removing Clips during applying clips to simulate resection volume phase
Clinical_Workflow | Clinical_procedure | CPR_47 | User shall right click on the model the apply the clip when adding and removing Clips during applying clips to simulate resection volume phase
Clinical_Workflow | Clinical_procedure | CPR_48 | User shall observe the resected organ volume and healthy tissue volume. If they are happy with the clip placement they may move on to place another clip when adding and removing Clips during applying clips to simulate resection volume phase
Clinical_Workflow | Clinical_procedure | CPR_49 | If the user is unhappy with the clip placement they chose, they shall either remove a single clip, or remove all clips when adding and removing Clips during applying clips to simulate resection volume phase
Clinical_Workflow | Clinical_procedure | CPR_50 | At any time, the user shall select any of the clips in the menu in order to review the volume simulation details for that clip when adding and removing Clips when saving the plan during applying clips to simulate resection volume phase
Clinical_Workflow | Clinical_procedure | CPR_51 | When clicking to a different module, user will first be prompted whether to save the plan. They may choose to save or not using the two buttons when saving the plan during applying clips to simulate resection volume phase
Clinical_Workflow | Clinical_procedure | CPR_52 | If they choose to save the plan, user names the plans and confirms. (Need to check) when saving the plan during applying clips to simulate resection volume phase
Clinical_Workflow | Clinical_procedure | CPR_53 | User shall navigate to "My Views" menu when saving current view for later access during saving current view phase
Clinical_Workflow | Clinical_procedure | CPR_54 | User shall select "SAVE NEW VIEW" and pop-up appears when saving current view for later access during saving current view phase
Clinical_Workflow | Clinical_procedure | CPR_55 | User may select to either save and replace a current view, or save it as a new view (and enter a new name), then press "CONFIRM" when saving current view for later access during saving current view phase
Clinical_Workflow | Clinical_procedure | CPR_56 | User shall press "CONFIRM" a second time and the view is saved when saving current view for later access during saving current view phase
DIR | Requirement | DIR_1 | Use of software via a non-purpose built computer: the system shall provide a method for the user to interact with Image Viewer via a non-purpose built computer.
DIR | Image Viewer | DIR_2 | The system shall provide a method to open a pre-operative patient image that is compliant with DICOM standard
DIR | Image Viewer | DIR_3 | The system shall have the capability to render 2D views of the preoperative patient images.
DIR | Image Viewer | DIR_4 | The system shall have the capability to render 3D views of standard 3D reconstructions from an MRI/CT study.
DIR | Image Viewer | DIR_5 | The system shall have the capability to render a view of 3D model segmentations
DIR | Image Viewer | DIR_6 | The system shall have the capability to render a colored 2D overlay of a 3D segmentation on top of an MPR view
DIR | Image Viewer | DIR_7 | The system shall provide a method to zoom in and out on the views.
DIR | Image Viewer | DIR_8 | The system shall provide a method to pan a viewing panel.
DIR | Image Viewer | DIR_9 | The system shall provide a method to rotate a 2D viewing panel.
DIR | Image Viewer | DIR_10 | The system shall provide a method to rotate a 3D viewing panel.
DIR | Image Viewer | DIR_11 | The system shall provide a method to toggle visibility of pre-operative image data viewed in the 3D panel.
DIR | Image Viewer | DIR_12 | The system shall provide a method to rotate the cut plane within an MPR view.
DIR | Image Viewer | DIR_13 | The system shall provide a method to undo and redo user interactions.
DIR | Image Viewer | DIR_14 | The system shall provide a method for resizing a viewing panel by maximizing, minimizing, or restoring to default size.
DIR | Image Viewer | DIR_15 | The system shall provide a method for hiding a viewing panel.
DIR | Image Viewer | DIR_16 | The system shall provide a method to adjust the Window Center and Window Width adjustments using presets.
DIR | Image Viewer | DIR_17 | The system shall provide a method for manually adjusting window width and center values for the intensity range
DIR | Image Viewer | DIR_18 | The system shall provide a method for a reset all viewing panels simultaneously
DIR | Image Viewer | DIR_19 | The system shall provide a method to select the layout for the viewing panels
DIR | Image Viewer | DIR_20 | The system shall provide a method to advance the multi-planer resection (MPR) cut planes through the image volume.
DIR | Image Viewer | DIR_21 | The system shall be able to display metadata contained in the DICOM data.
DIR | Image Viewer | DIR_22 | The system shall list all objects contained in the Visible Patient model container (3D model) and allow for user selection of objects.
DIR | Image Viewer | DIR_23 | The system shall display the volumetric sizing information provided by Visible Patient for each object (segmentation) which is part of the 3D model.
DIR | Image Viewer | DIR_24 | The system shall be able to display metadata contained in the 3D model container
DIR | Image Viewer | DIR_25 | The system shall display a semi-transparent overlay of the 3D model in the 2D viewing panels
DIR | Image Viewer | DIR_26 | The system shall use SI units and their derivatives
DIR | Image Viewer | DIR_27 | The system shall provide a method for performing one or more measurements within one of the viewing panels
DIR | Image Viewer | DIR_28 | The system shall label and list measurements automatically
DIR | Image Viewer | DIR_29 | The system shall provide the ability to place an annotation with its associated marker on a displayed image or 3D object.
DIR | Image Viewer | DIR_30 | The system shall provide the ability to edit the placement and content of annotations placed on a viewing panel.
DIR | Image Viewer | DIR_31 | The system shall allow for simulation of resection by applying a clip to approximated segments.
DIR | Image Viewer | DIR_32 | The system shall display the volumetric sizing information provided by Visible Patient for each simulated clip.
DIR | Image Viewer | DIR_33 | The system shall provide the ability to support individual language packs.
DIR | Image Viewer | DIR_34 | The system display shall provide the ability to display U/I's, messages, and other language specific content in more than one language.
DIR | Image Viewer | DIR_35 | The system shall, when appropriate and available, use standard internationally recognized symbols.
DIR | Image Viewer | DIR_36 | The system shall, when appropriate and available, use standard internationally recognized icons.
DIR | Image Viewer | DIR_37 | The System shall be capable of recording all preoperative, intraoperative, and postoperative functions performed by the System.
DIR | Image Viewer | DIR_38 | The System shall record all functions performed by the System pre, intra, and post operatively.
DIR | Image Viewer | DIR_39 | The System should record all functions performed by the System that are relevant to monitoring, diagnosis and repair of the System by service personnel.
DIR | Image Viewer | DIR_40 | The System shall record system functions and operational events generated by the System.
CIR | Image Viewer | Digital_TXT_11 | The purpose of this document is to define Customer Input Requirements (CIR), including clinical and business requirements, for Digital Solutions Viewer that enables pre-op planning.
CIR | Image Viewer | Digital_TXT_12 | The scope of this document covers the CIR for Release 1 of the Digital Solutions Viewer, which includes clinical pre-operative planning software used by the clinical user to create a pre-operative plan based on the completed 3D models and underlying pre-op patient images.
CIR | Image Viewer |  | The scope of this document does not cover how 3D model are ordered, created and delivered, as this is accomplished by other software components.
CIR | Image Viewer | Digital_TXT_13 | All references to specific policies, procedures, or work instructions in this document are to the latest revision of that document unless otherwise noted. All references to personnel in this document are to that person or his/her designated representative unless otherwise stated.
CIR | Image Viewer | Digital_TXT_14 | 3D Model: Anatomy segmentations, related measurements, and metadata created based on pre-op patient images by the 3D model creator.
CIR | Image Viewer |  | 3D Model Order: A request submitted by a clinical user to create a 3D model, which contains patient images and order metadata, e.g. selected model type, patient diagnosis information and any additional order-specific request from the surgeon.
CIR | Image Viewer |  | Application: The system comprised of 3D model ordering software, clinical pre-op planning software and underlying infrastructure components.
CIR | Image Viewer |  | DICOM: Digital Imaging and Communications in Medicine is the standard for the communication and management of medical imaging information and related data.
CIR | Image Viewer |  | PACS: Picture Archiving and Communication System, a system used in medical imaging to store, retrieve, distribute, analyze, and digitally process medical images.
CIR | Image Viewer |  | Pre-Op Patient Images: Pre-operative patient images such as CT and MRI studies in a DICOM format.
CIR | Image Viewer |  | Visible Patient (VP): External company, responsible for managing and delivering 3D models.
CIR | Image Viewer |  | Clinical User: A surgeon of any experience-level within the target specialties (e.g. THX, HPB, URO, COL) or surgeon’s assistant. This user intends to perform the following tasks: • create a 3D model order request • monitor 3D model status and address comments/problems • access/ download completed 3D models and report • develop pre-op plan • reference pre-op plan intraoperatively
CIR | Image Viewer |  | 3D Model Service Provider: An internal Visible Patient representatives (order admin) that manages 3D model orders and facilitates communication between the clinical user and the 3D model creator (annotator). 3D model service provider (order admin) intends to perform the following tasks: • update 3D model order status, problem indicator and comments • download de-identified 3D model orders • upload completed 3D models
CIR | Image Viewer |  | Healthcare Provider: A collective term to refer to clinical and non-clinical user associated with a hospital.
CIR | Image Viewer |  | Marketing and Sales Representative: Internal J&J users responsible for marketing and selling the product and monitoring key performance indicators related to product use to improve the product, price, place and promotion strategies.
CIR | Image Viewer | Digital_CIR_DS-1 | View Pre-Op Images: As a clinical user I need to manipulate visualization parameters associated with pre-op patient images to improve visualization of the anatomy and pathology of interest so that I can develop a patient-specific surgical plan and anticipate intra-op findings.
CIR | Image Viewer | Digital_CIR_DS-131 | Manipulate Pre-Op Images: As a clinical user I need to take point to point measurements using a 2D view of the pre-op patient images and a 2D view of a 3D model, if the 3D model is available so that I can develop a patient-specific surgical plan and anticipate intra-operative findings.
CIR | Image Viewer | Digital_CIR_DS-91 | 2D Measurements: As a clinical user I need to take point to point measurements using a 2D view of the pre-op patient images and a 2D view of a 3D model, if the 3D model is available so that I can develop a patient-specific surgical plan and anticipate intra-operative findings
CIR | Image Viewer | Digital_CIR_DS-5 | View 3D Models and Associated Volumes: As a clinical user I need to view 3D models of anatomy and pathology and associated volumes so that I can develop a patient-specific surgical plan and anticipate intra-op findings.
CIR | Image Viewer | Digital_CIR_DS-132 | Manipulate 3D Model: As a clinical user I need to manipulate visualization parameters associated with 3D models to improve visualization of the target anatomy and pathology so that I can develop a patient-specific surgical plan and anticipate intra-op findings
CIR | Image Viewer | Digital_CIR_DS-92 | Volume Calculations and Simulation of Resection: As a clinical user I need to access calculated volumes of the target anatomy and pathology so that I can simulate resection to determine the surgical approach.
CIR | Image Viewer | Digital_CIR_DS-93 | Configuring Display of Pre-op Images: As a clinical user I need to access and configure display of pre-op patient images and associated 3D models so that a clinical user can reference pre-op patient images and associated information as need arises before or during the procedure.
CIR | Image Viewer | Digital_CIR_DS-12 | Compatibility with OR Displays: As a clinical user I need the application to enable viewing of pre-op patient images and associated 3D models from existing OR displays so I can see all relevant patient information at the same time and minimizes interruptions/delays of having to reference multiple screens.
CIR | Image Viewer | Digital_CIR_DS-14 | Compatibility with Sterile and Non-Sterile Environment: As a clinical user I need the application to enable control of visualization parameters for pre-op patient images and 3D models from sterile and non-sterile environment so that I can minimize interruptions to my clinical workflow and reduce reliance on other personnel to control the device.
CIR | Image Viewer | Digital_CIR_DS-104 | Secure access to patient data: As a healthcare provider, I need to ensure a secure and compliant access to personal and patient data, so I can meet cyber-security and data privacy requirements of the region, country and state where I practice.
CIR | Image Viewer | Digital_CIR_DS-24 | Localization of Software: As a clinical user, I need to access the solution (clinical user-facing software) in the language I understand, so that I use the system correctly and efficiently.
CIR | Image Viewer | Digital_CIR_DS-109 | Indications for Use for 3D Model Viewing and Manipulation: As Ethicon marketing representative, I need the application to support indications for use of the 3D model service provider so that I can ensure that the indication for use is in alignment with indications for 3D models and related functionality supplied by the third party.
CIR | Image Viewer | Digital_CIR_DS-110 | Countries Supported for 3D Model Viewing and Manipulation: As Ethicon marketing representative, I need the application to be available in countries and geographic regions where 3D model service provider has obtained regulatory clearance and registration so that I can maximize customer and business value in according to the business plan/model.
CIR | Image Viewer | Digital_CIR_DS-113 | Capture and Storage of System Data: As Ethicon marketing representative, I need the application to capture and store application configuration/settings, user entered data and logs in a long term storage in accordance with data privacy regulations and agreements, so that I can enable use of the data for product improvement, development.
CIR | Image Viewer | Digital_CIR_DS-114 | Branding: As Ethicon marketing representative, I need the application to follow JnJ defined branding, so that I can maximize consistency in the look and feel of JnJ digital products that may be used by the same user base and reduce cognitive load on the users when using JnJ digital products.
CIR | Image Viewer | Digital_CIR_DS-115 | Demo Capabilities: As Ethicon marketing and sales representative, I need to be able to demonstrate the application to hospital stakeholders, so that I can communicate the value and expected workflow of the application effectively.
CIR | Image Viewer | Digital_CIR_DS-116 | Trial Capabilities: As Ethicon marketing and sales representative I need to be provide a trial functionality to the application so that I can enable clinical users evaluate the value and workflow of the application for their specific cases
Clinical_HA | Image Viewer |  | Major hepatic hemorrhage is an understood risk of performing a minimally invasive Hepatectomy procedure. Major hemorrhage of hepatic vessels during dissection or parenchymal division.
Clinical_HA | Image Viewer |  | Liver failure following surgery: Postoperative liver failure is an understood risk of performing a minimally invasive Hepatectomy procedure.
Clinical_HA | Image Viewer |  | Central Nervous System Event (CNS): Referring to any incidents that may cause damage to the CNS. CNS event is an understood risk of performing a laparoscopic colorectal procedure.
Clinical_HA | Image Viewer |  | Systemic Sepsis: A complication of an infection where chemicals released into the bloodstream to fight the infection trigger inflammatory responses throughout the body.
Clinical_HA | Image Viewer |  | Nerve Injury: Nerve injury – leading to incontinence or sexual dysfunction. Nerve Injury is an understood risk of performing a laparoscopic urologic procedure.
Clinical_HA | Image Viewer |  | Peripheral Neuropathy: Nerve damage (most often related to obturator nerve) cause by mechanical trauma to nerve fibers. Peripheral neuropathy is an understood risk of performing a laparoscopic urologic procedure.
Clinical_HA | Image Viewer |  | Acute Bleeding: Acute bleeding can lead to a hematoma or pooling of blood creating the appearance of bruising if visible. Acute bleeding is an understood risk of performing a laparoscopic urologic procedure.
Clinical_HA | Image Viewer |  | Major Vascular Injury can occur due to surgical manipulation of the pulmonary vessels. Aberrant pulmonary vasculature is well-documented in the literature. Major damage is a catastrophic complication and can lead to swift death due to blood loss. Major Vascular Injury is an understood risk of performing a minimally invasive Pulmonary Resection procedure.
Clinical_HA | Image Viewer |  | Bronchial Injury: Bronchial Injury can occur due to surgical manipulation of the bronchus. Bronchial Injury is an understood risk of performing a minimally invasive Pulmonary Resection procedure.
Clinical_HA | Image Viewer |  | Acute respiratory failure is the failure of pulmonary capillaries to perform air exchange, resulting in fluid buildup in the alveoli, systemic hypoxia and hypercapnia. Pulmonary edema caused by surgical manipulation can be responsible for acute respiratory failure. Acute Respiratory Failure is an understood risk of performing a minimally invasive Pulmonary Resection procedure.