## original fda lists
|  |  | 1 | 2 | 3 | 5
FDA # | Problem Category | Sub-Category | Device Problem | Code | FDA Code | Description
1 | Device Operational Issue | Device Operations Differently | device operates differently than expected | C62955 | FDA 2913 | Issue associated with any deviations from expected performance while operating and using the device.
2 | Device Operational Issue | Device Operations Differently | device displays error message | C63205 | FDA 2591 | Issue associated with a device prompting user with an error message in order to indicate a device problem
3 | Device Operational Issue | Device Operations Differently | device stops intermittently | C62924 | FDA 1599 | Issue associated with any deviation from device document specification relating to the irregular rate at which the device stops.
4 | Device Operational Issue | Device Operations Differently | failure to adhere or bond | C63032 | FDA 1031 | Issue associated with difficulties in attaching a device to another object including another device or device component or to a patient body part.
5 | Device Operational Issue | Device Operations Differently | failure to fire | C63177 | FDA 2610 | Issue associated with a therapy or algorithm not being delivered or executed at the expected time.
6 | Device Operational Issue | Device Operations Differently | failure to obtain samples | C63172 | FDA 2533
7 | Device Operational Issue | Device Operations Differently | failure to prime | C63167 | FDA 1492
8 | Device Operational Issue | Device Operations Differently | misassembled | C62958 | FDA 1398 | Issue associated with the use of the device characterized by incorrect assembly of device components, parts or constituents.
9 | Device Operational Issue | Device Operations Differently | prompts will not clear | C62861 | FDA 2281
10 | Device Operational Issue | Device Operations Differently | therapy delivered to incorrect body area | C62858 | FDA 1508 | Issue associated with the device causing unintended therapeutic action to an area of the body other than the intended area.
11 | Device Operational Issue | Device Operations Differently | unintended ejection | C62817 | FDA 1234 | Issue associated with unexpected discharge of device from expected location includes but not limited to devices such as clip appliers, film cartridge, staples
12 | Device Operational Issue | Calibration Issue | calibration issue | C63028 | FDA 2890 | Issue associated with the operation of the device, related to its accuracy, and associated with the calibration of the device
13 | Device Operational Issue | Calibration Issue | failure to calibrate | C62994 | FDA 2440 | Issue associated with the failure of the device to perform a self-calibration procedure or process designed to assure the accuracy and proper performance of the device.
14 | Device Operational Issue | Calibration Issue | calibration error | C63284 | FDA 1078
15 | Device Operational Issue | Calibration Issue | failure to recalibrate | C63241 | FDA 1517
16 | Device Operational Issue | Calibration Issue | imprecision | C63113 | FDA 1307
17 | Device Operational Issue | Calibration Issue | overcorrection | C72890 | FDA 3006 | Issue associated with an adjustment that surpasses a set of criterion.
18 | Device Operational Issue | Computer Software Issue | computer software issue | C63269 | FDA 1112 | Issue associated with written programs, codes, and/or software system that affects device performance or communication with another device.
19 | Device Operational Issue | Computer Software Issue | application network issue | C64349 | FDA 2879 | Issue associated with the deviations from documented system specifications that affects overall system performance and/or the performance of an individual device or collection of devices connected to that system.
20 | Device Operational Issue | Computer Software Issue | application program issue | C63305 | FDA 2880 | Issue associated with the requirement for software to fulfill its function within an intended use or application.
21 | Device Operational Issue | Computer Software Issue | application interface becomes nonfunctional or program exits abnormally | C63306 | FDA 1138
22 | Device Operational Issue | Computer Software Issue | application program version or upgrade problem | C63304 | FDA 2881
23 | Device Operational Issue | Computer Software Issue | incorrect error code | C63206 | FDA 2963
24 | Device Operational Issue | Computer Software Issue | problem with software installation | C67507 | FDA 3013 | Issue associated with installing the device software in a manner that allows full functioning of the device. Source of installation could be manufacturer or user.
25 | Device Operational Issue | Computer Software Issue | programming issue | C62839 | FDA 3014 | Issue associated with the written program code or application software used to satisfy a stated need or objective for functioning of the device. These do not include issues associated with the operating system
26 | Device Operational Issue | Computer Software Issue | incorrect software programming calculations | C63081 | FDA 1495
27 | Device Operational Issue | Computer Software Issue | dose calculation error due to software problem | C63220 | FDA 1189
28 | Device Operational Issue | Computer Software Issue | parameter calculation error due to software problem | C63083 | FDA 1449
29 | Device Operational Issue | Computer Software Issue | power calculation error due to software problem | C62916 | FDA 1473
30 | Device Operational Issue | Computer Software Issue | medication error | C91396 | FDA 3198 | Event in which the device software design results in errors of medication preparation or administration.
31 | Device Operational Issue | Computer Software Issue | data issue | C91397 | FDA 3196 | Event in which data (charting, orders, results) is not correctly stored, transferred, updated, or displayed.
32 | Device Operational Issue | Computer Software Issue | loss of data | C63257 | FDA 2903 | Event in which information is unintentionally permanently or temporarily lost, deleted, corrupted, or overwritten.
33 | Device Operational Issue | Computer Software Issue | patient data issue | C91398 | FDA 3197 | Event in which information is accessed by the healthcare provider and either the wrong patient or the wrong information is retrieved despite correct inquiry procedures.
34 | Device Operational Issue | Computer Software Issue | computer operating system issue | C63270 | FDA 2898 | Issue associated with the machinery operating system, a collection of software, firmware, and hardware elements that control the execution of computer programs and provides such services as computer resource allocation, job control, input/output control, and file management in a computer system.
35 | Device Operational Issue | Computer Software Issue | operating system becomes nonfunctional | C62894 | FDA 2996 | Issue associated with malfunction of the computer operating system as opposed to an application software issue.
36 | Device Operational Issue | Computer Software Issue | operating system version or upgrade problem | C62893 | FDA 2997 | Issue associated with replacing an older operation system to an up-to-date operation system.
37 | Device Operational Issue | Computer Software Issue | computer system security issue | C64348 | FDA 2899
38 | Device Operational Issue | Computer Software Issue | unauthorized access to computer system | C63259 | FDA 3025 | Issue associated with an access that was not permitted to the computer system that may lead to modification of program, corruption of data, or and break in network security. This concept is closely associated with computer integrity which is the degree to which a system or component prevents unauthorized access to, or modification of, computer programs or data.
39 | Device Operational Issue | Computer Software Issue | application security issue | C63042 | FDA 2882 | Issue associated with the acquisition of computer programming codes that can replicate and spread from one computer system to another thereby leading to damaged software, hardware and data.
40 | Device Operational Issue | Computer Software Issue | data back-up problem | C63258 | FDA 2902 | Issue associated with problems relating to a system, component, file, procedure, or person available to replace or help restore a primary item in the event of a failure or externally caused disaster.
41 | Device Operational Issue | Computer Software Issue | failure to back-up | C63195 | FDA 1047 | Issue associated with the inability to backup or to retrieve a backed up version (corrupted file) of device data or system files.
42 | Device Operational Issue | Computer Software Issue | failure to convert to back-up | C63189 | FDA 1048
43 | Device Operational Issue | Computer Software Issue | date-related software issue | C67508 | FDA 2582 | Issue associated with programming of calendar dates and/or time as a factor in the operation of a medical device.
44 | Device Operational Issue | Connection Issue | connection issue | C62952 | FDA 2900 | Issue associated with linking of device and/or device components and/or the functional units set up to provide means for a transfer of liquid, gas, electricity or data.
45 | Device Operational Issue | Connection Issue | blocked connection | C63288 | FDA 2888 | Issue associated with linking of device and/or device components whereby their functional units set up to provide means for a transfer of fluid, gas, or data is prevented or impeded.
46 | Device Operational Issue | Connection Issue | decoupling | C63256 | FDA 1145 | Issue associated with the device and/or device components being unassociated in such a way that fluid, gas, power or signal information may not be transferred from one to another.
47 | Device Operational Issue | Connection Issue | disconnection | C63223 | FDA 1171 | Issue associated with the linking of device and/or device components having a sufficient open space to prevent gas, liquid or electrical current flow between connectors.
48 | Device Operational Issue | Connection Issue | failure to disconnect | C63180 | FDA 2541 | Issue associated with the linking of device and/or device components whereby termination of the transfer of liquid, gas, electricity, or information cannot be accomplished, or linking components do not come apart, or disconnect, when expected.
49 | Device Operational Issue | Connection Issue | fitting problem | C63142 | FDA 2183 | Issue associated with the connection of device and/or device components whereby channels, switching systems, and other functional units set up to provide means for a transfer of liquid, gas, electricity, or information do not match or fit.
50 | Device Operational Issue | Connection Issue | loose or intermittent connection | C63055 | FDA 1371
51 | Device Operational Issue | Connection Issue | misconnection | C62915 | FDA 1399 | Issue associated with the connection of device and/or device components being improper or not in accordance with device specification, requirements and intended uses.
52 | Device Operational Issue | Deployment Issue | deployment issue | C63013 | FDA 2906 | Issue associated with any deviations from device documented performance specifications relating to the sequence of events for activation and positioning of the device or one of its components into a specific body location.
53 | Device Operational Issue | Deployment Issue | difficult to deploy | C63244 | FDA 1157 | Issue associated with users experiencing difficulty or uneasiness to deploy device and/or device components to a specified location.
54 | Device Operational Issue | Deployment Issue | difficult or delayed activation | C63035 | FDA 2577
55 | Device Operational Issue | Deployment Issue | failure to deploy | C63183 | FDA 1158 | Issue associated with the inability of device and/or device components to be activated and positioned in a specified location.
56 | Device Operational Issue | Deployment Issue | failure to expand | C96715 | FDA 3270 | Issue associated with the device or one of its components failing to expand.
57 | Device Operational Issue | Deployment Issue | failure to separate | C63159 | FDA 2547 | Issue associated with the device or one of its components failing to detach or separate as intended.
58 | Device Operational Issue | Deployment Issue | premature deployment | C62863 | FDA 1484 | Issue associated with an early and unexpected deployment of the device and/or device components from the system.
59 | Device Operational Issue | Deployment Issue | self-activation or keying | C62844 | FDA 1557 | Issue associated with the unintended activation of device, or device having been unexpectedly turned on during use.
60 | Device Operational Issue | Electrical Issue | electrical issue | C63007 | FDA 1198 | Issue associated with a failure of the electrical circuitry or components of the device
61 | Device Operational Issue | Electrical Issue | capturing issue | C63027 | FDA 2891 | Issue associated with the inability of device and/or device components to achieve successful depolarization and contraction of a cardiac chamber caused by a pacemaker output pulse.
62 | Device Operational Issue | Electrical Issue | failure to capture | C62993 | FDA 1081 | Issue associated with the failure to achieve effective and consistent depolarization of the heart resulting from the electrical stimulus of the pacemaker.
63 | Device Operational Issue | Electrical Issue | high capture threshold | C95875 | FDA 3266 | Issue with the amount of output energy needed to cause cardiac depolarization being higher than expected/desired.
64 | Device Operational Issue | Electrical Issue | intermittent capture | C63069 | FDA 1080 | Issue associated with the ineffective and inconsistent depolarization of the heart.
65 | Device Operational Issue | Electrical Issue | unstable capture threshold | C95876 | FDA 3269 | Issue with the amount of output energy needed to cause cardiac depolarization being unstable.
66 | Device Operational Issue | Electrical Issue | continuous firing | C63265 | FDA 1123 | Issue associated with the excessive production of electrical impulses over a period.
67 | Device Operational Issue | Electrical Issue | power source issue | C63025 | FDA 3010
68 | Device Operational Issue | Electrical Issue | battery issue | C63030 | FDA 2885
69 | Device Operational Issue | Electrical Issue | battery impedance issue | C64334 | FDA 2884
70 | Device Operational Issue | Electrical Issue | high battery impedance | C64336 | FDA 2947
71 | Device Operational Issue | Electrical Issue | low battery impedance | C64335 | FDA 2973
72 | Device Operational Issue | Electrical Issue | failure to run on portable mode | C62990 | FDA 1466
73 | Device Operational Issue | Electrical Issue | low battery | C63048 | FDA 2584
74 | Device Operational Issue | Electrical Issue | premature discharge of battery | C62864 | FDA 1057
75 | Device Operational Issue | Electrical Issue | charging issue | C63026 | FDA 2892 | Issue associated with the inability of device and/or device components to successfully charge an electrical source.
76 | Device Operational Issue | Electrical Issue | aborted charge | C63281 | FDA 2288 | Issue associated with the premature ending of the charging process (e.g. of a battery or other charge storage device).
77 | Device Operational Issue | Electrical Issue | delayed charge time | C63280 | FDA 2586 | Issue associated with an unexpected amount of time required to charge the device as required (for example, a delay in starting charging or a longer than expected charge time).
78 | Device Operational Issue | Electrical Issue | failure to charge | C63193 | FDA 1085 | Issue associated with inability to initiate the appropriate charging process (e.g. of a battery or other charge storage device)
79 | Device Operational Issue | Electrical Issue | failure to power-up | C62992 | FDA 1476
80 | Device Operational Issue | Electrical Issue | failure to run on AC/DC | C63162 | FDA 1001 | Failure to Run on AC-DC Medical Device Problem
81 | Device Operational Issue | Electrical Issue | unintended energization | C62818 | FDA 1162 | Problem associated with the device turning on when not intended.
82 | Device Operational Issue | Electrical Issue | device inoperable | C62979 | FDA 1663 | Issue associated with the device being in a nonfunctional or inoperable state and cannot be used unless the causes of the inoperability are located and fixed.
83 | Device Operational Issue | Electrical Issue | device sensing issue | C63238 | FDA 2917 | Issue associated with device features that are designed to respond to a physical stimulus (temperature, illumination, motion, cardiac rhythms) that do not transmit a resulting signal for interpretation or measurement.
84 | Device Operational Issue | Electrical Issue | failure to sense | C63160 | FDA 1559 | Issue associated with the failure of a device designed to respond to a physical stimulus (as temperature, illumination, motion) to transmit a resulting signal for interpretation or measurement.
85 | Device Operational Issue | Electrical Issue | invalid sensing | C63061 | FDA 2293 | Issue associated with the inaccuracy of a device designed to respond to an input (such as temperature, illumination, motion) to transmit a resulting signal that can not be used for interpretation or measurement.
86 | Device Operational Issue | Electrical Issue | decreased sensitivity | C63253 | FDA 2534 | Issue with the device being less sensitive to an input than intended or expected.
87 | Device Operational Issue | Electrical Issue | failure to analyze rhythm | C63196 | FDA 1539
88 | Device Operational Issue | Electrical Issue | failure to select signal | C63161 | FDA 1582 | Problem associated with the failure of the device to select the appropriate input signal.
89 | Device Operational Issue | Electrical Issue | high sensing threshold | C63121 | FDA 2574 | Issue associated with the amount of input required by the device to detect a signal being higher than expected/desired.
90 | Device Operational Issue | Electrical Issue | incorrect interpretation of rhythm | C63086 | FDA 1543
91 | Device Operational Issue | Electrical Issue | increased sensitivity | C63078 | FDA 2535 | Issue with the device being more sensitive to an input than intended or expected.
92 | Device Operational Issue | Electrical Issue | loss of threshold | C63050 | FDA 1633 | Issue associated with the loss of the minimum amount of energy, voltage, or current needed to consistently stimulate the heart muscle.
93 | Device Operational Issue | Electrical Issue | low sensing threshold | C63044 | FDA 2575 | Issue associated with the amount of an input required by the device to detect a signal being lower than expected/desired.
94 | Device Operational Issue | Electrical Issue | over-sensing | C62881 | FDA 1438
95 | Device Operational Issue | Electrical Issue | sensing intermittently | C63065 | FDA 1558
96 | Device Operational Issue | Electrical Issue | under-sensing | C62822 | FDA 1661 | Problem related to failure of the device to properly detect intrinsic cardiac activity and respond appropriately.
97 | Device Operational Issue | Electrical Issue | failure to conduct | C63191 | FDA 1114 | Issue associated with the inability of a device and/or device components to allow a current of electricity to pass or to conduct electricity continuously along an electrical path.
98 | Device Operational Issue | Electrical Issue | failure to convert rhythm | C63190 | FDA 1540 | Failure of a device therapy or set of therapies to terminate the arrhythmia that the therapy is meant to terminate.
99 | Device Operational Issue | Electrical Issue | failure to interrogate | C63173 | FDA 1332 | Issue associated with the device failure to appropriately respond to signals from a system designed to interrogate its status.
100 | Device Operational Issue | Electrical Issue | difficult to interrogate | C63232 | FDA 1331 | Issue associated with difficulty of a transponder system to trigger a response.
101 | Device Operational Issue | Electrical Issue | failure to pace or properly pace | C63169 | FDA 1439 | Issue associated with the inability of device and/or device components to generate a therapeutic simulated heart beat via electrical impulses once arrhythmia is detected.
102 | Device Operational Issue | Electrical Issue | failure to convert rhythm | C63190 | FDA 1540
103 | Device Operational Issue | Electrical Issue | failure to spike | C63157 | FDA 2536 | Failure of device to generate a correctly shaped pacing output, for example, a waveform that is too wide.
104 | Device Operational Issue | Electrical Issue | inaccurate synchronization | C63101 | FDA 1609 | Issue associated with an error due to imperfect timing of two operations this may or may not include signal transmission time.
105 | Device Operational Issue | Electrical Issue | no pacing | C95877 | FDA 3268 | Issue associated with the device ceasing to deliver paces.
106 | Device Operational Issue | Electrical Issue | pacer found in back-up mode | C62880 | FDA 1440
107 | Device Operational Issue | Electrical Issue | pacing asynchronously | C63298 | FDA 1441 | Issue associated with a pacing transmission process such that between any two significant instants in the same group, there is always an integral number of unit intervals. Between two significant instants located in different groups, there are not always an integral number of unit intervals.
108 | Device Operational Issue | Electrical Issue | pacing inadequately | C63097 | FDA 1442 | Pacing voltage or pulse width is less than desired.
109 | Device Operational Issue | Electrical Issue | pocket stimulation | C62872 | FDA 1463 | Issue associated with a pocket of skin in which the pulse generator is housed.
110 | Device Operational Issue | Electrical Issue | pacing intermittently | C63066 | FDA 1443 | Issue associated with the failure of pacing device for a limited period of time, following which the item recovers its ability to perform its required function without being subjected to any external corrective action. Note: such as failure is often recurrent.
111 | Device Operational Issue | Electrical Issue | failure to shock or properly shock | C63158 | FDA 1573 | Issue associated with the inability of device and/or device components to provide an appropriate or successful electrical shock
112 | Device Operational Issue | Electrical Issue | failure to deliver counter-shock | C63185 | FDA 1133 | Issue associated with the effective electromotive force within a system that fails to oppose the passage of current in a specified direction.
113 | Device Operational Issue | Electrical Issue | failure to discharge | C63181 | FDA 1169 | Issue associated with the failure of a battery or other charge storage device to appropriately discharge as intended.
114 | Device Operational Issue | Electrical Issue | inappropriate shock | C63093 | FDA 1574 | Issue associated with the inappropriate delivery of an electrical energy
115 | Device Operational Issue | Electrical Issue | intermittent shock | C63064 | FDA 2287 | Issue associated with the failure to deliver shock for a limited period of time, following which the item recovers its ability to perform its required function without being subjected to any external corrective action. Note: such as failure is often recurrent.
116 | Device Operational Issue | Electrical Issue | grounding malfunction | C63125 | FDA 1271 | Issue associated with the inability to connect conductors of an electronic system for the purpose of controlling or impeding ground currents and voltages.
117 | Device Operational Issue | Incompatibility Problem | incompatibility problem | C62983 | FDA 2960 | Issue associated with the device not being compatible with another device component or substance (medication, body fluid etc) that it contains or transports. Does not refer to device issues related to disconnection.
118 | Device Operational Issue | Incompatibility Problem | component or accessory incompatibility | C63020 | FDA 2897 | Issue associated with the incompatibility of any device and/or device components while being operated in the same use environment thereby leading to a dysfunction between the device and its components.
119 | Device Operational Issue | Incompatibility Problem | accessory incompatible | C63036 | FDA 1004 | An accessory required for the intended purpose of the device appears incompatible with device, thus compromising the intended function of the device.
120 | Device Operational Issue | Incompatibility Problem | component incompatible | C63022 | FDA 1108
121 | Device Operational Issue | Incompatibility Problem | device-device incompatibility | C63236 | FDA 2919 | Issue associated with the incompatibility of two or more devices while being operated in the same use environment thereby leading to a dysfunction of more than one device.
122 | Device Operational Issue | Incompatibility Problem | patient-device incompatibility | C62919 | FDA 2682 | Issue associated with the interaction between the patient's physiology or anatomy and the device that affects patient and/or device. (e.g. biocompatibility Issues, patient's weight exceeds safe working load of a lifting device.)
123 | Device Operational Issue | Incompatibility Problem | rejection | C62853 | FDA 1524
124 | Device Operational Issue | Incompatibility Problem | isize ncorrect for patient | C62947 | FDA 1583
125 | Device Operational Issue | Incompatibility Problem | measurement system incompatibility | C63040 | FDA 2982 | Issue associated with the incompatibility of the measurement systems between and/or within device systems that are inherent to the individual device thereby leading to miscalculated or mismatched measurements from those devices, e.g., international metric system versus U.S. measurement system.
126 | Device Operational Issue | Infusion or Flow Issue | infusion or flow issue | C63075 | FDA 2964 | Issue associated with the device failing to deliver liquids or gases as intended (e.g. delivering drugs at incorrect rate, issues with drawing fluid from a system etc)
127 | Device Operational Issue | Infusion or Flow Issue | deflation issue | C63249 | FDA 1149 | Issue associated with the inability of device and/or device components to release its contents
128 | Device Operational Issue | Infusion or Flow Issue | excess flow or over-infusion | C62882 | FDA 1311 | Issue associated with an overdose of delivery therapy, e.g., air, gas, drugs or fluids being delivered into a device or a patient under positive pressure that is being generated by a pump.
129 | Device Operational Issue | Infusion or Flow Issue | filling problem | C63144 | FDA 1233 | Issue associated with the method or amount of time associated with the delivery of a fluid. Time to delivery or amount of delivered entity may be affected.
130 | Device Operational Issue | Infusion or Flow Issue | inability to auto-fill | C63106 | FDA 1044
131 | Device Operational Issue | Infusion or Flow Issue | overfill | C62884 | FDA 2404
132 | Device Operational Issue | Infusion or Flow Issue | short fill | C62841 | FDA 1575
133 | Device Operational Issue | Infusion or Flow Issue | volume accuracy issue | C62805 | FDA 1675
134 | Device Operational Issue | Infusion or Flow Issue | filtration issue | C62986 | FDA 2941 | Issue associated with the process of passing a substance through a porous medium, e.g., a blood clot filter for the removal of suspended matter.
135 | Device Operational Issue | Infusion or Flow Issue | inadequate filtration process | C62985 | FDA 2308
136 | Device Operational Issue | Infusion or Flow Issue | ultra-filtration | C62827 | FDA 1656 | Issue associated with the transfer of fluid between the blood and dialysate through the dialysis membrane due to a pressure gradient (trans-membrane pressure) existing between the blood and dialysate compartments.
137 | Device Operational Issue | Infusion or Flow Issue | improper flow or infusion | C63110 | FDA 2954 | Issue associated with the unsubstantiated regulation and delivery of therapy, e.g., air, gas, drugs or fluids into a device or a patient under positive pressure that is being generated by a pump.
138 | Device Operational Issue | Infusion or Flow Issue | bleed back | C63290 | FDA 1064
139 | Device Operational Issue | Infusion or Flow Issue | free or unrestricted flow | C63130 | FDA 2945 | Issue associated with uncontrolled flow of infusion of air, gas or fluids.
140 | Device Operational Issue | Infusion or Flow Issue | gradient increase | C63126 | FDA 1270 | Issue associated with the increased rate of change in temperature, pressure, or other variables as a function of distance, time, etc.
141 | Device Operational Issue | Infusion or Flow Issue | inaccurate delivery | C63104 | FDA 2339
142 | Device Operational Issue | Infusion or Flow Issue | inaccurate dispensing | C63103 | FDA 2955
143 | Device Operational Issue | Infusion or Flow Issue | inaccurate flowrate | C63102 | FDA 1249
144 | Device Operational Issue | Infusion or Flow Issue | intermittent infusion | C63067 | FDA 2341
145 | Device Operational Issue | Infusion or Flow Issue | reflux within device | C62854 | FDA 1522 | Issue associated with a backward or return flow.
146 | Device Operational Issue | Infusion or Flow Issue | restricted flowrate | C62849 | FDA 1248
147 | Device Operational Issue | Infusion or Flow Issue | tidal volume fluctuations | C63138 | FDA 1634 | Issue associated with the amount of gas that is inspired and expired during one respiratory cycle.
148 | Device Operational Issue | Infusion or Flow Issue | inflation issue | C63076 | FDA 1310 | Issue associated with the inability of a device and/or device components to expand or enlarge with the intended inflation agent (e.g. saline or air)
149 | Device Operational Issue | Infusion or Flow Issue | insufficient flow or underinfusion | C62823 | FDA 2182 | Issue associated with an under dose of intravenous therapy, i.e., drugs or fluids being delivered into a patient under positive pressure generated by a pump.
150 | Device Operational Issue | Infusion or Flow Issue | no flow | C62902 | FDA 2991 | Issue arising from the device failing to deliver the specified liquid or gas.
151 | Device Operational Issue | Infusion or Flow Issue | failure to deliver | C63186 | FDA 2338
152 | Device Operational Issue | Infusion or Flow Issue | failure to infuse | C63174 | FDA 2340
153 | Device Operational Issue | Infusion or Flow Issue | inability to irrigate | C63105 | FDA 1337
154 | Device Operational Issue | Infusion or Flow Issue | obstruction within device | C62897 | FDA 2423
155 | Device Operational Issue | Infusion or Flow Issue | occlusion within device | C62896 | FDA 1423 | Issue associated with an obstruction or blockage within any of the device components (e.g. tube, opening, pipe) that results in a restriction of flow.
156 | Device Operational Issue | Infusion or Flow Issue | blockage within device or device component | C63289 | FDA 1065
157 | Device Operational Issue | Infusion or Flow Issue | device clogged | C63024 | FDA 1094 | Problem related to an obstruction or blockage within the device component (e.g. tube, opening, pipe) that results in no flow.
158 | Device Operational Issue | Infusion or Flow Issue | difficult to flush | C63317 | FDA 1251
159 | Device Operational Issue | Infusion or Flow Issue | pressure issue | C62934 | FDA 3012 | Issue associated with the application of a force either internal of external to the device that compromises the flow of fluid or gas.
160 | Device Operational Issue | Infusion or Flow Issue | decrease in pressure | C63255 | FDA 1490
161 | Device Operational Issue | Infusion or Flow Issue | increase in pressure | C63080 | FDA 1491
162 | Device Operational Issue | Infusion or Flow Issue | no pressure | C64243 | FDA 2994
163 | Device Operational Issue | Infusion or Flow Issue | pumping issue | C62860 | FDA 3016 | Issue associated with pump performance deviating from specifications in a way to compromise flow or infusion
164 | Device Operational Issue | Infusion or Flow Issue | decreased pump speed | C63252 | FDA 1500
165 | Device Operational Issue | Infusion or Flow Issue | failure to pump | C63165 | FDA 1502
166 | Device Operational Issue | Infusion or Flow Issue | increased pump speed | C63077 | FDA 1501
167 | Device Operational Issue | Infusion or Flow Issue | pre- or post-pumping problem | C62867 | FDA 1477
168 | Device Operational Issue | Infusion or Flow Issue | pumping stopped | C62859 | FDA 1503
169 | Device Operational Issue | Infusion or Flow Issue | suction issue | C62833 | FDA 2170 | Issue associated with suction equipment, which may be a manual, electrical, vacuum or pressure source operated to evacuate and remove undesired substances (air, gas, fluid, or particulates via tubing and collection bag.
170 | Device Operational Issue | Infusion or Flow Issue | aspiration issue | C63299 | FDA 2883
171 | Device Operational Issue | Infusion or Flow Issue | decrease in suction | C63254 | FDA 1146 | Issue associated with the removal by suction of excess fluid or gas from a body cavity.
172 | Device Operational Issue | Infusion or Flow Issue | increase in suction | C63079 | FDA 1604
173 | Device Operational Issue | Output Issue | output issue | C62941 | FDA 3005 | Issue associated with any deviation from device documented performance specifications relating to the end result, data, or test results provided by the device.
174 | Device Operational Issue | Output Issue | improper device output | C63108 | FDA 2953 | Issue associated with the measurements/values/data obtained from the device being different from displayed, printed, stored, or exported measurements/values/data.
175 | Device Operational Issue | Output Issue | false device output | C63150 | FDA 1226
176 | Device Operational Issue | Output Issue | image orientation incorrect | C62887 | FDA 1305 | Issue associated with an incorrect image orientation on the device display.
177 | Device Operational Issue | Output Issue | image reversal | C63115 | FDA 1358
178 | Device Operational Issue | Output Issue | improper gas output | C63109 | FDA 1266
179 | Device Operational Issue | Output Issue | inappropriate prompts | C63094 | FDA 2280
180 | Device Operational Issue | Output Issue | inaudible voice prompts | C63092 | FDA 2283
181 | Device Operational Issue | Output Issue | no voice prompts | C62899 | FDA 2282
182 | Device Operational Issue | Output Issue | incorrect display | C63088 | FDA 1184 | Issue associated with device failure and malfunction when device does not display adequate results or have proper display.
183 | Device Operational Issue | Output Issue | erratic display | C63207 | FDA 1182
184 | Device Operational Issue | Output Issue | image display error | C63117 | FDA 1304 | Event in which measurement functions produce erroneous results or the image display is corrupted.
185 | Device Operational Issue | Output Issue | issue with displayed error message | C76129 | FDA 2967 | Issue associated with error messages that are undecipherable or which fail to prompt the user toward appropriate corrective action.
186 | Device Operational Issue | Output Issue | unable to obtain readings | C62826 | FDA 1516
187 | Device Operational Issue | Output Issue | output above specifications | C62943 | FDA 1432
188 | Device Operational Issue | Output Issue | radiation overexposure | C72892 | FDA 3017 | Issue associated with excessive radiation emitted from radiological or diagnostic devices.
189 | Device Operational Issue | Output Issue | output below specifications | C62942 | FDA 3004
190 | Device Operational Issue | Output Issue | radiation underexposure | C72893 | FDA 3018 | Issue associated with too little radiation emitted from radiological or diagnostic devices. Results in insufficient production of normal image contrast.
191 | Device Operational Issue | Output Issue | output energy incorrect | C63002 | FDA 1431
192 | Device Operational Issue | Output Issue | energy output to patient tissue incorrect | C63001 | FDA 1209
193 | Device Operational Issue | Output Issue | energy spectrum incorrect | C63000 | FDA 1210
194 | Device Operational Issue | Output Issue | failure to deliver energy | C63184 | FDA 1211
195 | Device Operational Issue | Output Issue | poor quality image | C62869 | FDA 1408 | Issue associated with an image or any visual representation displayed by the device with non-distinctive characteristics.
196 | Device Operational Issue | Output Issue | image resolution poor | C62871 | FDA 1306
197 | Device Operational Issue | Output Issue | therapeutic or diagnostic output failure | C62828 | FDA 3023 | Issue associated with the failure of the device to deliver any end results relating to treatment or identification of a disease.
198 | Device Operational Issue | Output Issue | incorrect or inadequate result | C62848 | FDA 1535 | Issue associated with a nonconforming end result, data, or test results provided by the device to its performance specifications.
199 | Device Operational Issue | Output Issue | display misread | C63221 | FDA 1181 | Issue associated with not correctly reading the patient or test result information provided by the device.
200 | Device Operational Issue | Output Issue | false negative result | C63149 | FDA 1225 | Issue associated with the device incorrectly reporting that something has not been detected and misleads the operator into not taking certain actions when action should be taken.
201 | Device Operational Issue | Output Issue | false positive result | C63147 | FDA 1227 | Issue associated with the device incorrectly reporting that something has been detected and misleads the operator to take certain actions.
202 | Device Operational Issue | Output Issue | false reading from device noncompliance | C63145 | FDA 1228 | Issue associated with the noncompliance of the device to meet its specific performance characteristics thereby causing a false reading.
203 | Device Operational Issue | Output Issue | high readings | C63123 | FDA 2459 | Reading provided by the device is too high or higher than expected.
204 | Device Operational Issue | Output Issue | incorrect measurement | C63085 | FDA 1383
205 | Device Operational Issue | Output Issue | incorrect or inadequate test results | C62829 | FDA 2456
206 | Device Operational Issue | Output Issue | artifact | C63300 | FDA 1036 | Issue associated with impurities or interference in a signal or image (e.g. ECG artifact).
207 | Device Operational Issue | Output Issue | high test results | C63122 | FDA 2457
208 | Device Operational Issue | Output Issue | low test results | C63045 | FDA 2458
209 | Device Operational Issue | Output Issue | missing test results | C95878 | FDA 3267 | Issue associated with the results of a test or measurement not appearing.
210 | Device Operational Issue | Output Issue | low readings | C63046 | FDA 2460
211 | Device Operational Issue | Output Issue | unexpected therapeutic results | C62821 | FDA 1631 | Issue associated with the use of a medical device for therapeutic purposes.
212 | Device Operational Issue | Output Issue | no device output | C62900 | FDA 1435 | Issue associated with no measurement outcome, value, or data obtained from the device.
213 | Device Operational Issue | Output Issue | no display or display failure | C62904 | FDA 1183 | Issue associated with the absence of visual indicator.
214 | Device Operational Issue | Protective Measure Issue | protective measure issue | C62932 | FDA 3015 | Issue associated with any deviations from device documented performance specifications relating to the implemented and inherited design features specific to devices used for reducing risks to patient or caregiver or maintaining risks within specified levels.
215 | Device Operational Issue | Protective Measure Issue | device alarm system issue | C63033 | FDA 1012 | Issue associated with electrical equipment or a system that detects alarm conditions (a state of the alarm system when it has determined that a potential or actual hazard exists) and, as appropriate, generates alarm signals.
216 | Device Operational Issue | Protective Measure Issue | false alarm | C63152 | FDA 1013 | Issue associated with device providing incorrect alarm warning or alert to user.
217 | Device Operational Issue | Protective Measure Issue | improper alarm | C63112 | FDA 2951
218 | Device Operational Issue | Protective Measure Issue | alarm not visible | C63309 | FDA 1022
219 | Device Operational Issue | Protective Measure Issue | defective alarm | C63251 | FDA 1014
220 | Device Operational Issue | Protective Measure Issue | delayed alarm | C63247 | FDA 1011
221 | Device Operational Issue | Protective Measure Issue | low audible alarm | C63049 | FDA 1016
222 | Device Operational Issue | Protective Measure Issue | not audible alarm | C63310 | FDA 1019
223 | Device Operational Issue | Protective Measure Issue | fail-safe mechanism issue | C62997 | FDA 2936 | Problem associated with the feature that prevents the unsafe use of the device.
224 | Device Operational Issue | Protective Measure Issue | fail-safe design failure | C63199 | FDA 1222 | Issue associated with safety features malfunctioning in such a way that results in harm to the patient or device operator.
225 | Device Operational Issue | Protective Measure Issue | failure of device to self-test | C62989 | FDA 2937 | Issue associated with the device failing to perform an internal self-diagnostic process to ensure normal operation during or prior to use.
226 | Device Operational Issue | Protective Measure Issue | no fail-safe mechanism | C62903 | FDA 2990
227 | Device Operational Issue | Protective Measure Issue | failure to auto stop | C62995 | FDA 2938 | Issue associated with the inability of a device to turn itself off when the device is not in an operable condition.
228 | Device Operational Issue | Protective Measure Issue | plunge | C62873 | FDA 1462
229 | Device Operational Issue | Protective Measure Issue | premature indicator activation | C62865 | FDA 3011
230 | Device Operational Issue | Protective Measure Issue | premature elective replacement indicator | C63315 | FDA 1483
231 | Device Operational Issue | Protective Measure Issue | premature end-of-life indicator | C63320 | FDA 1480
232 | Device Operational Issue | Protective Measure Issue | reset issue | C62852 | FDA 3019 | Issue associated with setting a variable, register, or other storage location back to a prescribed state.
233 | Device Operational Issue | Protective Measure Issue | failure to reset | C62991 | FDA 1532
234 | Device Operational Issue | Protective Measure Issue | failure to zero | C62987 | FDA 1683
235 | Device Operational Issue | Protective Measure Issue | inappropriate or unexpected reset | C63031 | FDA 2959
236 | Device Operational Issue | Protective Measure Issue | shielding failure | C62842 | FDA 1568 | Issue associated with device inability to act as a barrier for absorption of radiation energy in X-rays, gamma rays, etc.
237 | Facilities Issue |  | facilities issue - LEVEL 1 - |  |  | Issue associated with the physical establishment and environment in which medical devices are transported, stored, processed, serviced or used within the user facility
238 | Facilities Issue | Device Maintenance Issue | device maintenance issue | C62977 | FDA 1379 | Issue associated with the servicing of a device.
239 | Facilities Issue | Device Maintenance Issue | failure to service | C62928 | FDA 1563 | Issue associated with the lack of periodic preventative maintenance or performance assurance checks.
240 | Facilities Issue | Device Maintenance Issue | inadequate service | C63098 | FDA 1564 | Issue associated with inadequate periodic preventative maintenance or performance assurance checks.
241 | Facilities Issue | Device Maintenance Issue | lack of maintenance documentation or guidelines | C72671 | FDA 2971 | Issue associated with user facility not receiving adequate service documentation, guidelines, or recommendations to perform preventative and corrective maintenance and performance assurance checks.
242 | Facilities Issue | Device Maintenance Issue | maintenance does not comply to manufacturers recommendations | C62978 | FDA 2974 | Issue associated with the noncompliance or nonperformance to manufacturer specifications relating to device routine maintenance, i.e., periodic inspection, failure detection, repair, and care of the device to sustain or restore acceptable operating conditions.
243 | Facilities Issue | Disinfection or Sterilization Issue at User Location | disinfection or sterilization issue at user location | C63276 | FDA 2895 | Issue associated with the undesired introduction of impurities to a device, or the insufficient removal of any visible soil, foreign material or organism deposits on the external surfaces, crevices, and joints of a device by a mechanical and/or manual process intended to render the device sterile, safe for handling, and/or for further processes to decontaminate. This issue is restricted to happening at the location where the device is used as opposed to during manufacturer or shipping.
244 | Facilities Issue | Disinfection or Sterilization Issue at User Location | contamination during use | C63018 | FDA 1120 | Issue associated with the undesired introduction of impurities either chemical or microbiological in nature, or of foreign matter into or onto the device at the user facility.
245 | Facilities Issue | Disinfection or Sterilization Issue at User Location | contamination of device ingredient or reagent | C63266 | FDA 2901
246 | Facilities Issue | Disinfection or Sterilization Issue at User Location | device contamination with biological material | C63017 | FDA 2908 | Medical Device Contamination with Biological Material
247 | Facilities Issue | Disinfection or Sterilization Issue at User Location | bacterial contamination of device | C63297 | FDA 2303
248 | Facilities Issue | Disinfection or Sterilization Issue at User Location | biofilm coating in device | C63293 | FDA 1062 | Problem associated with the undesired introduction of a biofilm coating into or onto the device.
249 | Facilities Issue | Disinfection or Sterilization Issue at User Location | device contamination with blood or blood product | C63016 | FDA 2317 | Problem associated with the undesired presence of body fluid in/on the device, which are not part of the documented device specifications and requirements.
250 | Facilities Issue | Disinfection or Sterilization Issue at User Location | foreign material present in device | C63133 | FDA 2944 | Issue associated with the presence of materials, which are not part of the documented device specifications and requirements.
251 | Facilities Issue | Disinfection or Sterilization Issue at User Location | device cleaning issue | C63099 | FDA 1091 | Issue associated with the insufficient removal of unwanted visible soil, foreign material or organism deposits on the external surfaces, crevices, joints of a device by a mechanical and/or manual process intended to render the device sterile, safe for handling, and/or for further processes to decontaminate.
252 | Facilities Issue | Disinfection or Sterilization Issue at User Location | device disinfection or sterilization issue | C63008 | FDA 2909
253 | Facilities Issue | Disinfection or Sterilization Issue at User Location | failure to disinfect | C63179 | FDA 1175
254 | Facilities Issue | Disinfection or Sterilization Issue at User Location | failure to sterilize | C63156 | FDA 1596
255 | Facilities Issue | Disinfection or Sterilization Issue at User Location | device rinsing issue | C63107 | FDA 2309
256 | Facilities Issue | Disinfection or Sterilization Issue at User Location | failure to flush | C63176 | FDA 1252
257 | Facilities Issue | Disinfection or Sterilization Issue at User Location | failure to remove enzymatic cleaner | C63163 | FDA 1213
258 | Facilities Issue | Disinfection or Sterilization Issue at User Location | residue after decontamination | C62850 | FDA 2325 | Issue associated with the decontamination process not adequately removing unwanted visible soil, foreign material, or organism deposits.
259 | Facilities Issue | Environmental Control or Utility Issue | environmental control or utility issue | C63209 | FDA 2929 | Issue associated with the surrounding conditions in which the device is being used such as temperature, noise, lighting, ventilation, or other external factors such as power supply.
260 | Facilities Issue | Environmental Control or Utility Issue | ambient noise issue | C63308 | FDA 2877 | Issue associated with any undesired acoustic energy or vibration that tends to interfere with the operation of the device.
261 | Facilities Issue | Environmental Control or Utility Issue | ambient temperature issue | C63307 | FDA 2878 | Issue associated with compromised device performance at the ambient room temperature
262 | Facilities Issue | Environmental Control or Utility Issue | changes in ambient temperature in device environment | C63282 | FDA 1027
263 | Facilities Issue | Environmental Control or Utility Issue | biological environmental factor | C63292 | FDA 2887
264 | Facilities Issue | Environmental Control or Utility Issue | fungus in device environment | C63128 | FDA 2316 | Issue associated with the visibility of molds, mildews, yeasts, and/or mushrooms in the immediate environment in which a device is being used.
265 | Facilities Issue | Environmental Control or Utility Issue | device expiration issue | C62998 | FDA 1216 | Issue associated with the expiration date or shelf life of a product.
266 | Facilities Issue | Environmental Control or Utility Issue | expiration date error | C62999 | FDA 2528 | Issue associated with errors in identification of expiration date.
267 | Facilities Issue | Environmental Control or Utility Issue | shelf life exceeded | C62926 | FDA 1567 | Issue associated with the using a device after the marked or known shelf life.
268 | Facilities Issue | Environmental Control or Utility Issue | device unsafe to use in environment | C73118 | FDA 2918
269 | Facilities Issue | Environmental Control or Utility Issue | electrical power problem | C63217 | FDA 2925 | Issue associated with the quality of the facility-supplied power.
270 | Facilities Issue | Environmental Control or Utility Issue | emergency power failure | C63212 | FDA 1205 | Issue associated with the failure of the facility's emergency power backup system(s) including generators and/or interruptible power systems (UPS).
271 | Facilities Issue | Environmental Control or Utility Issue | loss of power | C63052 | FDA 1475 | Issue associated with the failure of primary power supplied by the facility.
272 | Facilities Issue | Environmental Control or Utility Issue | power conditioning issue | C62868 | FDA 1474 | Issue associated with a momentary overpower/over voltage from the utility and electrical systems of user facilities
273 | Facilities Issue | Environmental Control or Utility Issue | environmental particulates | C72674 | FDA 2930 | Issue associated with fine solids or liquid particles such as dust, smoke, fume, and/or mist suspended in the immediate atmosphere in which the device is being used.
274 | Facilities Issue | Environmental Control or Utility Issue | inadequate lighting | C62870 | FDA 2957 | Issue associated with an inappropriate level of light or illumination in the immediate environment in which the device is being used or stored.
275 | Facilities Issue | Environmental Control or Utility Issue | medical gas supply problem | C72887 | FDA 2985 | Issue associated with the facility-supplied medical gases such as medical air, oxygen, nitrous oxide, and nitrogen.
276 | Facilities Issue | Environmental Control or Utility Issue | moisture or humidity problem | C62909 | FDA 2986 | Issue associated with an unsatisfactory humidity level in the storage or use environment which affects device performance.
277 | Facilities Issue | Environmental Control or Utility Issue | moisture damage | C62910 | FDA 1405 | Issue associated with damage inflicted upon the device or device components from water vapor or water in the immediate environment in which the device is being used.
278 | Facilities Issue | Environmental Control or Utility Issue | ventilation issue in device environment | C62807 | FDA 3027
279 | Facilities Issue | Environmental Control or Utility Issue | fogging | C63135 | FDA 1253 | Issue associated with the visibility of water vapor in the immediate atmosphere in which the device is being used.
280 | Facilities Issue | Environmental Control or Utility Issue | fumes or vapors | C63129 | FDA 2529 | Issue associated with the visibility, odor, or toxicity of an ambient vapor or gas which affects the operation of the device.
281 | Facilities Issue | Inadequate Storage | inadequate storage | C63095 | FDA 1600 | Issue associated with inadequate or inappropriate storage of the device, not otherwise described.
282 | Facilities Issue | Installation-Related Problem | installation-related problem | C63074 | FDA 2965 | Issue associated with unsatisfactory installation, configuration, and/or setup of a specific device or technology.
283 | Human Factors Issue |  | human factors issue - LEVEL 1 - |  |  | Issue associated with the interaction and interfacing between devices and users in terms of users' abilities, expectations, and limitations with work environments and system design.
284 | Human Factors Issue | Human-Device Interface Issue | human-device interface issue | C76121 | FDA 2949
285 | Human Factors Issue | Human-Device Interface Issue | device difficult to maintain | C79146 | FDA 3134
286 | Human Factors Issue | Human-Device Interface Issue | device difficult to setup or prepare | C63010 | FDA 1487 | Issue associated with the use of the device in terms of user experiencing difficulty in preparing the device for use, even if the operation is being performed according to labeled instructions for use.
287 | Human Factors Issue | Human-Device Interface Issue | difficult to program or calibrate | C64339 | FDA 1496 | Issue associated with the user's ability to cause device settings or actions to change to the state desired by the user.
288 | Human Factors Issue | Human-Device Interface Issue | inadequate user interface | C63100 | FDA 2958 | Issue associated with the means by which the operator and the equipment communicate or interact.
289 | Human Factors Issue | Inadequate Training | inadequate training | C73117 | FDA 1643 | Issue associated with facility not providing satisfactory initial and/or periodic user training covering operation of the device
290 | Human Factors Issue | Instruction for Use Issue | instruction for use issue | C63072 | FDA 1318 | Issue associated with the accuracy and appropriateness of any written, printed, graphic or audio/visual matter that is supplied with a medical device or its containers, wrappers with any matter that accompanies a medical device including instructions related to identification, technical description and use of the medical device provided by the device manufacturer.
291 | Human Factors Issue | Instruction for Use Issue | inadequate instructions for healthcare professional | C63087 | FDA 1319 | Issue associated with inaccuracies in any written, printed, or graphic matter that is affixed to a medical device or its containers, wrappers with any matter that accompanies a medical device including verbal instructions related to identification, technical description and use of the medical device provided by the device manufacturers.
292 | Human Factors Issue | Instruction for Use Issue | nadequate instructions for non-healthcare professional | C63267 | FDA 2956 | Issue associated with users being unclear and not able to follow any written, printed, or graphic matter that is affixed to a medical device or its containers, wrappers with any matter that accompanies a medical device including verbal instructions related to identification, technical description and use of the medical device provided by the device manufactures that vary from the standard of medical care in a given environment.
293 | Human Factors Issue | Use of Device Issue | use of device issue | C63318 | FDA 1670 | Issue associated with the user's failure to process, service, or operate the device according to the manufacturer's recommendations or recognized best practices.
294 | Human Factors Issue | Use of Device Issue | device handling issue | C95879 | FDA 3265 | User handling not in accordance with specification.
295 | Human Factors Issue | Use of Device Issue | improper or incorrect procedure or method | C62862 | FDA 2017 | Issue associated with the use of the device in terms of nonconforming to that device's intended use, specifications, procedure and process or service instructions and information provided by the device manufacturers.
296 | Human Factors Issue | Use of Device Issue | misassembled by users | C79147 | FDA 3133
297 | Human Factors Issue | Use of Device Issue | use of incorrect control settings | C63264 | FDA 1126 | Issue associated with the use of the device in terms of inappropriate and false control setting for the device's specified operation and/or intended use.
298 | Human Factors Issue | Use of Device Issue | user used incorrect product for intended use | C63082 | FDA 1494
299 | Physical Property Issue |  | physical property issue - LEVEL 1 - |  |  | Issue associated with any deviations from specifications relating to device physical properties in terms of material integrity, mechanical, manufacturing/shipping, communication/transmission, optical, temperature, chemical, and electronic issues.
300 | Physical Property Issue | Chemical Issue | chemical issue | C62945 | FDA 2893 | Issue associated with any deviations from device documented performance specifications relating to any chemical characterization, i.e., element, compound, or mixture.
301 | Physical Property Issue | Chemical Issue | chemical spillage | C63278 | FDA 2894 | Issue associated with unintentional pouring out or releasing of chemical substances.
302 | Physical Property Issue | Chemical Issue | device emits odor | C63004 | FDA 1425
303 | Physical Property Issue | Chemical Issue | device ingredient or reagent issue | C62982 | FDA 2910
304 | Physical Property Issue | Chemical Issue | biocompatibility issue | C63294 | FDA 2886
305 | Physical Property Issue | Chemical Issue | high pH | C62981 | FDA 2426
306 | Physical Property Issue | Chemical Issue | sediment, precipitate or deposit in device or device ingredient | C62845 | FDA 3021
307 | Physical Property Issue | Chemical Issue | clumping in device or device ingredient | C63275 | FDA 1095 | Issue associated with the aggregation of particles into irregular masses.
308 | Physical Property Issue | Chemical Issue | coagulation in device or device ingredient | C63274 | FDA 1096 | Issue associated with the undesired characterization of congealing, solidifying, thickening, curdling.
309 | Physical Property Issue | Chemical Issue | cross reactivity | C63261 | FDA 1137 | Issue associated with the degree to which an antibody or antigen participates in cross reactions.
310 | Physical Property Issue | Chemical Issue | particulates | C62879 | FDA 1451 | Substances that consist of separate particles that are introduced by the device during use.
311 | Physical Property Issue | Chemical Issue | precipitate in device or device ingredient | C62866 | FDA 1478 | Issue associated with the separation of solid particles from a liquid as the result of a chemical or physical change.
312 | Physical Property Issue | Chemical Issue | flashpoint threshold met | C63139 | FDA 2943 | Issue associated with approaching certain temperatures at which a combustible liquid ignites.
313 | Physical Property Issue | Chemical Issue | improper chemical reaction | C63111 | FDA 2952 | Issue associated with an unexpected or inappropriate chemical reaction or effect.
314 | Physical Property Issue | Communication or Transmission Issue | communication or transmission issue | C63271 | FDA 2896 | Issue associated with the device sending or receiving signals or data. This includes transmission among internal components of the device and other external devices to which the device is intended to communicate.
315 | Physical Property Issue | Communication or Transmission Issue | failure to read input signal | C63164 | FDA 1581
316 | Physical Property Issue | Communication or Transmission Issue | failure to transmit record | C63155 | FDA 1521
317 | Physical Property Issue | Communication or Transmission Issue | radio signal problem | C62856 | FDA 1511 | Issue associated with the wireless transmission signal in the lowest range of the electromagnetic spectrum.
318 | Physical Property Issue | Communication or Transmission Issue | telemetry discrepancy | C62830 | FDA 1629 | Issue associated with variability of the transmission of signals, which can be characterized as telemetry channel coding, a method of processing data sent from a source to a destination so that distinct messages are created which are easily distinguishable from one another.
319 | Physical Property Issue | Communication or Transmission Issue | wireless communication issue | C104820 | FDA 3283 | Issue related to the transmission and reception of wireless medical device signals and data and the effects on the safety, effectiveness and functions of the medical device. These include issues related to wireless transmission and reception such as data integrity, timeliness, and reliability of transmitted signals and data. It can include issues with the RF wireless technology characteristics and performance (e.g., frequency, output power, range, reception), wireless quality of service, wireless coexistence, security of wireless signals and data, and electromagnetic compatibility.
320 | Physical Property Issue | Electronic Property Issue | electronic property issue | C63005 | FDA 2928 | Issue associated with any deviations from device documented performance specifications relating to devices, circuits, or systems utilizing electrons and in association with components to convert, control and condition of electric power (alternating or direct current), and of voltage and current.
321 | Physical Property Issue | Electronic Property Issue | arcing | C63301 | FDA 2583 | Issue associated with electrical current flowing through a gap between two conductive surfaces, typically resulting in a visible flash of light.
322 | Physical Property Issue | Electronic Property Issue | arcing at electrodes | C63303 | FDA 2289
323 | Physical Property Issue | Electronic Property Issue | arcing at paddles | C63302 | FDA 1032
324 | Physical Property Issue | Electronic Property Issue | circuit failure | C63277 | FDA 1089 | Issue associated with a failure of the internal network paths or electrical circuitry (i.e. electrical components, circuit boards, wiring)
325 | Physical Property Issue | Electronic Property Issue | capacitive coupling | C63283 | FDA 1079
326 | Physical Property Issue | Electronic Property Issue | electrical shorting | C63216 | FDA 2926
327 | Physical Property Issue | Electronic Property Issue | intermittent continuity | C63068 | FDA 1121
328 | Physical Property Issue | Electronic Property Issue | electrical overstress | C63218 | FDA 2924 | Issue associated with an electrical activity that exceeded the specified threshold limit of the internal integrated circuitry.
329 | Physical Property Issue | Electronic Property Issue | electromagnetic compatibility issue | C63215 | FDA 2927 | Issue associated with the ability of a system to function in its electromagnetic environment without introducing intolerable disturbances to anything in its environment.
330 | Physical Property Issue | Electronic Property Issue | electromagnetic interference (EMI) | C63214 | FDA 1194 | Issue associated with a measure of electromagnetic radiation from equipment.
331 | Physical Property Issue | Electronic Property Issue | electro-static discharge | C63213 | FDA 2149 | Issue associated with the discharge of electricity between two bodies previously electrically charged.
332 | Physical Property Issue | Electronic Property Issue | radiofrequency interference (RFI) | C62855 | FDA 2314 | Issue associated with the degradation of the reception of a wanted signal caused by RF disturbance.
333 | Physical Property Issue | Electronic Property Issue | failure to shut off | C62988 | FDA 2939 | Issue associated with the device not powering off when a shut down was requested.
334 | Physical Property Issue | Electronic Property Issue | device remains activated | C62931 | FDA 1525 | Issue associated with the device continuing to be in an active state after deactivation was requested.
335 | Physical Property Issue | Electronic Property Issue | impedance issue | C63114 | FDA 2950 | Issue associated with electrical impedance levels between device components or device and patient connections
336 | Physical Property Issue | Electronic Property Issue | high impedance | C63124 | FDA 1291 | Issue associated with higher than intended electrical impedance levels between device components or device and patient connections
337 | Physical Property Issue | Electronic Property Issue | low impedance | C63047 | FDA 2285 | Issue associated with lower than intended electrical impedance levels between device components or device and patient connections.
338 | Physical Property Issue | Electronic Property Issue | spark | C62837 | FDA 2595 | Issue associated with a flash of light related to an electrical discharge in an unexpected location.
339 | Physical Property Issue | Manufacturing or Shipping Issue Associated with Device | manufacturing or shipping issue associated with device | C63041 | FDA 2975 | Issue associated with any deviations from device documented performance specifications relating to nonconformity during manufacture to the design of an item or to specified manufacturing processes.
340 | Physical Property Issue | Manufacturing or Shipping Issue Associated with Device | component missing | C63021 | FDA 2306 | Issue associated with the absence of any part, substance, software, hardware, or raw material that is designed to be put together as part of the finished device or product.
341 | Physical Property Issue | Manufacturing or Shipping Issue Associated with Device | device markings issue | C62976 | FDA 2911 | Issue associated with the written, printed or graphic material that is affixed to a medical device or any of its containers or wrappers or accompanying the device including verbal instructions, relating to identification, technical description, and usage which are provided by the device manufacturers. Issue can include but is not limited to this material being unclear, missing worn out, incorrect or inaccurate.
342 | Physical Property Issue | Manufacturing or Shipping Issue Associated with Device | defective item | C63250 | FDA 2588 | Issue associated with having flaws or dimensional deviations greater than acceptable for the intended use of the device.
343 | Physical Property Issue | Manufacturing or Shipping Issue Associated with Device | defective component | C63023 | FDA 2292 | Issue associated with a device component having flaws of dimensional deviations greater than acceptable for the intended use.
344 | Physical Property Issue | Manufacturing or Shipping Issue Associated with Device | device damaged prior to use | C63015 | FDA 2284 | Problem associated with packaging or shipping damage prior to the use of a medical device.
345 | Physical Property Issue | Manufacturing or Shipping Issue Associated with Device | shipping damage or problem | C62925 | FDA 1570
346 | Physical Property Issue | Manufacturing or Shipping Issue Associated with Device | device misassembled during manufacturing or shipping | C62957 | FDA 2912
347 | Physical Property Issue | Manufacturing or Shipping Issue Associated with Device | incorrect device or component shipped | C63084 | FDA 2962 | Issue associated with the delivery of an inappropriate device or device component to a user facility.
348 | Physical Property Issue | Manufacturing or Shipping Issue Associated with Device | item contaminated during manufacturing or shipping | C63019 | FDA 2969 | Issue associated with the presence of any unexpected foreign substance found on the surface or in the package materials, which may affect optimal performance for its intended use.
349 | Physical Property Issue | Manufacturing or Shipping Issue Associated with Device | delivered as unsterile product | C63014 | FDA 1421 | Issue associated with the device being received in such a manner to indicate that its sterility has been compromised (e.g. sterile packaging breached, visible contaminate present)
350 | Physical Property Issue | Manufacturing or Shipping Issue Associated with Device | packaging issue | C62939 | FDA 3007 | Issue associated with the materials used to construct the cover or outer wrapping of the device.
351 | Physical Property Issue | Manufacturing or Shipping Issue Associated with Device | device packaging compromised | C63240 | FDA 2916 | Issue associated with the nonconformance to device specifications and minimum packaging requirements as the device may not be operating and functioning as intended.
352 | Physical Property Issue | Manufacturing or Shipping Issue Associated with Device | incomplete or missing packaging | C62940 | FDA 2312
353 | Physical Property Issue | Manufacturing or Shipping Issue Associated with Device | tear, rip or hole in device packaging | C62831 | FDA 2385
354 | Physical Property Issue | Manufacturing or Shipping Issue Associated with Device | unsealed device packaging | C62810 | FDA 1444
355 | Physical Property Issue | Manufacturing or Shipping Issue Associated with Device | difficult to open or remove packaging material | C63231 | FDA 2922 | Issue associated with difficulty for end-users to operate device, specifically as it relates to the opening or removal of the outer wrapping.
356 | Physical Property Issue | Manufacturing or Shipping Issue Associated with Device | failure to unwrap | C63154 | FDA 2940 | Issue associated with problems removing the materials used to construct the cover or outer wrapping of the device.
357 | Physical Property Issue | Manufacturing or Shipping Issue Associated with Device | product quality issue | C62948 | FDA 1506 | Issue associated with an inherent device and/or device component characteristic that is not satisfactory as specified or delivered
358 | Physical Property Issue | Manufacturing or Shipping Issue Associated with Device | dull | C63219 | FDA 2407
359 | Physical Property Issue | Manufacturing or Shipping Issue Associated with Device | nonstandard device or device component | C62898 | FDA 1420 | Issue associated with a material that does not meet the specifications or requirements for which it was manufactured, and has not been subjected to use or processing other than that required for its initial manufacture.
360 | Physical Property Issue | Manufacturing or Shipping Issue Associated with Device | out-of-box failure | C62944 | FDA 2311
361 | Physical Property Issue | Material Integrity Issue | material integrity issue | C62968 | FDA 2978 | Issue associated with any deviations from device documented performance specifications relating to the limited durability of all material used to construct the device.
362 | Physical Property Issue | Material Integrity Issue | break | C62973 | FDA 1069 | Issue associated with undesired damage or breakage of those materials used in device construction.
363 | Physical Property Issue | Material Integrity Issue | fracture | C63132 | FDA 1260 | Issue associated with a partial or full-thickness crack in the device materials.
364 | Physical Property Issue | Material Integrity Issue | loss of or failure to bond | C63194 | FDA 1068 | Issue associated with lack or loss of adherence between materials intended to be joined together by an adhesive.
365 | Physical Property Issue | Material Integrity Issue | solder joint failure | C62838 | FDA 2324
366 | Physical Property Issue | Material Integrity Issue | material fragmentation | C62969 | FDA 1261 | Issue associated with small pieces of the device breaking off unexpectedly
367 | Physical Property Issue | Material Integrity Issue | buckled material | C63287 | FDA 2889 | Issue associated with an undesired bulge, bend, bow, kink, or wavy condition observed in the device material resulting from compressive stresses.
368 | Physical Property Issue | Material Integrity Issue | burst | C62972 | FDA 1074 | Issue associated with the pressure inside a vessel or container rising to such a degree that the container ruptures.
369 | Physical Property Issue | Material Integrity Issue | crack | C62971 | FDA 1135 | Issue associated with an undesired separation and/or a visible opening along the length or width in the materials that are used in device construction.
370 | Physical Property Issue | Material Integrity Issue | degraded | C62970 | FDA 1153 | Issue associated with a deleterious change in the chemical structure, physical properties, or appearance in the materials that are used in device construction.
371 | Physical Property Issue | Material Integrity Issue | calcified | C63285 | FDA 1077 | Issue associated with buildup of calcium salts on the device or its components.
372 | Physical Property Issue | Material Integrity Issue | corrosion | C63263 | FDA 1131 | Issue associated with the chemical or electrochemical reaction between materials, usually a metal and its environment that produces a deterioration of the metal and its properties.
373 | Physical Property Issue | Material Integrity Issue | flaked | C63141 | FDA 1246 | Issue associated with the detachment of small pieces of the coating film of a material.
374 | Physical Property Issue | Material Integrity Issue | material erosion | C63208 | FDA 1214 | Issue associated with a progressive loss of a material from a solid surface.
375 | Physical Property Issue | Material Integrity Issue | naturally worn | C62906 | FDA 2988 | Issue associated with material damage to a surface, usually involving progressive loss or displacement of material, due to relative motion between that surface and a contacting substance or substances.
376 | Physical Property Issue | Material Integrity Issue | peeled | C62878 | FDA 1454 | Issue associated with stresses of sufficient magnitude to exceed the strength of adhesion between the device material and its coating, i.e., insulation pulling away from a wire, label peeling off.
377 | Physical Property Issue | Material Integrity Issue | delamination | C50433 | FDA 2904 | Issue associated with peeling of composite materials. Occurs when layers are separated as a result of continuous stress or impact. Results in loss of mechanical toughness.
378 | Physical Property Issue | Material Integrity Issue | pitted | C62874 | FDA 1460 | Issue associated with the corrosion of a material's surface, confined to a point or small area that takes the form of cavities.
379 | Physical Property Issue | Material Integrity Issue | unraveled material | C62811 | FDA 1664
380 | Physical Property Issue | Material Integrity Issue | bent | C63248 | FDA 2976 | Issue associated with an undesired material change in shape or property caused by external forces.
381 | Physical Property Issue | Material Integrity Issue | dent in material | C63245 | FDA 2526
382 | Physical Property Issue | Material Integrity Issue | failure to fold | C63175 | FDA 1255
383 | Physical Property Issue | Material Integrity Issue | failure to unfold or unwrap | C62949 | FDA 1669 | Issue associated with the comprising materials' deformation in that a device fails to open its wrapping or open/extend to a certain fashion or form i.e. balloon or lens.
384 | Physical Property Issue | Material Integrity Issue | material frayed | C63131 | FDA 1262 | Issue associated with the comprising materials having damaged edges.
385 | Physical Property Issue | Material Integrity Issue | material invagination | C63063 | FDA 1336
386 | Physical Property Issue | Material Integrity Issue | material rigid or stiff | C62847 | FDA 1544
387 | Physical Property Issue | Material Integrity Issue | material twisted | C62962 | FDA 2981
388 | Physical Property Issue | Material Integrity Issue | bent | C63296 | FDA 1059
389 | Physical Property Issue | Material Integrity Issue | folded | C63134 | FDA 2630
390 | Physical Property Issue | Material Integrity Issue | kinked | C63060 | FDA 1339
391 | Physical Property Issue | Material Integrity Issue | wrinkled | C62804 | FDA 2614
392 | Physical Property Issue | Material Integrity Issue | coiled | C63273 | FDA 1098
393 | Physical Property Issue | Material Integrity Issue | knotted | C63059 | FDA 1340 | None
394 | Physical Property Issue | Material Integrity Issue | material torqued | C64340 | FDA 2980 | Issue associated with a movement of forces that produces or tends to produce rotation or torsion of a material.
395 | Physical Property Issue | Material Integrity Issue | uncoiled | C62825 | FDA 1659
396 | Physical Property Issue | Material Integrity Issue | melted | C62908 | FDA 1385 | Issue associated with a solid device or device components being transformed into a molten or liquid state
397 | Physical Property Issue | Material Integrity Issue | mushroomed | C62907 | FDA 2987
398 | Physical Property Issue | Material Integrity Issue | stretched | C62835 | FDA 1601 | Issue associated with an increase or elongation in a materials' dimension.
399 | Physical Property Issue | Material Integrity Issue | material discolored | C63224 | FDA 1170 | Issue associated with an undesired streak, pattern, and/or a noticeable change in color from the rest of the materials used in device construction.
400 | Physical Property Issue | Material Integrity Issue | material disintegration | C63222 | FDA 1177 | Issue associated with material breaking into small fragments or particles
401 | Physical Property Issue | Material Integrity Issue | metal shedding debris | C62918 | FDA 1804
402 | Physical Property Issue | Material Integrity Issue | material distortion | C72894 | FDA 2977 | Issue associated with undesired physical appearance of device material resulting from excessive stress or impact.
403 | Physical Property Issue | Material Integrity Issue | material perforation | C62967 | FDA 2205 | Issue associated with an undesired material damage characterized by closely spaced punched or drilled holes.
404 | Physical Property Issue | Material Integrity Issue | cut in material | C63260 | FDA 2454
405 | Physical Property Issue | Material Integrity Issue | hole in material | C63120 | FDA 1293 | Issue associated with an opening not characteristic of the normal material.
406 | Physical Property Issue | Material Integrity Issue | material puncture | C62966 | FDA 1504 | Issue associated with the comprising material(s) being penetrated, pierced or punctured, especially with a sharp or pointed object.
407 | Physical Property Issue | Material Integrity Issue | material rupture | C62965 | FDA 1546 | Issue associated with the bursting or tearing apart of a material.
408 | Physical Property Issue | Material Integrity Issue | material protrusion | C72891 | FDA 2979 | Issue associated with undesired physical appearance of device material, specifically when material extends beyond or above the device surface.
409 | Physical Property Issue | Material Integrity Issue | material opacification | C62895 | FDA 1426 | Issue associated with an undesirable opaqueness or cloudiness.
410 | Physical Property Issue | Material Integrity Issue | material separation | C62964 | FDA 1562 | Issue associated with an undesired disassociation or breaking apart of device materials
411 | Physical Property Issue | Material Integrity Issue | scratched material | C62846 | FDA 3020 | Issue associated with an undesirable shallow cut or narrow groove in the surface of the device materials.
412 | Physical Property Issue | Material Integrity Issue | device abrasion from instrument or another object | C63037 | FDA 1387
413 | Physical Property Issue | Material Integrity Issue | split | C62963 | FDA 2537 | Issue associated with an undesired tearing of the material membrane that may be due to exceeding the tensile stress limits belonging to the materials used in device construction.
414 | Physical Property Issue | Material Integrity Issue | torn material | C62832 | FDA 3024 | Issue associated with comprising material(s) being pulled apart or into pieces by force, wrenching, or laceration.
415 | Physical Property Issue | Mechanical Issue | mechanical issue | C62961 | FDA 1384 | Issues associated with any deviations from device documented performance specifications relating to mechanical defects, including moving parts or subassemblies, etc.
416 | Physical Property Issue | Mechanical Issue | delivery system failure | C63246 | FDA 2905
417 | Physical Property Issue | Mechanical Issue | detachment of device or device component | C63242 | FDA 2907 | Issue associated with the separation of the device and/or device components from its physical construct, integrity, or chassis.
418 | Physical Property Issue | Mechanical Issue | detachment of device component | C63243 | FDA 1104 | Issue associated with the unintentional separation of the device and/or its components from something to which it is connected or attached.
419 | Physical Property Issue | Mechanical Issue | disassembly | C63225 | FDA 1168
420 | Physical Property Issue | Mechanical Issue | device or device component damaged by another device | C62951 | FDA 2915 | Issue associated with one device and/or device component causing harm to another device and/or component
421 | Physical Property Issue | Mechanical Issue | dislodged or dislocated | C62950 | FDA 2923 | Issue associated with mechanical force that displaces device and/or device components from an intended location.
422 | Physical Property Issue | Mechanical Issue | component falling | C63153 | FDA 1105 | Issue associated with an undesirable descent of a device component due to the force of gravity, usually from a higher to a lower level, usually the ground or floor.
423 | Physical Property Issue | Mechanical Issue | device or device fragments location unknown | C63038 | FDA 2590
424 | Physical Property Issue | Mechanical Issue | escape | C63204 | FDA 2931 | Escape Medical Device Problem
425 | Physical Property Issue | Mechanical Issue | tip-over | C62921 | FDA 2589
426 | Physical Property Issue | Mechanical Issue | slippage of device or device component | C62840 | FDA 1584
427 | Physical Property Issue | Mechanical Issue | leak | C63058 | FDA 1354 | Issue associated with the escape of a liquid or gas from the vessel or container in which it is housed.
428 | Physical Property Issue | Mechanical Issue | fluid leak | C63137 | FDA 1250
| Physical Property Issue | Mechanical Issue | fluid injection |  |  |  | endologix
429 | Physical Property Issue | Mechanical Issue | hot oil leak | C63119 | FDA 1298
430 | Physical Property Issue | Mechanical Issue | peri-valvular leak | C62876 | FDA 1457 | Issue associated with the escape of blood around a heart valve, particularly around its leaflets.
431 | Physical Property Issue | Mechanical Issue | strikethrough | C62834 | FDA 2538
432 | Physical Property Issue | Mechanical Issue | gas leak | C63127 | FDA 2946 | Issue associated with the unintended escape of a gas from the container in which it is housed.
433 | Physical Property Issue | Mechanical Issue | air leak | C63311 | FDA 1008
434 | Physical Property Issue | Mechanical Issue | gel leak | C63316 | FDA 1267
435 | Physical Property Issue | Mechanical Issue | radiation leak | C62857 | FDA 1357
436 | Physical Property Issue | Mechanical Issue | mechanical jam | C63039 | FDA 2983 | Issue associated with a problem that prevents or restricts the motion of the device or its components.
437 | Physical Property Issue | Mechanical Issue | looping | C63056 | FDA 1370
438 | Physical Property Issue | Mechanical Issue | mechanics altered | C62960 | FDA 2984 | Issue associated with device mechanical functioning of machinery, moving parts or tools of the device being changed or modified.
439 | Physical Property Issue | Mechanical Issue | failure to align | C63197 | FDA 2522 | Issue associated with a circuit, equipment, or system whereby its functions fail to be properly synchronized or its relative positions properly oriented.
440 | Physical Property Issue | Mechanical Issue | failure to cut | C63188 | FDA 2587
441 | Physical Property Issue | Mechanical Issue | failure to cycle | C63187 | FDA 1142 | Issue associated with the device failing to complete a series of processes or events
442 | Physical Property Issue | Mechanical Issue | failure to form staple | C64328 | FDA 2579
443 | Physical Property Issue | Mechanical Issue | misfire | C62914 | FDA 2532 | Issue associated with failure of device to discharge its load (e.g. surgical stapler failed to partially or completely deploy its staples)
444 | Physical Property Issue | Mechanical Issue | audible noise | C99179 | FDA 3273 | Issue associated with any unintended sound which emanates from a device (for example, squeaking from two parts rubbing together or buzzing sounds from electrical components).
445 | Physical Property Issue | Mechanical Issue | osseointegration issue | C62886 | FDA 3003 | Issue associated with interconnection between bone tissue and implanted device.
446 | Physical Property Issue | Mechanical Issue | failure to osseointegrate | C63171 | FDA 1863 | Issue associated with the failure to see direct anchorage of an implant by the formation of bony tissue around the implant without the growth of fibrous tissue at the bone-implant interface.
447 | Physical Property Issue | Mechanical Issue | loss of osseointegration | C63053 | FDA 2408
448 | Physical Property Issue | Mechanical Issue | positioning issue | C63034 | FDA 3009 | Issue associated with the movement of the device and/or device components to an intended location.
449 | Physical Property Issue | Mechanical Issue | difficult to advance | C63235 | FDA 2920 | Issue associated with difficulty moving the device or its components to an intended location (e.g. difficulty in advancing guidewire)
450 | Physical Property Issue | Mechanical Issue | difficult to insert | C63233 | FDA 1316 | Issue associated with problems introducing or inserting the device, even if the user is operating the device in accordance with the instructions for use or labeling.
451 | Physical Property Issue | Mechanical Issue | difficult to position | C63230 | FDA 1467 | Issue associated with the use of the device in terms of user experiencing difficulty to put device and/or device components in place, even if the operation is being performed according to labeled instructions for use.
452 | Physical Property Issue | Mechanical Issue | difficult to remove | C63228 | FDA 1528 | Issue associated with the use of the device in terms of user experiencing difficulty to take out or get rid of a device and/or device components, even if the operation is being performed according to labeled instructions for use.
453 | Physical Property Issue | Mechanical Issue | entrapment of device or device component | C63210 | FDA 1212 | Issue associated with the device and/or device accessories caught within patient vasculature, tissue, or other devices or device components.
454 | Physical Property Issue | Mechanical Issue | failure to advance | C63198 | FDA 2524 | Issue associated with failure to move the device or its components to an intended location
455 | Physical Property Issue | Mechanical Issue | malposition of device | C63043 | FDA 2616 | Issue associated with device being positioned in a location other than intended or specified.
456 | Physical Property Issue | Mechanical Issue | physical resistance | C62875 | FDA 2578
457 | Physical Property Issue | Mechanical Issue | sticking | C62836 | FDA 1597
458 | Physical Property Issue | Mechanical Issue | retraction problem | C62929 | FDA 1536 | Issue associated with drawing back the device and/or device component to an intended location.
459 | Physical Property Issue | Mechanical Issue | structural problem | C62923 | FDA 2506 | Issue associated with the basic physical construction or physical make up of the device.
460 | Physical Property Issue | Mechanical Issue | collapse | C63272 | FDA 1099 | Issue associated with the buckling or crushing of material from external forces.
461 | Physical Property Issue | Mechanical Issue | difficult to fold or unfold | C63234 | FDA 1254 | Issue associated with the use of the device in terms of user experiencing difficulty to close or to spread out/extend length of a device, even if the operation is being performed according to labeled instructions for use.
462 | Physical Property Issue | Mechanical Issue | difficult to open or close | C63012 | FDA 2921 | Issue associated with the use of the device in terms of user experiencing difficulty opening and closing the device, even if the operation is being performed according to labeled instructions for use.
463 | Physical Property Issue | Mechanical Issue | incomplete coaptation | C63090 | FDA 2507 | Issue associated with the heart valve leaflet not closing properly.
464 | Physical Property Issue | Mechanical Issue | unintended movement | C62814 | FDA 3026 | Issue associated with an undesired movement of device, which may be related to device malfunction, misdiagnosis, or mistreatment.
465 | Physical Property Issue | Mechanical Issue | migration of device or device component | C62917 | FDA 1395 | Issue associated with an undesired movement of device and/or device components, related to its movement away from or dislodging from a source.
466 | Physical Property Issue | Mechanical Issue | expulsion | C50554 | FDA 2933 | Issue associated with the unintended ejection of device component materials.
467 | Physical Property Issue | Mechanical Issue | unintended arm motion | C62820 | FDA 1033
468 | Physical Property Issue | Mechanical Issue | unintended collision | C62819 | FDA 1429 | Issue associated with device or device components impacting with another object.
469 | Physical Property Issue | Mechanical Issue | unintended head motion | C62816 | FDA 1284
470 | Physical Property Issue | Mechanical Issue | unintended system motion | C62813 | FDA 1430 | Issue associated with any motion of the system or components that was not initiated by the user.
471 | Physical Property Issue | Mechanical Issue | unstable | C62809 | FDA 1667 | Issue associated with the connection of device and/or device components not possessing stability, or is unbounded.
472 | Physical Property Issue | Mechanical Issue | vibration | C62806 | FDA 1674 | Issue associated with an oscillation wherein the quantity is a parameter that defines the motion of a mechanical system.
473 | Physical Property Issue | Mechanical Issue | extrusion | C50557 | FDA 2934 | Issue associated with the unintended projection of device component materials.
474 | Physical Property Issue | Optical Issue | optical issue | C62953 | FDA 3001 | Issue associated with problems transmitting visible light affecting the quality of the image transmitted or otherwise affecting the intended application of the visible light path.
475 | Physical Property Issue | Optical Issue | light interference | C63057 | FDA 2972
476 | Physical Property Issue | Optical Issue | optical decentration | C62891 | FDA 1360 | Issue associated with being off-center of optical lenses.
477 | Physical Property Issue | Optical Issue | misfocusing | C62913 | FDA 1401
478 | Physical Property Issue | Optical Issue | optical discoloration | C62890 | FDA 2999 | Issue associated with an optical streak or other pattern on the signal surface that causes a noticeable change of color from the rest of the pultruded surface.
479 | Physical Property Issue | Optical Issue | optical distortion | C62889 | FDA 3000 | Issue associated with an optical defect in an image-forming system whereby the image is not the shape of an ideal image of the object.
480 | Physical Property Issue | Optical Issue | optical obstruction | C62888 | FDA 3002 | Issue associated with the blocking of optical devices, e.g., visual pathways.
481 | Physical Property Issue | Temperature Issue | temperature issue | C62922 | FDA 3022 | Issue associated with the device producing unintended temperatures (for issues related to environmental/ambient temperatures refer to code C63307)
482 | Physical Property Issue | Temperature Issue | burn of device or device component | C63286 | FDA 1071 | Issues associated with a discoloration or destruction as a result of thermal decomposition of the device or its components
483 | Physical Property Issue | Temperature Issue | charred | C63279 | FDA 1086
484 | Physical Property Issue | Temperature Issue | excessive cooling | C63202 | FDA 2932 | Issue associated with the device or its components producing temperatures that are lower than specified.
485 | Physical Property Issue | Temperature Issue | fire | C63143 | FDA 1245 | Issues associated with the combustion of device components, resulting in any of the following: light, flame, smoke.
486 | Physical Property Issue | Temperature Issue | flare or flash | C63140 | FDA 2942 | Issue associated with device-related burn with an unsteady flame.
487 | Physical Property Issue | Temperature Issue | insufficient cooling | C63071 | FDA 1130 | Issue associated with device or device parts being insufficiently cool in device active (working) or/and non-active nonworking) state.
488 | Physical Property Issue | Temperature Issue | insufficient heating | C63070 | FDA 1287 | Issue associated with the device or its components producing temperatures that are not as high as what is specified.
489 | Physical Property Issue | Temperature Issue | overheating of device or device component | C62883 | FDA 1437 | Issue associated with the device producing high temperatures, such that its operation is compromised (e.g. overheating that produces melting of components or automatic shutdown)
490 | Physical Property Issue | Temperature Issue | unintended magnet quench | C62815 | FDA 1377
491 | Physical Property Issue | Temperature Issue | smoking | C63003 | FDA 1585 | Issue associated with a cloud of vapor or gas generated from the device, generally associated after a fire or a burn.
492 | Missing Value Reason | No Code Available | no code available | C64343 | FDA 3191 | The device problem is not adequately described by any other term. Note: this code must not be used unless there is no other feasible code. The preferred term should be documented when submitting an adverse event report. This information will be used to determine if a new term should be added to the code table.
493 | Missing Value Reason | No Information | no information | C53269 | FDA 3190 | An adverse event appears to have occurred but there is not yet enough information available to classify the device problem.
494 | Missing Value Reason | No Known Device Problem | no known device problem | C76126 | FDA 2993
495 | Missing Value Reason | Not Applicable | not applicable | C48660 | FDA 3189

## hazard lists
Therapy
Muscle_Stimulator
Hazard | Failure_Cause | Hazardous_Situation | Severity_Level | Harm | Mitigation
Aborted Charge
Accessory Incompatible
Air Leak
Alarm not Visible
Ambient Noise Issue
Ambient Temperature Issue
Application Interface becomes nonfunctional or program exits abnormally
Application Network Issue
Application Program Issue
Application Program Version or Upgrade Problem
Application Security Issue
Arcing
Arcing at Electrodes
Arcing at Paddles
Artifact
Aspiration Issue
Bacterial Contamination of Device
Battery Impedance Issue
Battery Issue
Bent
Biocompatibility Issue
Biofilm Coating in Device
Biological Environmental Factor
Bleed Back
Blockage within Device or Device Component
Blocked Connection
Break
Buckled Material
Burn of Device or Device Component
Burst
Calcified
Calibration Error
Calibration Issue
Capacitive Coupling
Capturing Issue
Changes in Ambient Temperature in Device Environment
Charging Issue
Charred
Chemical Issue
Chemical Spillage
Circuit Failure
Clumping in Device or Device Ingredient
Coagulation in Device or Device Ingredient
Coiled
Collapse
Communication or Transmission Issue
Component Falling
Component Incompatible
Component Missing
Component or Accessory Incompatibility
Computer Operating System Issue
Computer Software Issue
Computer System Security Issue
Connection Issue
Contamination During Use
Contamination of Device Ingredient or Reagent
Continuous Firing
Corrosion
Crack
Cross Reactivity
Cut in Material
Data Back-Up Problem
Data Issue
Date-related software issue
Decoupling
Decrease in Pressure
Decrease in Suction
Decreased Pump Speed
Decreased Sensitivity
Defective Alarm
Defective Component
Defective Item
Deflation Issue
Degraded
Delamination
Delayed Alarm
Delayed Charge Time
Delivered As Unsterile Product
Delivery System Failure
Dent in Material
Deployment Issue
Detachment of Device Component
Detachment of Device or Device Component
Device Abrasion from Instrument or another Object
Device Alarm System Issue
Device Cleaning Issue
Device Clogged
Device Contamination with Biological Material
Device Contamination with Blood or Blood Product
Device Damaged Prior to Use
Device Difficult to Maintain
Device Difficult to Setup or Prepare
Device Disinfection or Sterilization Issue
Device displays error message
Device Emits Odor
Device Expiration Issue
Device Handling Issue
Device Ingredient or Reagent Issue
Device Inoperable
Device Maintenance Issue
Device Markings Issue
Device Misassembled during Manufacturing or Shipping
Device Operates Differently than Expected
Device or Device Component Damaged by another Device
Device or Device Fragments Location Unknown
Device Packaging Compromised
Device Remains Activated
Device Rinsing Issue
Device Sensing Issue
Device Stops Intermittently
Device Unsafe to Use in Environment
Device-Device Incompatibility
Difficult or Delayed Activation
Difficult to Advance
Difficult to Deploy
Difficult to Flush
Difficult to Fold or unfold
Difficult to Insert
Difficult to Interrogate
Difficult to Open or Close
Difficult to Open or Remove Packaging Material
Difficult to Position
Difficult to Program or Calibrate
Difficult to Remove
Disassembly
Disconnection
Disinfection or Sterilization Issue at User Location
Dislodged or Dislocated
Display Misread
Dose Calculation Error due to Software Problem
Dull
Electrical Issue
Electrical Overstress
Electrical Power Problem
Electrical Shorting
Electromagnetic Compatibility Issue
Electromagnetic Interference (EMI)
Electronic Property Issue
Electro-Static Discharge
Emergency Power Failure
Energy Output to Patient Tissue Incorrect
Energy Spectrum Incorrect
Entrapment of Device or Device Component
Environmental Control or Utility Issue
Environmental Particulates
Erratic Display
Escape
Excess Flow or Over-infusion
Excessive Cooling
Expiration Date Error
Expulsion
Extrusion
Facilities Issue
Fail-Safe Design Failure
Fail-Safe Mechanism Issue
Failure of device to Self-Test
Failure to adhere or bond
Failure to Advance
Failure to Align
Failure to Analyze Rhythm
Failure to Auto Stop
Failure to Back-Up
Failure to Calibrate
Failure to Capture
Failure to Charge
Failure to Conduct
Failure to Convert Rhythm
Failure to Convert to Back-Up
Failure to Cut
Failure to Cycle
Failure to Deliver
Failure to Deliver Counter-shock
Failure to Deliver Energy
Failure to Deploy
Failure to Discharge
Failure to Disconnect
Failure to Disinfect
Failure to Expand
Failure to Fire
Failure to Flush
Failure to Fold
Failure to form staple
Failure to Infuse
Failure to Interrogate
Failure to Obtain Samples
Failure to Osseointegrate
Failure to Pace or Properly Pace
Failure to Power-Up
Failure to Prime
Failure to Pump
Failure to Read Input Signal
Failure to Recalibrate
Failure to Remove Enzymatic Cleaner
Failure to Reset
Failure to Run on AC/DC
Failure to Run on Portable Mode
Failure to Select Signal
Failure to Sense
Failure to Separate
Failure to Service
Failure to Shock or Properly Shock
Failure to Shut Off
Failure to Spike
Failure to Sterilize
Failure to Transmit Record
Failure to Unfold or Unwrap
Failure to Unwrap
Failure to Zero
False Alarm
False Device Output
False Negative Result
False Positive Result
False Reading from Device Noncompliance
Filling Problem
Filtration Issue
Fire
Fitting Problem
Flaked
Flare or Flash
Flashpoint Threshold Met
Fluid Leak
Fogging
Folded
Foreign Material Present in Device
Fracture
Free or Unrestricted Flow
Fumes or Vapors
Fungus in Device Environment
Gas Leak
Gel Leak
Gradient Increase
Grounding Malfunction
High Battery Impedance
High Capture Threshold
High Impedance
High pH
High Readings
High Sensing Threshold
High Test Results
Hole in Material
Hot Oil Leak
Human Factors Issue
Human-Device Interface Issue
Image Display Error
Image Orientation Incorrect
Image Resolution Poor
Image Reversal
Impedance Issue
Imprecision
Improper Alarm
Improper Chemical Reaction
Improper Device Output
Improper Flow or Infusion
Improper Gas Output
Improper or Incorrect Procedure or Method
Inability to Auto-fill
Inability to Irrigate
Inaccurate Delivery
Inaccurate Dispensing
Inaccurate Flowrate
Inaccurate Synchronization
Inadequate Filtration Process
Inadequate instructions for healthcare professional
Inadequate instructions for non-healthcare professional
Inadequate Lighting
Inadequate Service
Inadequate Storage
Inadequate Training
Inadequate User Interface
Inappropriate or Unexpected Reset
Inappropriate Prompts
Inappropriate Shock
Inaudible Voice Prompts
Incompatibility Problem
Incomplete Coaptation
Incomplete or Missing Packaging
Incorrect Device or Component Shipped
Incorrect Display
Incorrect Error Code
Incorrect Interpretation of Rhythm
Incorrect Measurement
Incorrect or Inadequate Result
Incorrect or Inadequate Test Results
Incorrect Software Programming Calculations
Increase in Pressure
Increase in Suction
Increased Pump Speed
Increased Sensitivity
Inflation Issue
Infusion or Flow Issue
Installation-Related Problem
Instruction for Use Issue
Insufficient Cooling
Insufficient Flow or Underinfusion
Insufficient Heating
Intermittent Capture
Intermittent Continuity
Intermittent Infusion
Intermittent Shock
Invalid Sensing
Issue with displayed error message
Item Contaminated during Manufacturing or Shipping
Kinked
Knotted
Lack of Maintenance documentation or Guidelines
Leak
Light Interference
Looping
Loose or Intermittent Connection
Loss of Data
Loss of or Failure to Bond
Loss of Osseointegration
Loss of Power
Loss of Threshold
Low Audible Alarm
Low Battery
Low Battery Impedance
Low Impedance
Low Readings
Low Sensing Threshold
Low Test Results
Maintenance Does not comply to Manufacturers Recommendations
Malposition of Device
Manufacturing or Shipping Issue Associated with Device
Material Discolored
Material Disintegration
Material Distortion
Material Erosion
Material Fragmentation
Material Frayed
Material Integrity Issue
Material Invagination
Material Opacification
Material Perforation
Material Protrusion
Material Puncture
Material Rigid or Stiff
Material Rupture
Material Separation
Material Torqued
Material Twisted
Measurement System Incompatibility
Mechanical Issue
Mechanical Jam
Mechanics Altered
Medical Gas Supply Problem
Medication Error
Melted
Metal Shedding Debris
Migration of Device or Device Component
Misassembled
Misassembled by Users
Misconnection
Misfire
Misfocusing
Missing Test Results
Moisture Damage
Moisture or Humidity Problem
Mushroomed
Naturally Worn
No Code Available
No Device Output
No Display or Display Failure
No Fail-Safe Mechanism
No Flow
No Information
No Known Device Problem
No Pacing
No Pressure
No Voice Prompts
Noise, Audible
Nonstandard Device or Device Component
Not Applicable
Not Audible Alarm
Obstruction within Device
Occlusion within Device
Operating System Becomes Nonfunctional
Operating System Version or Upgrade Problem
Optical Decentration
Optical Discoloration
Optical Distortion
Optical Issue
Optical Obstruction
Osseointegration Issue
Out-Of-Box Failure
Output above Specifications
Output below Specifications
Output Energy Incorrect
Output Issue
Overcorrection
Overfill
Overheating of Device or Device Component
Over-sensing
Pacer Found in Back-Up Mode
Pacing Asynchronously
Pacing Inadequately
Pacing Intermittently
Packaging Issue
Parameter Calculation Error due to Software Problem
Particulates
Patient Data Issue
Patient-Device Incompatibility
Peeled
Peri-valvular Leak
Physical Property Issue
Physical Resistance
Pitted
Plunge
Pocket Stimulation
Poor Quality Image
Positioning Issue
Power Calculation Error due to Software Problem
Power Conditioning Issue
Power Source Issue
Pre- or Post-Pumping Problem
Precipitate in Device or Device Ingredient
Premature Deployment
Premature Discharge of Battery
Premature Elective Replacement Indicator
Premature End-of-Life Indicator
Premature Indicator Activation
Pressure Issue
Problem with Software Installation
Product Quality Issue
Programming Issue
Prompts will not Clear
Protective Measure Issue
Pumping Issue
Pumping Stopped
Radiation Leak
Radiation overexposure
Radiation underexposure
Radio Signal Problem
Radiofrequency Interference (RFI)
Reflux within Device
Rejection
Reset Issue
Residue after Decontamination
Restricted Flowrate
Retraction Problem
Scratched Material
Sediment, Precipitate or Deposit in Device or Device Ingredient
Self-Activation or Keying
Sensing Intermittently
Shelf Life Exceeded
Shielding Failure
Shipping Damage or Problem
Short Fill
Size Incorrect for patient
Slippage of Device or Device Component
Smoking
Solder Joint Failure
Spark
Split
Sticking
Stretched
Strikethrough
Structural Problem
Suction Issue
Tear, Rip or Hole in Device Packaging
Telemetry Discrepancy
Temperature Issue
Therapeutic or Diagnostic Output Failure
Therapy delivered to incorrect body area
Tidal Volume Fluctuations
Tip-over
Torn Material
Ultra-filtration
Unable to Obtain Readings
Unauthorized Access to Computer System
Uncoiled
Under-sensing
Unexpected Therapeutic Results
Unintended Arm Motion
Unintended Collision
Unintended Ejection
Unintended Energization
Unintended Head Motion
Unintended Magnet Quench
Unintended Movement
Unintended System Motion
Unraveled Material
Unsealed Device Packaging
Unstable
Unstable Capture Threshold
Use of Device Issue
Use of Incorrect Control Settings
User used incorrect product for intended use
Ventilation Issue in Device Environment
Vibration
Volume Accuracy Issue
Wireless Communication Issue
Wrinkled

## hazard-harm training data
Product_Category | Product_SubCategory | Hazard | Failure_Cause | Hazardous_Situation | Harm | Severity | Mitigation
CRM | Pacer | Ambient Noise Issue |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | Ambient Noise Issue |  | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Application Program Issue |  | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Application Program Issue |  | Abnormal Cardiac Pacing | Discomfort | 2
CRM | Pacer | Application Program Issue |  | Abnormal Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Application Program Issue |  | Loss of Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | Application Program Issue |  | Loss of Cardiac Pacing | Prolonged Surgery | 2
CRM | Pacer | Application Program Issue |  | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Application Program Issue |  | Loss of Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Application Program Version Or Upgrade Problem |  | Abnormal Cardiac Pacing | Arrhythmia | 4
CRM | Pacer | Application Program Version Or Upgrade Problem |  | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Application Program Version Or Upgrade Problem |  | Abnormal Cardiac Pacing | Dizziness | 2
CRM | Pacer | Application Program Version Or Upgrade Problem |  | Abnormal Defibrillation Therapy | Arrhythmia | 4
CRM | Pacer | Application Program Version Or Upgrade Problem |  | Incomplete or Misleading Diagnostic Data | Death | 5
CRM | Pacer | Application Program Version Or Upgrade Problem |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | Application Program Version Or Upgrade Problem |  | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Application Program Version Or Upgrade Problem |  | Loss of Defibrillation Therapy | Arrhythmia | 4
CRM | Pacer | Artifact |  | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Artifact |  | Inappropriate Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Artifact |  | Inappropriate Defibrillation Therapy | Discomfort | 2
CRM | Pacer | Artifact |  | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Battery Issue |  | Device Heating | Burn, Thermal | 3
CRM | Pacer | Battery Issue |  | Device Heating | Burn, Thermal | 2
CRM | Pacer | Battery Issue |  | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Biocompatibility Issue |  | Biocompatibility Issue | Hypersensitivity | 3
CRM | Pacer | Biocompatibility Issue |  | Biocompatibility Issue | Hematoma | 2
CRM | Pacer | Biofilm Coating In Device |  | Surgical Complications | Seroma | 2
CRM | Pacer | Burst |  | Abnormal Defibrillation Therapy | Arrhythmia | 2
CRM | Pacer | Capturing Issue |  | Loss of Cardiac Pacing | Arrhythmia | 5
CRM | Pacer | Capturing Issue |  | Loss of Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | Capturing Issue |  | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Capturing Issue |  | Unintentional Extra Cardiac Stimulation | Discomfort | 2
CRM | Pacer | Circuit Failure |  | Abnormal Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Circuit Failure |  | Loss of Cardiac Pacing | Additional Surgical Procedure | 3
CRM | Pacer | Communication or Transmission Issue |  | Inaccessible Diagnostic Data | Additional Surgical Procedure | 3
CRM | Pacer | Communication or Transmission Issue |  | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Communication or Transmission Issue |  | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Component Or Accessory Incompatibility |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | Computer Operating System Issue |  | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Computer Operating System Issue |  | Loss of Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | Computer Operating System Issue |  | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Computer Software Issue |  | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Crack |  | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Date Related Software Issue |  | Inaccessible Diagnostic Data | Additional Surgical Procedure | 3
CRM | Pacer | Defective Component |  | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Defective Component |  | Application of Excessive Force | Prolonged Surgery | 1
CRM | Pacer | Defective Component |  | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Defective Component |  | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Degraded |  | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Degraded |  | Unintentional Extra Cardiac Stimulation | Additional Surgical Procedure | 3
CRM | Pacer | Degraded |  | Unintentional Extra Cardiac Stimulation | Burn, Thermal | 2
CRM | Pacer | Delayed Charge Time |  | Abnormal Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Delayed Charge Time |  | Abnormal Defibrillation Therapy | Arrhythmia | 2
CRM | Pacer | Device Abrasion From Instrument Or Another Object |  | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Device Alarm System Issue |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | Device Alarm System Issue |  | Loss of Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | Device Alarm System Issue |  | Loss of Patient Notification | Additional Surgical Procedure | 3
CRM | Pacer | Device Damaged Prior To Use |  | Abnormal Cardiac Pacing | Prolonged Surgery | 1
CRM | Pacer | Device Damaged Prior To Use |  | Abnormal Defibrillation Therapy | Prolonged Surgery | 1
CRM | Pacer | Device Damaged Prior To Use |  | Loss of Cardiac Pacing | Additional Surgical Procedure | 3
CRM | Pacer | Device Difficult To Maintain |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | Device Difficult To Setup Or Prepare |  | Abnormal Cardiac Pacing | Prolonged Surgery | 1
CRM | Pacer | Device Difficult To Setup Or Prepare |  | Confusion | Prolonged Surgery | 1
CRM | Pacer | Device Difficult To Setup Or Prepare |  | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Device Maintenance Issue |  | Inappropriate Defibrillation Therapy | Discomfort | 2
CRM | Pacer | Device Or Device Component Damaged By Another Device |  | Application of Excessive Force | Prolonged Surgery | 1
CRM | Pacer | Device Or Device Component Damaged By Another Device |  | Loss of Cardiac Pacing | Arrhythmia | 5
CRM | Pacer | Device Or Device Component Damaged By Another Device |  | Loss of Cardiac Pacing | Additional Surgical Procedure | 3
CRM | Pacer | Device Or Device Component Damaged By Another Device |  | Loss of Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Device Sensing Issue |  | Abnormal Cardiac Pacing | Additional Surgical Procedure | 3
CRM | Pacer | Device Sensing Issue |  | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Device Sensing Issue |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | Device Sensing Issue |  | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Device Sensing Issue |  | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Difficult To Advance |  | Surgical Complications | Embolism | 3
CRM | Pacer | Difficult To Deploy |  | Confusion | Additional Surgical Procedure | 3
CRM | Pacer | Difficult To Insert |  | Application of Excessive Force | Prolonged Surgery | 1
CRM | Pacer | Difficult To Insert |  | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Disinfection Or Sterilization Issue At User Location |  | Lack of Sterility | Infection | 3
CRM | Pacer | Dislodged or Dislocated |  | Abnormal Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Dislodged or Dislocated |  | Loss of Cardiac Pacing | Arrhythmia | 5
CRM | Pacer | Dislodged or Dislocated |  | Loss of Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | Dislodged or Dislocated |  | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Electrical Shorting |  | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Electrical Shorting |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | Electrical Shorting |  | Loss of Cardiac Pacing | Additional Surgical Procedure | 3
CRM | Pacer | Electrical Shorting |  | Loss of Cardiac Pacing | Prolonged Surgery | 1
CRM | Pacer | Electrical Shorting |  | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Electrical Shorting |  | Loss of Defibrillation Therapy | Arrhythmia | 4
CRM | Pacer | Electrical Shorting |  | Loss of Defibrillation Therapy | Additional Surgical Procedure | 4
CRM | Pacer | Electro Magnetic Interference |  | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Electro Magnetic Interference (EMI) |  | Abnormal Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | Electro Magnetic Interference (EMI) |  | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Electro Magnetic Interference (EMI) |  | Inappropriate Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Electro Magnetic Interference (EMI) |  | Inappropriate Defibrillation Therapy | Discomfort | 2
CRM | Pacer | Electro Magnetic Interference (EMI) |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | Electro Magnetic Interference (EMI) |  | Loss of Cardiac Pacing | Arrhythmia | 5
CRM | Pacer | Electro Magnetic Interference (EMI) |  | Loss of Cardiac Pacing | Arrhythmia | 4
CRM | Pacer | Electro Magnetic Interference (EMI) |  | Loss of Cardiac Pacing | Additional Surgical Procedure | 3
CRM | Pacer | Electro Magnetic Interference (EMI) |  | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Electro Magnetic Interference (EMI) |  | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Electro Magnetic Interference (EMI) |  | Loss of Defibrillation Therapy | Arrhythmia | 4
CRM | Pacer | Electro Magnetic Interference (EMI) |  | Loss of Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Electro Magnetic Interference (EMI) |  | Unintentional Extra Cardiac Stimulation | Discomfort | 2
CRM | Pacer | Electro Static Discharge |  | Abnormal Cardiac Pacing | No consequences or impact to patient | 1
CRM | Pacer | Electro Static Discharge |  | Abnormal Defibrillation Therapy | No consequences or impact to patient | 1
CRM | Pacer | Electronic Property Issue |  | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Energy Output To Patient Tissue Incorrect |  | Device Heating | Burn, Thermal | 2
CRM | Pacer | Energy Output To Patient Tissue Incorrect |  | Unintentional Extra Cardiac Stimulation | Discomfort | 2
CRM | Pacer | Entrapment Of Device Or Device Component |  | Surgical Complications | Hemorrhage Serious | 3
CRM | Pacer | Entrapment Of Device Or Device Component |  | Surgical Complications | Pneumothorax | 2
CRM | Pacer | Environmental Control Or Utility Issue |  | Abnormal Cardiac Pacing | No consequences or impact to patient | 1
CRM | Pacer | Failure To Align |  | Inappropriate Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Failure To Align |  | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Failure To Analyze Rhythm |  | Abnormal Defibrillation Therapy | Arrhythmia | 2
CRM | Pacer | Failure To Analyze Rhythm |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | Failure To Analyze Rhythm |  | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Failure To Auto Stop |  | Inappropriate Defibrillation Therapy | Discomfort | 2
CRM | Pacer | Failure To Capture |  | Abnormal Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | Failure To Capture |  | Abnormal Cardiac Pacing | Heart Failure | 3
CRM | Pacer | Failure To Capture |  | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Failure To Capture |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | Failure To Capture |  | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Failure To Discharge |  | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Failure To Interrogate |  | Application of Excessive Force | Prolonged Surgery | 1
CRM | Pacer | Failure To Interrogate |  | Inaccessible Diagnostic Data | Additional Surgical Procedure | 3
CRM | Pacer | Failure To Interrogate |  | Inaccessible Diagnostic Data | Intervention or Procedure | 2
CRM | Pacer | Failure To Interrogate |  | Inappropriate Defibrillation Therapy | Discomfort | 2
CRM | Pacer | Failure To Interrogate |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | Failure To Pace Or Properly Pace |  | Abnormal Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | Failure To Pace Or Properly Pace |  | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Failure To Pace Or Properly Pace |  | Abnormal Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Failure To Pace Or Properly Pace |  | Abnormal Cardiac Pacing | Discomfort | 1
CRM | Pacer | Failure To Pace Or Properly Pace |  | Inconvenience | Prolonged Surgery | 1
CRM | Pacer | Failure To Sense |  | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Failure To Sense |  | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Failure To Service |  | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Failure To Shock Or Properly Shock |  | Inappropriate Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Failure To Spike |  | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Failure To Spike |  | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Failure To Sterilize |  | Lack of Sterility | Infection | 5
CRM | Pacer | Failure To Sterilize |  | Lack of Sterility | Infection | 3
CRM | Pacer | False Positive Result |  | Incomplete or Misleading Diagnostic Data | Additional Surgical Procedure | 3
CRM | Pacer | False Positive Result |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | False Reading From Device Non Compliance |  | Inappropriate Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | False Reading From Device Non Compliance |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | Fluid Leak |  | Extravascular Foreign Body | Additional Surgical Procedure | 4
CRM | Pacer | Fluid Leak |  | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Fracture |  | Loss of Cardiac Pacing | Heart Failure | 3
CRM | Pacer | Fracture |  | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | High Capture Threshold |  | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | High Capture Threshold |  | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | High Impedance |  | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | High Sensing Threshold |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | Human Device Interface Issue |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | Human Factors Issue |  | Abnormal Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Human Factors Issue |  | Application of Excessive Force | Dissection | 3
CRM | Pacer | Human Factors Issue |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | Human Factors Issue |  | Loss of Cardiac Pacing | Arrhythmia | 5
CRM | Pacer | Human Factors Issue |  | Loss of Cardiac Pacing | Arrhythmia | 4
CRM | Pacer | Human Factors Issue |  | Loss of Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | Human Factors Issue |  | Loss of Cardiac Pacing | Heart Failure | 3
CRM | Pacer | Human Factors Issue |  | Loss of Cardiac Pacing | Additional Surgical Procedure | 3
CRM | Pacer | Human Factors Issue |  | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Human Factors Issue |  | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Human Factors Issue |  | Surgical Complications | Cardiac Perforation | 4
CRM | Pacer | Human Factors Issue |  | Unintentional Extra Cardiac Stimulation | Discomfort | 2
CRM | Pacer | Image Display Error |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | Impedance Issue |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | Improper Alarm |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | Improper Or Incorrect Procedure Or Method |  | Loss of Cardiac Pacing | Heart Failure | 3
CRM | Pacer | Improper Or Incorrect Procedure Or Method |  | Loss of Cardiac Pacing | Additional Surgical Procedure | 3
CRM | Pacer | Improper Or Incorrect Procedure Or Method |  | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Improper Or Incorrect Procedure Or Method |  | Unintentional Extra Cardiac Stimulation | Discomfort | 2
CRM | Pacer | Inaccurate Synchronization |  | Abnormal Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | Inaccurate Synchronization |  | Abnormal Cardiac Pacing | Additional Surgical Procedure | 3
CRM | Pacer | Inaccurate Synchronization |  | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Inaccurate Synchronization |  | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Inadequate Instructions For Healthcare Professional |  | Inappropriate Programmed Settings | Prolonged Surgery | 1
CRM | Pacer | Inadequate Service |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | Inadequate Service |  | Loss of Cardiac Pacing | Arrhythmia | 5
CRM | Pacer | Inadequate Service |  | Loss of Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | Inadequate Service |  | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Inadequate Service |  | Loss of Defibrillation Therapy | Additional Surgical Procedure | 3
CRM | Pacer | Inadequate Service |  | Unintentional Extra Cardiac Stimulation | Discomfort | 2
CRM | Pacer | Inadequate Training |  | Inappropriate Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Inadequate Training |  | Inappropriate Defibrillation Therapy | Discomfort | 2
CRM | Pacer | Inadequate Training |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | Inappropriate Shock |  | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Incorrect Interpretation Of Rhythm |  | Abnormal Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Incorrect Interpretation Of Rhythm |  | Inappropriate Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Incorrect Interpretation Of Rhythm |  | Inappropriate Defibrillation Therapy | Discomfort | 2
CRM | Pacer | Incorrect Interpretation Of Rhythm |  | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Incorrect Measurement |  | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Incorrect Measurement |  | Incomplete | Additional Surgical Procedure | 3
CRM | Pacer | Incorrect Measurement |  | Incomplete or Misleading Diagnostic Data | Additional Surgical Procedure | 3
CRM | Pacer | Incorrect Measurement |  | Incomplete or Misleading Diagnostic Data | Intervention or Procedure | 1
CRM | Pacer | Incorrect Measurement |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | Incorrect Measurement |  | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Incorrect Measurement |  | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Incorrect Software Programming Calculations |  | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Incorrect Software Programming Calculations |  | Incomplete or Misleading Diagnostic Data | Additional Surgical Procedure | 3
CRM | Pacer | Incorrect Software Programming Calculations |  | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Intermittent Continuity |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | Invalid Sensing |  | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Invalid Sensing |  | Inappropriate Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Invalid Sensing |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | Invalid Sensing |  | Loss of Defibrillation Therapy | Arrhythmia | 2
CRM | Pacer | Loose Or Intermittent Connection |  | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Loose Or Intermittent Connection |  | Loss of Cardiac Pacing | Arrhythmia | 5
CRM | Pacer | Loose Or Intermittent Connection |  | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Loose Or Intermittent Connection |  | Loss of Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Loss Of Data |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | Low Audible Alarm |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | Low Battery |  | Loss of Cardiac Pacing | Arrhythmia | 5
CRM | Pacer | Low Battery |  | Loss of Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | Maintenance Does Not Comply To Manufacturers Recommendations |  | Abnormal Misuse | Death | 5
CRM | Pacer | Maintenance Does Not Comply To Manufacturers Recommendations |  | Device Tissue Pressure | Additional Surgical Procedure | 3
CRM | Pacer | Maintenance Does Not Comply To Manufacturers Recommendations |  | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Manufacturing Or Shipping Issue Associated With Device |  | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Manufacturing Or Shipping Issue Associated With Device |  | Device Tissue Pressure | Pocket Erosion | 3
CRM | Pacer | Material Deformation |  | Loss of Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Material Separation |  | Loss of Cardiac Pacing | Arrhythmia | 5
CRM | Pacer | Material Separation |  | Loss of Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | Material Separation |  | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Material Separation |  | Loss of Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Misconnection |  | Application of Excessive Force | Prolonged Surgery | 1
CRM | Pacer | Misconnection |  | Device Heating | Burn, Thermal | 2
CRM | Pacer | Misconnection |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | Misconnection |  | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Misconnection |  | Loss of Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Missing Test Results |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | MRI Environment Compatibility Issue |  | Abnormal Cardiac Pacing | Arrhythmia | 4
CRM | Pacer | MRI Environment Compatibility Issue |  | Abnormal Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | MRI Environment Compatibility Issue |  | Abnormal Cardiac Pacing | Heart Failure | 3
CRM | Pacer | MRI Environment Compatibility Issue |  | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | MRI Environment Compatibility Issue |  | Abnormal Cardiac Pacing | Heart Failure | 2
CRM | Pacer | MRI Environment Compatibility Issue |  | Device Heating | Burn, Thermal | 2
CRM | Pacer | MRI Environment Compatibility Issue |  | Device Tissue Pressure | Injury | 2
CRM | Pacer | MRI Environment Compatibility Issue |  | Device Tissue Pressure | No consequences or impact to patient | 1
CRM | Pacer | MRI Environment Compatibility Issue |  | Inappropriate Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | MRI Environment Compatibility Issue |  | Inappropriate Defibrillation Therapy | Discomfort | 2
CRM | Pacer | MRI Environment Compatibility Issue |  | Incomplete or Misleading Diagnostic Data | Intervention or Procedure | 2
CRM | Pacer | MRI Environment Compatibility Issue |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | MRI Environment Compatibility Issue |  | Loss of Cardiac Pacing | Arrhythmia | 4
CRM | Pacer | MRI Environment Compatibility Issue |  | Loss of Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | MRI Environment Compatibility Issue |  | Loss of Cardiac Pacing | Additional Surgical Procedure | 3
CRM | Pacer | MRI Environment Compatibility Issue |  | Loss of Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | MRI Environment Compatibility Issue |  | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | MRI Environment Compatibility Issue |  | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | MRI Environment Compatibility Issue |  | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | MRI Environment Compatibility Issue |  | Loss of Defibrillation Therapy | Additional Surgical Procedure | 3
CRM | Pacer | MRI Environment Compatibility Issue |  | Unintentional Extra Cardiac Stimulation | Discomfort | 1
CRM | Pacer | No Pacing |  | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Overheating Of Device Or Device Component |  | Device Heating | Burn, Thermal | 3
CRM | Pacer | Overheating Of Device Or Device Component |  | Device Heating | Burn, Thermal | 2
CRM | Pacer | Oversensing |  | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Oversensing |  | Abnormal Defibrillation Therapy | Arrhythmia | 2
CRM | Pacer | Oversensing |  | Inappropriate Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Oversensing |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | Oversensing |  | Loss of Cardiac Pacing | Arrhythmia | 5
CRM | Pacer | Oversensing |  | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Oversensing |  | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Pacer Found In Back Up Mode |  | Abnormal Cardiac Pacing | Additional Surgical Procedure | 3
CRM | Pacer | Pacer Found In Back Up Mode |  | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Pacer Found In Back Up Mode |  | Abnormal Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Pacer Found In Back Up Mode |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | Pacer Found In Back Up Mode |  | Unintentional Extra Cardiac Stimulation | Additional Surgical Procedure | 3
CRM | Pacer | Pacer Found In Back Up Mode |  | Unintentional Extra Cardiac Stimulation | Discomfort | 2
CRM | Pacer | Pacing Asynchronously |  | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Pacing Inadequately |  | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Pacing Inadequately |  | Abnormal Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Pacing Inadequately |  | Inappropriate Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Pacing Inadequately |  | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Parameter Calculation Error Due To Software Problem |  | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Parameter Calculation Error Due To Software Problem |  | Abnormal Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Physical Property Issue |  | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Poor Quality Image |  | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Premature Discharge of Battery |  | Inappropriate Programmed Settings | Intervention or Procedure | 2
CRM | Pacer | Premature Discharge of Battery |  | Loss of Cardiac Pacing | Arrhythmia | 4
CRM | Pacer | Premature Discharge of Battery |  | Loss of Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | Premature Discharge of Battery |  | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Premature Discharge of Battery |  | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Premature Discharge of Battery |  | Loss of Defibrillation Therapy | Additional Surgical Procedure | 3
CRM | Pacer | Problem With Software Installation |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | Product Quality Issue |  | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Product Quality Issue |  | Abnormal Cardiac Pacing | Discomfort | 1
CRM | Pacer | Product Quality Issue |  | Device Heating | Burn, Thermal | 2
CRM | Pacer | Product Quality Issue |  | Inappropriate Defibrillation Therapy | Discomfort | 2
CRM | Pacer | Product Quality Issue |  | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Product Quality Issue |  | Loss of Patient Notification | Intervention or Procedure | 2
CRM | Pacer | Product Quality Issue |  | Unintentional Extra Cardiac Stimulation | Discomfort | 2
CRM | Pacer | Product Quality Issue |  | Unintentional Extra Cardiac Stimulation | Discomfort | 1
CRM | Pacer | Programming Issue |  | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Programming Issue |  | Abnormal Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Programming Issue |  | Confusion | Arrhythmia | 2
CRM | Pacer | Programming Issue |  | Inappropriate Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Programming Issue |  | Inappropriate Defibrillation Therapy | Discomfort | 2
CRM | Pacer | Programming Issue |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | Programming Issue |  | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Radiation Overexposure |  | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Radio Signal Problem |  | Abnormal Cardiac Pacing | Additional Surgical Procedure | 3
CRM | Pacer | Radio Signal Problem |  | Inappropriate Defibrillation Therapy | Discomfort | 2
CRM | Pacer | Radio Signal Problem |  | Inappropriate Programmed Settings | Arrhythmia | 2
CRM | Pacer | Radiofrequency Interference (RFI) |  | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Sensing Intermittently |  | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Sensing Intermittently |  | Loss of Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Shielding Failure |  | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Shielding Failure |  | Abnormal Defibrillation Therapy | Arrhythmia | 2
CRM | Pacer | Shielding Failure |  | Loss of Cardiac Pacing | Arrhythmia | 5
CRM | Pacer | Therapeutic Or Diagnostic Output Failure |  | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Unauthorized Access To Computer System |  | Abnormal Cardiac Pacing | Arrhythmia | 5
CRM | Pacer | Unauthorized Access To Computer System |  | Inappropriate Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Unauthorized Access To Computer System |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | Unauthorized Access To Computer System |  | Loss of Cardiac Pacing | Arrhythmia | 5
CRM | Pacer | Unauthorized Access To Computer System |  | Loss of Cardiac Pacing | Additional Surgical Procedure | 3
CRM | Pacer | Unauthorized Access To Computer System |  | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Undersensing |  | Abnormal Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | Undersensing |  | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Undersensing |  | Inappropriate Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Undersensing |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | Undersensing |  | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Undersensing |  | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Unexpected Therapeutic Results |  | Inappropriate Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Unstable Capture Threshold |  | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Unstable Capture Threshold |  | Loss of Cardiac Pacing | Arrhythmia | 5
CRM | Pacer | Unstable Capture Threshold |  | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Unstable Capture Threshold |  | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Use Of Device Issue |  | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Use Of Device Issue |  | Abnormal Cardiac Pacing | Prolonged Surgery | 1
CRM | Pacer | Use Of Device Issue |  | Inappropriate Defibrillation Therapy | Discomfort | 2
CRM | Pacer | Use Of Device Issue |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | Use Of Device Issue |  | Loss of Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Use Of Device Issue |  | Surgical Complications | Hematoma | 2
CRM | Pacer | Use Of Incorrect Control Settings |  | Abnormal Cardiac Pacing | Arrhythmia | 3
CRM | Pacer | Use Of Incorrect Control Settings |  | Abnormal Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Use Of Incorrect Control Settings |  | Abnormal Cardiac Pacing | Discomfort | 2
CRM | Pacer | Use Of Incorrect Control Settings |  | Abnormal Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Use Of Incorrect Control Settings |  | Abnormal Cardiac Pacing | Discomfort | 1
CRM | Pacer | Use Of Incorrect Control Settings |  | Abnormal Defibrillation Therapy | Arrhythmia | 2
CRM | Pacer | Use Of Incorrect Control Settings |  | Inappropriate Defibrillation Therapy | Arrhythmia | 3
CRM | Pacer | Use Of Incorrect Control Settings |  | Inappropriate Defibrillation Therapy | Discomfort | 2
CRM | Pacer | Use Of Incorrect Control Settings |  | Incomplete or Misleading Diagnostic Data | Arrhythmia | 2
CRM | Pacer | Use Of Incorrect Control Settings |  | Inconvenience | No consequences or impact to patient | 1
CRM | Pacer | Use Of Incorrect Control Settings |  | Loss of Cardiac Pacing | Additional Surgical Procedure | 3
CRM | Pacer | Use Of Incorrect Control Settings |  | Loss of Cardiac Pacing | Arrhythmia | 2
CRM | Pacer | Use Of Incorrect Control Settings |  | Loss of Cardiac Pacing | Heart Failure | 2
CRM | Pacer | Use Of Incorrect Control Settings |  | Loss of Defibrillation Therapy | Arrhythmia | 5
CRM | Pacer | Use Of Incorrect Control Settings |  | Unintentional Extra Cardiac Stimulation | Discomfort | 2
CRM | Pacer | Wireless Communication Issue |  | Inaccessible Diagnostic Data | Additional Surgical Procedure | 3
CRM | Pacer | Wireless Communication Issue |  | Inconvenience | No consequences or impact to patient | 1
Note on A1: RDQCC: Therapy; Diagnostic
Note on B1: RDQCC: Therapy: CRM (pacer, ICD) muscle_stimulator ventilator infusion_pump diebete_pump